Microchimica Acta (2021) 188: 184
https://doi.org/10.1007/s00604-021-04835-9
A voltammetric sensor based on mixed proton-electron conducting composite including metal-organic framework JUK-2
for determination of citalopram
Maria Madej1	& Dariusz Matoga2 & Klaudia Skaźnik1 & Radosław Porada3 & Bogusław Baś 3 & Jolanta Kochana1
Received: 9 February 2021 / Accepted: 23 April 2021 / Published online: 11 May 2021
# The Author(s) 2021
Abstract
A voltammetric sensor has been developed based on glassy carbon electrode (GCE) modification with nanocomposite consisting of manganese-based metal-organic framework (JUK-2), multi-walled carbon nanotubes (MWCNTs), and gold nanoparticles (AuNPs) for detection of citalopram (CIT). The composition and morphology of JUK-2-MWCNTs-AuNPs/GCE were characterized by X-ray powder diffraction (XRD), Fourier-transform infrared spectroscopy (FT-IR), energy dispersion spectroscopy (EDS), and scanning electron microscopy (SEM). The electrochemical properties investigated using electro- chemical impedance spectroscopy (EIS) and cyclic voltammetry (CV) indicated that the fabricated hybrid material exhibits the properties of mixed ion-electron conductor (MIEC). Using staircase voltammetry (SCV), under optimized conditions, the fabricated sensor shows a linear response in three CIT concentration ranges, 0.05–1.0, 1.0–10.0 and 15.0–
115.0 μmol L−1, with a detection limit of 0.011 μmol L−1. The JUK-2-MWCNTs-AuNPs/GCE was successfully employed
for the determination of CIT in pharmaceutical, environmental waters, and biological samples with satisfactory recoveries
(98.6–104.8%).
Keywords Citalopram . Metal-organic framework . Electrochemical sensor . Mixed ion-electron conductor . Multi-walled carbon nanotubes . Gold nanoparticles
Introduction
(CIT), as the SSRIs representative, occurs in commercially available formulations as a racemic mixture, but its pharma- cological effect is related to the S-enantiomer. CIT is metab- olized mainly in the liver via oxidation and excreted in urine, with approx. 20% in unmetabolized form [2, 3]. Numerous studies prove the presence of antidepressants, especially citalopram, venlafaxine, sertraline, and fluoxetine, in various elements of the environment, such as surface water, municipal sewage, or bottom sediments at the concentration level
reaching μg L−1 [4, 5]. Due to the relatively high persistence in the aquatic environment, even such small amounts of psy-
choactive substances pose a risk of bioaccumulation in aquatic organisms and may affect their basic life functions [6]. As a result, many analytical methods have been introduced to de- termine CIT in tablets, water, and biological samples, most notably chromatographic [7] and capillary electrophoretic [8]. Also, a few works based on voltammetric techniques have been reported for CIT determination using bulk glassy carbon electrode [9] or electrodes modified with carbon nanostruc- tures, metal or metal oxide nanoparticles, as well as conducting polymers [10–13].
Antidepressants are a group of compounds worthy of atten- tion, due to a constant increase in their consumption in most well-developed countries [1] and hence rising environmental pollution caused by these active substances and their metabo- lites. One of the most prescribed antidepressants belongs to selective serotonin reuptake inhibitors (SSRI), whose mode of action is based on blocking the uptake of serotonin, thus in- ducing the neurotransmission between neurons. Citalopram

* Maria Madej marysia.madej@doctoral.uj.edu.pl
1
Faculty of Chemistry, Department of Analytical Chemistry, Jagiellonian University, Gronostajowa 2, 30-387 Kraków, Poland
Faculty of Chemistry, Department of Inorganic Chemistry, Jagiellonian University, Gronostajowa 2, 30-387 Kraków, Poland
Faculty of Materials and Ceramics, Department of Analytical Chemistry, AGH University of Science and Technology, A. Mickiewicza 30, 30-059 Kraków, Poland
2
3
ORIGINAL PAPER

184 Page 2 of 15
Microchim Acta (2021) 188: 184
Metal-organic frameworks (MOFs) are relatively new class of materials, which belong to a wide group of porous coordi- nation polymers with unique, two- or three-dimensional struc- tures [14]. Their excellent physicochemical properties, such as
high surface area (up to 7800 m2 g−1 [15]), high porosity (with
pore volumes up to 93% of the total volume of the material
[16]), flexibility, and good thermal stability, made MOFs use- ful as drug carriers, heterogeneous catalysts, materials used for gas storage and separation of hazardous compounds [14]. Because of these attractive features, metal-organic frame- works have been also explored for the development of sensors providing a very good alternative to carbon materials, molec- ularly imprinted polymers, and zeolites commonly used in this field [17, 18].
A significant limitation of using MOFs to construct elec- trochemical sensors is that most of these materials are elec- trically insulating, as composed of hard metal centers and redox-inactive organic linkers with hard donor atoms that cannot provide efficient electron transfer pathways. The improvement of electric conductivity properties of MOF- based sensors can be achieved through a variety of strate- gies, including incorporation of guest molecules, doping with redox-active compounds (like ferrocene, iodine or me- tallic species), as well as hybridization with materials exhibiting high electron conductivity, like carbon nano- structures or polymers [19]. On the other hand, conductive MOFs have recently been developed as proton-conducting materials [20], which can also be used for the design of nanostructured electrodes and sensing applications. Finally, in general, mixed ionic-electronic conductors (MIECs), capable of conducting both electrons and ions, are materials that can serve a unique function for electro- chemical devices including sensors [21]. Since pristine MIEC materials are rare, they typically come in blended architectures, composed of two types of conductors such as the PEDOT/PSS composite, used for sensitive sensing of glucose and other biomolecules from saliva [22]. To our knowledge, however, literature examples of using MOF- based MIEC composites for electrochemical sensing are lacking.
In this work, we demonstrate using a proton-conducting MOF for the preparation of a MIEC composite for electro- chemical sensing application. For this purpose, we selected
{(NH4)2[Mn(ina)2(NCS)2]}n·xH2O (JUK-2; Hina =
isonicotinic acid) which exhibits two-dimensional layered mi- croporous structure with interlayer separation of 9.63 Å,
solvent-free mechanochemistry has increasingly become a green alternative for the preparation of various materials including MOFs [24]. The mechanochemical preparation of JUK-2 relies on grinding a non-conducting 2D MOF precursor, {[Mn2(ina)4(H2O)2]·2EtOH}n (JUK-1) [25] with ammonium thiocyanate. This solid-state reaction involves
the NCS− complexation, a significant rearrangement of intralayer coordination bonds (unzipping of JUK-1 bilayers
to JUK-2 monolayers) and introducing of extraframework NH4+ ions as proton carriers [26].
The main aim of this work was to develop a simple voltammetric sensor employing the manganese-based met- al-organic framework (JUK-2) for CIT determination. The electrode composite, consisting of layered proton- conducting JUK-2 and nanomaterials exhibiting good elec- trical properties ( multi-walled carbon nanotubes (MWCNTs) and gold nanoparticles (AuNPs)), demonstrat- ed mixed proton-electron conductivity. Optimization of electrode-modifying mixture composition and measure- ment conditions allowed to develop a sensor exhibited good analytical parameters: high sensitivity, wide linear ranges and low detection limit. The usefulness of proposed JUK-2-MWCNTs-AuNPs/GCE was confirmed by CIT quantification in pharmaceutical, environmental waters, and biological samples.
To our knowledge, the application of MOF-based sensor for citalopram determination, as well as for any other antide- pressant, has not been reported in literature so far.
Experimental
Apparatus
For voltammetric measurements, standard three-electrode sys- tem, consisting of Pt wire as the auxiliary electrode, silver chloride electrode (Ag/AgCl, 3 M KCl, Mineral, Poland) as a reference electrode, and bare or modified glassy carbon electrode (GCE, φ = 3 mm, BASi, USA) as the working elec- trode, was utilized and controlled by the M161 multipurpose electrochemical analyzer combined with M164 electrode stand (both mtm-anko, Poland) with measurement accuracy of ±1 mV for potential and ± 0.1 nA for current. The imped- ance spectra were recorded with the use of μAUTOLAB III analyzer (EcoChemie, the Netherlands) with NOVA 2.0 soft- ware. The phase composition of optimal hybrid material was verified by measuring XRD patterns on a Rigaku Miniflex
600 diffractometer with Cu-Kα radiation (λ = 1.5418 Å) in a 2θ range from 3° to 80° with a 0.05° step and 0.5° min−1 scan speed. Scanning electron microscope Tescan Vega3
(ThermoFisher Scientific, USA) with energy-dispersive X- ray spectroscopy EDS attachment was used for the surface and elemental analysis of GCE covered with final
high proton conductivity above 4 × 10−4 S	(in the 70–90% relative humidity range at 25 °C), and good stabil-
ity in aqueous media [23]. Interestingly, JUK-2 can only be obtained by a solvent-free mechanochemical method, and it
cm−1
is achievable neither by conventional solvothermal methods nor by immersion in solutions. Generally, in con- trast to costly and non-ecological syntheses in solution, a

Microchim Acta (2021) 188: 184
Page 3 of 15 184
nanocomposite. FTIR spectra were recorded on a Thermo Fisher Scientific Nicolet iS10 FT-IR spectrophotometer equipped with an iD7 diamond ATR attachment.
were mixed with distilled water and sonicated with an ultra- sonic bath for 15 min. After that, the appropriate volume of the obtained MWCNTs suspension was transferred to the Eppendorf containing JUK-2 and sonicated again for 15 min. In the case of the final nanocomposite, an appropriate volume of AuNP suspension was added to the suspension consisting of JUK-2 and MWCNTs and homogenized using vortex for 5 min. All tested suspensions were stored in room temperature without access to light.
The GC electrode modification was preceded by polishing its surface on a polishing cloth (Leco, USA) with 0.05-μm alumina suspension (Buechler, USA) for 30 s and rinsing with strong stream of distilled water. Cleaned GCE was air dried, and then 5.0 μL of tested one- or multi-component suspen- sion, previously homogenized for 5 min using vortex, was applied on the surface of the GCE and left to dry at room temperature.
Reagents and materials
The JUK-2 MOF was synthesized by a mechanochemical method according to the procedure described in [23]. Multi-walled carbon nanotubes (length: 0.1–10 μm; inter- nal diameter: 2–6 nm) were purchased from Sigma- Aldrich (USA). The gold nanoparticle (AuNPs) suspen- sion (average diameter ~ 80 nm; 4.8 × 109 particles
mL−1) was prepared by tetrachloroauric acid reduction with citric acid, according to the procedure described in [27]. The
CIT standard stock solution (1.0 mmol L−1) was obtained by dissolving of an appropriate amount of citalopram
hydrobromide (LGC, Great Britain) in water and then stored in the fridge. Acetic acid, ascorbic acid, disodium hydrogen phosphate, potassium chloride, potassium hexacyanoferrate(II), potassium hexacyanoferrate(III), and sodium dihydrogen phos- phate were obtained from Chempur (Poland). Ammonia, am- monium chloride, and sodium acetate were obtained from Merck (Germany). The phosphate, acetate, and ammonium buffer solutions were prepared as described in [9]. Glucose (Micropharm, Poland), hexadecyltrimethylammonium bromide, lactose, magnesium stearate, sodium dodecyl sulfate, starch, talc, titanium dioxide, and Triton X-100 (Sigma-Aldrich, USA) were used for interference studies. Medicaments contain- ing citalopram hydrobromide Oropram (Pluspharma, Austria) and Cital (Glenmark, India) were purchased in local pharmacy. Certified reference materials (CRMs) of synthetic urine Seronorm Trace Elements Urine (Sero, Norway), synthetic se- rum Nortrol (Thermo Scientific, Finland), Surface Water (SPS– SW1 Batch no. 114), and Waste Water (SPS–WW2 Batch no.
113) (Spectrapure Standards AS, Norway) were used as samples with complex matrix. For that purpose, also water sample from Vistula River was collected in Krakow. All aqueous solutions were prepared with chemicals of analytical grade purity and distilled water from the HLP 5 system (Hydrolab, Poland).
Procedure of measurements
The performance of bare GCE, JUK-2/GCE, MWCNTs/ GCE, JUK-2-MWCNTs/GCE, and JUK-2-MWCNTs-
AuNPs/GCE was verified on the basis of cyclic voltammetry measurements conducted for [Fe(CN)6]3-\4- reversible redox system in 0.1 mol L−1 KCl, with potential range between
−0.4 and 0.8 V, using the scan rate of 0.05 V s−1. The CV
experiments were also performed for CIT in 0.1 mol L−1
phosphate buffer solution (PBS) pH 7 in the potential range 0.3–1.2 V, scan rate 0.1 V s−1, accumulation time (tacc) of 60 s, and accumulation potential (Eacc) of 0 V. The electro-
chemical behavior of CIT on JUK-2-MWCNTs-AuNPs/ GCE was examined under the same conditions.
The impedance spectra were recorded in the frequency range 100 kHz – 25 mHz, using sinusoidal signals with an amplitude of 10 mV imposed on the formal potential of the
redox pair [Fe(CN)6]3−/4- at concentration of 0.2 mmol L−1 in
0.1 mol L−1 KCl solution. To obtain the key figures of merit,
the equivalent electrical circuit was fitted to every recorded
impedance spectrum.
For quantitative determination of CIT, staircase voltamm- etry in both anodic and cathodic direction from 0.3 to 1.2 V was applied, with potential step (Es) 8 mV and step width (ts) 80 ms (60 ms of waiting time and 20 ms of current sampling
time), which corresponded to the scan rate of 0.1 V s−1. Before
each measurement, Eacc of 0 mV for 180 s was applied to the
working electrode. For each CIT concentration, three voltam- mograms were recorded, and for data analysis, the mean value of peak current was taken into consideration.
Standard procedures
Metal-organic framework-based sensors fabrication
Due to the high stability of JUK-2 in an aqueous environment, all suspensions used for GCE surface modification were pre- pared with distilled water as a solvent. For that purpose, one-
component suspensions containing 4 mg mL−1 of JUK-2 and 1 mg mL−1 of MWCNTs were prepared by weighting appro- priate amounts of each nanomaterial, mixing them with 2 mL
of distilled water and then homogenized with the use of ultra- sonic bath for 15 min. For the preparation of the composite consisting of both JUK-2 and MWCNTs, carbon nanotubes
Samples preparation
The stock solution with expected CIT concentration of
1.250 mmol L−1 for each commercial medicament was

184 Page 4 of 15
Microchim Acta (2021) 188: 184
prepared according to the procedure described in previous work [9]. For determination of CIT, 40.0 μL of selected stock solution was added to 10 mL of 0.1 mol L−1 PBS pH 7, and
then measurements were conducted according to the standard
addition calibration method. For each concentration, three subsequent voltammograms were recorded, and after back- ground subtraction, the mean value of achieved peak current values was calculated.
The sample of river water was subjected to filtration and stored in the fridge before analysis. For river water samples and CRMs of surface water and waste water, SCV measure- ments were performed, and no signals indicating the presence
of the analyte in the samples were observed. Thus, they were enriched with CIT, to obtain concentrations of 5.0 μmol L−1, diluted twice with the supporting electrolyte and analyzed
analogously as pharmaceutical samples.
The synthetic urine and blood CRMs were spiked with CIT in order to obtain concentration of 100.0 and 400.0 μmol L−1, respectively. For SCV determination, 50 μL of spiked urine or
12.5 μL of spiked serum was added to the 0.1 mol L−1 PBS pH 7, and then measurements were conducted as described for
previous samples.
component suspensions. Accordingly, the deposition of a composite from multi-component suspension seemed to be the best way for sensor preparation.
Electrochemistry of studied modified electrodes
The electrochemical impedance spectroscopy (EIS) and cyclic voltammetry (CV) were used to describe the surface modifi- cation of GCE with selected nanomaterials, namely, JUK-2, MWCNTs, and AuNPs. In order to investigate the influence of particular nanocomponents on the metrological parameters of the working electrode, the Nyquist plots for bare GCE, JUK-2/GCE, MWCNTs/GCE, JUK-2-MWCNTs/GCE, MWCNTs-AuNPs/GCE, and JUK-2-MWCNTs-AuNPs/
GCE in 0.1 mol L- 1 KCl solution containing 0.2 mmol L−1 [Fe(CN)6]3−/4- were recorded (Fig. 1a).
Based on the obtained EIS spectra, the characteristic pa-
rameters such as the charge transfer resistance (Rct), Warburg impedance modulus (σ), differential capacity of the double layer (Cdl), as well as heterogeneous rate constant of the elec- trode reaction (ks) can be determined. For the impedance spec- tra recorded at the formal potential of a given depolarizer, the charge-transfer resistance and Warburg impedance modulus are expressed by the following formulas [28]:
Results and discussion
1=4
2ðDODRÞ
RT
Choice of materials
Rct ¼ n2 F2A k
ð1Þ

ð2Þ

1=2
1=2
DO C þ D	CR
eff s
O
R
At the stage of an electrochemical sensor constructing, the selection of materials used for electrode modification plays a crucial role for achieving high sensitivity and stability of its performance. Since the goal of our research was to develop a MOF-based sensor, various MOFs were first investigated with regard to its stability in aqueous media, durability on the GCE surface, and its electrochemical performance in both model
redox probe [Fe(CN)6]3−/4- and CIT solution. Out of several MOFs, with different active centers (e.g., Mn2+, Cd2+, Zn2+)
as well as various organic linkers, JUK-2 exhibited the best performance, due to its proton-conductive properties, high stability in the aquatic media, and ability to form a highly homogeneous layer on the GCE surface. Since the JUK-2 MOF itself shows neither electron conductivity nor electrocat- alytic properties, it was combined with materials that exhibit such features. For this purpose, components like mesoporous carbon CMK-3, MWCNTs, Nafion®, and AuNPs were inves- tigated. From all the tested materials, MWCNTs and AuNPs gave the best results, exhibiting extraordinary properties of a hybrid composite, which were not observed with the other tested materials. Thus, they were used in further investiga- tions. Additionally, various methods of materials deposition onto the GCE surface were tested, including the deposition of a composite from one suspension containing several materials or the deposition of materials separately from single-
RT
4
σ ¼
Aeff p2ﬃﬃﬃ
2
1=2
1=2
n
2
F
DO C þ D	CR
O
R
where Ris the universal gas constant (R = 8.314 J mol−1 K−1), T is the temperature, n is the number of the exchanged elec- trons, F is the Faraday constant (F = 96,485.3 C mol−1), Aeff is the effective surface area, ksis the heterogenic rate constant, D
is the diffusion coefficient and c is the concentration of the
reduced (c ) and oxidized form (c ).
R
O
By dividing these equations, the expression for the heterogenic rate constant ks can be obtained:
1=4
σ	ðDODRÞ
ks ¼ R	pﬃﬃﬃ
ð3Þ
2
ct
According to literature, diffusion coefficients DO and DR
for the [Fe(CN)6]4−/3- redox system are equal 7.2 ×
10−6 cm2 s−1 [29]. EIS characteristic parameters determined based on the above-mentioned equations were summarized in
Table 1.
The relatively small dispersion of σ, as well as characteris-
tic, asymmetric shape of redox peaks for model probe [Fe(CN)6]3−/4- observed on CV curves, indicate that for all tested electrodes, mass transfer is controlled by linear diffu- sion. The value of Cdl for GCE is typical for 0.1 mol L−1 KCl,

Microchim Acta (2021) 188: 184
Page 5 of 15
184
0.1 V s−1, Eacc = 0 mV, and tacc = 60 s. The electrodes were marked as follows: GCE, black line; JUK-2/GCE, orange line; MWCNTs/GCE, red
line; JUK-2-MWCNTs/GCE, blue line; MWCNTs-AuNPs/GCE, violet line; JUK-2-MWCNTs-AuNPs/GCE, green line
Fig. 1 The impedance spectra (a) recorded for [Fe(CN)6]3-\4- (0.2 mmol L−1) in 0.1 mol L−1 KCl solution and corresponding cyclic voltammograms (b) (scan rate 0.05 V s−1); (c) CV curves recorded for
0.1 mmol L−1 of CIT in PBS (pH 7, 0.1 mmol L−1) with scan rate
which proves that no additional adsorption occurs at bare electrode. The recorded EIS spectra show that the modifica- tion of GCE surface with JUK-2 contributes to a distinct in- crease of Rct and a decrease of rate constant ks, which proves that the employed metal-organic framework hinders electron transfer due to its poor electron conductivity. Moreover, JUK- 2 causes a significant increase in Cdl, which is a highly unde- sirable effect. The modification of GCE with MWCNTs re- sults in a slight increase of Rct and a decrease of ks, but to a much lesser extent than in the case of JUK-2. Compared with the JUK-2/GCE, the electrode modified with both JUK-2 and MWCNTs is characterized by much better metrological pa- rameters, thus confirming the legitimacy of using JUK-2 as an
electrode modifier. However, such a two-component modify- ing layer causes a significant increase in the Cdl, and conse- quently the growth of capacitive current, which is the main component of the background current in voltammetric tech- niques. The introduction of AuNPs into the modifying layer lowers the Cdl value twice, which is undoubtedly a favorable phenomenon. In addition, the electrode prepared in this way has the lowest Rct value (approx. 2.5 times lower compared to the unmodified GCE), which is also reflected in a high value of the constant ks. Meanwhile, the MWCNTs-AuNPs/GCE exhibits a very high Cdl value and a moderate ks, lower than for JUK-2-MWCNTs/GCE and JUK-2/MWCNTs-AuNPs/ GCE, which confirms the crucial role of JUK-2 in the
Table 1 The comparison of tested electrodes based on EIS measurements (n = 3)
Cdl/μF cm−2
ks/× 10−5 m s−1
Rct/kΩ
σ/kΩ s-1/2
Electrode
GCE
JUK-2/GCE MWCNTs/GCE
JUK-2-MWCNTs/GCE MWCNTs-AuNPs/GCE
JUK-2-MWCNTs-AuNPs/GCE
1.25 ± 0.05
14.7 ± 0.3
2.73 ± 0.15
2.50 ± 0.05
3.00 ± 0.50
0.50 ± 0.01
9.6 ± 0.2
10.5 ± 0.1
8.8 ± 0.1
10.4 ± 0.2
11.1 ± 0.4
9.0 ± 0.1
11.5 ± 0.3
278 ± 18
10.5 ± 0.7
2220 ± 40
5480 ± 80
1030 ± 25
(14.6±1.0)
(1.30±0.08)
(6.18±0.06)
(7.89±0.05)
(6.96±0.05)
(34.2±0.8)

184 Page 6 of 15
Microchim Acta (2021) 188: 184
constructed sensor. The obtained results undoubtedly prove that the combination of JUK-2 with MWCNTs and AuNPs improves the metrological parameters and electrical properties of the bare GCE and provides the most favorable signal-to- noise (S/N) ratio.
In order to confirm the electrode behavior observed in EIS measurements, CVs in 0.1 mol KCl containing
1.0 mmol L−1 of [Fe(CN)6]3−/4- were recorded (see Fig. 1b). Based on Randles-Ševčík equation (Eq. 4), the electrochemi-
cal active surface area of each tested electrode was calculated [30, 31]:
outstanding increase of recorded current and thus higher sensitivity.
The performance of tested modified electrodes was inves- tigated by CV in 0.1 mol PBS (pH 7) containing 100 μmol L−1 of CIT (Fig. 1c). Before surface modification,
the GCE showed a poor, one anodic wave (40.3 μA cm−2) at
the potential of 0.91 V. For JUK-2/GCE, the oxidation peak current density slightly decreased to 38.5 μA cm−2 with minor shift of peak position towards positive potentials due to low
electric conductivity of the JUK-2. These observations are consistent with the JUK-2/GCE properties determined in EIS measurements. The modification of electrode with MWCNTs resulted in more than twofold increase of recorded
anodic current density (106.1 μA cm−2) accompanied by a
shift of the peak potential in the cathodic direction up to
0.82 V. Such a large shift of the CIT oxidation peak indicates a significant reduction of the activation energy of the drug electrooxidation process, and thus the creation of many new active centers on the sensor surface. This favorable effect is amplified by the decoration of MWCNTs with Au nanoparti- cles. However, both electrodes, i.e., MWCNTs/GCE and MWCNTs-AuNPs/GCE, are characterized by poor reproduc- ibility and low mechanical durability. Further improvement of the sensor performance was achieved by GCE modification with JUK-2-MWCNT composite. The observed favorable en- hancement of the signal at JUK2-MWCNTs/GCE is almost twice higher than the predicted growth resulting from the combination of individual components properties, which indi- cates the presence of synergistic effect of JUK-2 and MWCNTs. Among all, the best electrochemical and function- al parameters were obtained for JUK2-MWCNTs-AuNPs/ GCE, which showed a tenfold increase of anodic peak current density (402.6 μA cm-2) in comparison to bare GCE.
The excellent performance of JUK-2-MWCNTs-AuNPs/ GCE towards CIT could be contributed to several factors. The combination of proton conductive JUK-2 crystallites with electron conductive MWCNTs and AuNPs leads to enlarge- ment of the active surface area and formation of a hybrid composite, which acts as mixed ion-electron conductor (MIEC) [32]. Such a feature should be especially favorable for electrochemical processes in which electron transfer is accompanied by deprotonation, as in the case of CIT. This phenomenon could also explain the synergistic effect of JUK-2, MWCNTs, and AuNPs.
L−1
L−1
Ip ¼ (2:69⋅105)n3=2AelD1=2v1=2c
ð4Þ
where Ip is the peak current, Ael the electrochemical active
surface area of electrode, and v the scan rate (v = 0.05 V s−1), D = 7.2 × 10−6 cm2 s−1 [29].
A well-defined voltammogram with symmetric anodic and
cathodic peaks and the peak separation (ΔE) of 75 mV was observed at the bare GCE. After modification of electrode with JUK-2, the anodic and cathodic current density de- creased, with a growth of ΔE to 160 mV, which is justified by the low electrical conductivity of JUK-2 and the resulting significant increase in the values of Rct and Cdl (Table 1). The deposition of MWCNTs on the GCE caused a distinct growth of Rct due to the weak physical adhesion of MWCNTs to the electrode surface. Also, a slight increase of the modulus σ and capacity Cdl, as well as an almost twofold decrease of rate constant ks in reference to bare GCE was observed. As a result, for MWCNTs/GCE, a slight increase of redox peaks and a growth of ΔE to 116 mV, compared with bare electrode, was noticed. The JUK-2-MWCNTs/GCE showed a distinct enhancement of the anodic and cathodic peak current density (above 3 times compared with the bare GCE, JUK-2/GCE and MWCNTs/GCE). However, the high value of the differential capacity Cdl (almost 100 times higher than for the GCE) makes the S/N ratio unfavorable and the capacitive current is not compensated. The best electrical properties from all con- sidered sensors ensure the JUK-2-MWCNTs-AuNPs/GCE. For this electrode, the lowest Rct value and the highest value of the rate constant reaction ks were obtained; hence the best shaped voltammograms and the most favorable S/N ratio were achieved.
The calculated electrochemical active surface area of JUK-2-MWCNTs-AuNPs/GCE (0.19 cm2) is comparable to JUK-2-MWCNTs/GCE (0.20 cm2), ca. 50% higher than for MWCNTs-AuNPs/GCE (0.13 cm2) and more than twice as large as for MWCNTs/GCE (0.093 cm2), JUK-2/GCE (0.079 cm2), and bare GCE (0.070 cm2). The significant expansion of electrochemical active surface area of JUK- 2-MWCNTs-AuNPs/GCE, accompanied by facilitated electron transfer, enables to increase the efficiency of the occurring electrode reaction, which manifests in an
Composition and morphology of the electrode nanocomposite
The amounts of chosen components, JUK-2, MWCNTs, and AuNPs, played a significant role in the electrochemical per- formance of a modified electrode. Therefore, different mass ratios of JUK-2 and MWCNTs in composite suspensions (from 1:3 to 9:2) were evaluated on the basis of CV

Microchim Acta (2021) 188: 184
Page 7 of 15 184
measurements conducted in 0.1 mol L−1 KCl for increasing concentration of [Fe(CN)6]3−/4- (Fig. S1a). It was observed that the sensitivity of modified electrodes to the model redox
probe was increasing with a growing amount of JUK-2 rela- tive to MWCNTs, reached a maximum for mass ratio of 4:1 and then slightly decreased. As a result, the optimal suspen-
sion containing 4 mg mL−1 of the JUK-2 and 1 mg mL−1 of
MWCNTs was selected for subsequent studies.
After introduction of AuNPs (5% vol.) to the electrode nanocomposite, the sensitivity of measurements conducted for the model redox probe fell by about 12% and it remained constant regardless of the amount of AuNPs.
Hence, the measurements in the presence of target analyte CIT were performed. For that purpose, CV in 0.1 mol L−1 PBS (pH 7.0) for increasing concentration of CIT were
recorded (Fig. S1b). It is clearly visible that sensitivity for CIT determination is strongly affected by the amount of AuNPs. The highest sensitivity was obtained using elec- trode modified with nanocomposite containing 20% AuNPs by volume. In comparison to JUK-2-MWCNTs/ GCE, the addition of AuNPs resulted in 3.5 times improve- ment of sensitivity. Thus, the optimal nanocomposite
distinct elements (Mn and S, C, and Au) of each of the three phases (Fig. 2f).
In order to estimate the mass fraction of the individual components in the hybrid material, the approximate quantita- tive analysis of the composite content was investigated, on the basis of AuNP concentration and average nanoparticle size in initial AuNPs suspension, as well as JUK-2 and MWCNTs amount. Performed calculations suggest that optimal nano- composite contained 80% of JUK-2, 20% of MWCNTs, and 0.1% of AuNPs, respectively. Calculated percentage compo- sition of the optimal nanocomposite is consistent with the results obtained in EDS analysis.
Investigation of the electrode reaction mechanism
The electrode reactions are governed by either adsorption or diffusion-controlled processes and the systematic analysis of the impact of scan rate (ν) on the recorded peak current (Ip) can provide evidence to distinguish between these two pro- cesses. In order to understand the mechanism of CIT oxidation at JUK-2/MWCNTs-AuNPs/GCE, CV measurements at scan
rates ranging from 6.25 to 500 mV s−1 were performed (Fig. 3a). In the presence of 30 μmol of CIT, the
anodic peak current increased proportionally with scan rate, in accordance with regression equation of Ip (μA) = 0.227 ν – 3.852 (R2 = 0.992, see Fig. S2).
Furthermore, a relationship between logarithm of peak current on the logarithm of scan rate (Fig. 3b) exhibited an excellent linearity (R2 = 0.991) with a slope of 0.914, which is close to theoretical value of 1.0, corresponding to adsorption-controlled process [33].
In addition to peak current changes, it can be observed that the oxidation peak potential (Ep) shifts positively as the scan rate increases (Fig. 3a), which indicates the irreversibility of electrode reaction [34]. According to Laviron equation, for an irreversible, adsorption-controlled process, the dependence between Ep and scan rate can be described as follows [35]:
mL− 1
mL− 1
consisted of 4 mg
of JUK-2, 1 mg
of
L−1
MWCNTs and 20% vol. of AuNPs. No influence of AuNPs on the sensitivity of measurements conducted in the presence of Fe2+/Fe3+ in comparison to significant increase of sensitivity observed for CIT can be related to different reac- tion mechanisms. The CIT oxidation is controlled by adsorp- tion, and in this case, AuNPs could act as active centers on which CIT molecules easily adsorb, which distinctly improve
the sensitivity. Whereas for Fe[(CN)6]3−/4- both oxidation and
reduction processes are controlled by diffusion, so electrode
reactions do not run directly on the electrode surface; therefore the presence of AuNPs has no impact on their course.
The morphology of the final nanocomposite was specified using powder X-ray diffraction (XRD), scanning electron mi- croscopy (SEM), and energy dispersive spectroscopy (EDS). The XRD patterns of pristine JUK-2 and the JUK-2/ MWCNTs/AuNPs composite clearly demonstrate the coexis- tence of three components in the hybrid material (Fig. 2a). Apart from predominant characteristic peaks of JUK-2 phase in the composite, the extra diffraction peak at 26.0o indicates the presence of MWCNTs with d-spacing of 3.42 Å between graphene sheets. The third component is corroborated by the diffraction peak at 38.6o that corresponds to the (111) plane in the fcc structure of gold nanoparticles. IR spectroscopy re- veals that the vibrations of the skeletal atoms remained intact during the composite formation, confirming the retention of the layered JUK-2 structure. Moreover, the presence of a
strong CN band at 2100 cm−1 proves that thiocyanates remain coordinated as terminal ligands in JUK-2 (Fig. 2b). SEM im-
ages (Fig. 2c, d) and EDS elemental mapping (Fig. 2e) of the studied composite additionally confirm the presence of
 RT
E ¼ A þ	ln v
ð5Þ
p
ð1−αÞnF
where A is the constant value.
The value of (1 – α) n calculated from the slope of Ep versus ln v dependance (Fig. 3c) was equal to 0.955. For irreversible electrode processes, the charge transfer coefficient (α) is assumed to be 0.5 [35]; hence the number of electrons was estimated as 1.91, which indicates that two electrons are involved in the citalopram oxidation.
In order to investigate the effect of the pH on the oxidation of CIT at the developed sensor, CV were recorded in
0.1 mol L−1 acetate (AcBS), phosphate (PBS), and ammoni-
um (AmBS) buffer solutions with pH values ranging from 4.0
to 10.0 (Fig. 4). The obtained voltammograms demonstrate

184 Page 8 of 15
Microchim Acta (2021) 188: 184
Fig. 2 XRD patterns (a) and FT-IR spectra (b) for pristine JUK-2 (black) and the JUK-2-MWCNTs-AuNPs composite (red); SEM images of glassy carbon electrode surface modified with final nanocomposite layer
(c, d) with corresponding EDS mapping (e) and elemental composition estimated by EDS analysis (f)
that with increasing pH value, the anodic peak potential shifted linearly in negative direction, indicating the involve- ment of protons in the electrochemical reaction. The
relationship between oxidation peak potential and pH can be described as Ep (mV) = −59.5 pH + 1281.9 with R2 = 0.987. The achieved directional coefficient is close to theoretical

Microchim Acta (2021) 188: 184
Page 9 of 15
184
peak potential (c) performed in 0.1 mol L−1 PBS (pH 7) containing 30 μmol L−1 of CIT (Eacc = 0 V, tacc = 60 s); n =3
Fig. 3 CV curves recorded for scan rate values in the range of 6.25– 500 mV s−1 (a) and their impact on CIT anodic peak current (b) and
value of 59.0 mV pH −1, suggesting the participation of equal number of protons and electrons in CIT oxidation process
[36].
On the basis of obtained results and previous reports concerning the electrooxidation of amine derivatives, it can
be concluded that observed anodic peak is related to the irre- versible, adsorption-controlled oxidation of tertiary amine group present in CIT molecule through the formation of a cation radical which, after deprotonation and loss of two elec- trons, transforms into secondary amine and aldehyde [37, 38].
Fig. 4 CV curves (subjected to background reduction) recorded in various buffers (a) (each at
concentration of 0.1 mol L−1) containing 30 μmol L−1 of CIT (ν = 0.1 V s−1, Eacc =0 mV, tacc = 60 s) and the influence of buffer
pH on CIT oxidation peak poten- tial (inset). AmBS ammonium buffer solution, PBS phosphate buffer solution, AcBS acetate buffer solution

184 Page 10 of 15
Microchim Acta (2021) 188: 184
Analytical performance of the sensor
The created amine may undergo further oxidation under ap- propriate conditions, which may explain the presence of the second, residual anodic wave observed for high concentra- tions of CIT. These conclusions are consistent with previous reports describing the CIT oxidation at modified electrodes [11, 12]. The proposed mechanism for the electrode reaction is presented in Fig. S3.
Under the optimal experimental conditions, the sensitivity of CIT determination in 0.1 mol L−1 PBS (pH 7) at JUK-2/ MWCNTs-AuNPs/GCE was investigated by SCV (Fig. 5).
A good linear relation between the oxidation peak current and the CIT concentration was found. The anodic peak current of CIT increased linearly in three concentration ranges of
0.05–1.0, 1.0–10.0, and 15.0–115.0 μmol L−1, with sensitiv-
ities of 38.00 ± 0.79 μA L μmol−1 cm−2, 15.89 ±
0 . 6 8 μ A L μ m o l − 1 c m − 2 , a n d 3 . 3 2 ±
0.47 μA L μmol−1 cm−2 (mean values for 3 sensors), respec-
tively. The limit of detection (LOD), defined as 3.3 times
standard deviation of peak current recorded for
0.05 μmol L−1 CIT (n = 5) over the calibration slope, was calculated as 0.011 μmol L−1, whereas the limit of quantifica- tion (LOQ), evaluated as 3LOD, was equal to 0.033 μmol L−1 [39].
The reproducibility and stability of the proposed sensor were also examined. The relative standard deviation (RSD) of peak current values recorded in PBS containing 10 μmol L−1 of
CIT using six independently fabricated electrodes was 2.3%,
indicating good reproducibility. Additionally, the short-term stability of the sensor was investigated after storage at air at room temperature for 24 h (Fig. S5). After one day, the record- ed signal decreased to almost 50% of the initial peak current, which suggests that the sensor shows its best performance im- mediately after preparation. Despite the short stability of the sensor, it is noteworthy that the stability of the electrode mod- ifying suspension itself is very high. Even after 6 months, it allows for fabrication of sensors with excellent reproducibility (3.5%). Considering that, the electrode preparation procedure consists of simple and fast one-step modification by drop cast- ing of multi-component suspension onto GCE. The sensor can be easily fabricated just before analysis; thus its short stability does not appear to be a significant limitation.
Comparing to other electrochemical methods developed recently for CIT determination (Table 2), the proposed sensor is characterized by almost the lowest LOD value described in literature so far, with a wide linear range. In contrast to all the modified electrodes presented in Table 2, JUK-2/MWCNTs- AuNPs/GCE is easy to fabricate and does not require the use of harmful solvents, like toluene or DMF. Moreover, at all stages, from synthesis of both JUK-2 [23] and its precursor JUK-1 [25] to CIT quantification, only reagents neutral for environment are used, which promotes the principles of green analytical chemistry (GAC) [40].
Optimization of experimental parameters for quantitative analysis
The electrochemical processes involving organic compounds are mostly strictly related to the pH of supporting electrolyte. Hence, the influence of pH of different buffers (PBS, AmBS, AcBS) on the oxidation of CIT was investigated not only in terms of the peak position but also its height (Fig. 4). The recorded CV voltammograms demonstrate that the anodic peak current increases rapidly from pH 4.0 to a maximum at pH 7.0, and then it distinctly declines. The observed tendency is consistent with the proposed mechanism of CIT oxidation, which includes proton elimination, so in acidic media the electrode reaction should be obstructed [10]. The dissociation constant (pKa) of CIT is close to 9, so below pH 9 CIT mol- ecules are predominantly protonated, which facilitates the course of electrode reaction [11]. With further pH increase, the concentration of the protonated form declines, and thus the peak current also decreases. The above considerations confirm that pH 7 constitutes the best environment for CIT oxidation; therefore all quantitative electrochemical experiments were carried out in PBS with this pH value.
In order to develop a methodology for quantitative analysis of CIT, various voltammetric techniques, such as cyclic volt- ammetry (CV), staircase voltammetry (SCV), differential pulse voltammetry (DPV), and square wave voltammetry (SWV) were considered. For final procedure, SC voltammetry was selected, since this technique showed a higher sensitivity in comparison to CV and reduced capacitive current, but at the same time it is not a pulse technique such as DPV and SWV, whose application resulted in a drastic, 15-fold decrease in the measured current values, regardless of the used parameters (data not shown). This phenomenon is possibly related with the nature of the electrical double layer that forms at the phase boundary between modified electrode and solution. The opti- mal conditions for CIT determination using SCV were select- ed through the examination of the relationship between the recorded oxidation peak current and measurement parameters, such as potential step (Es), step width (ts), accumulation time (tacc) and accumulation potential (Eacc) (see Supplementary Material for details). The anodic peak with satisfactory height and good repeatability was obtained for Es of 8 mV and ts of
80 ms, which corresponds to scan rate equal 0.1 V s−1, with tacc equal to 180 s and Eacc of 0 mV. Therefore, these param-
eters were chosen for further studies.
Interference studies
The ability of the fabricated sensor for the determination of CIT in real samples such as environmental water, wastewater samples, and pharmaceuticals was evaluated. For that

Microchim Acta (2021) 188: 184
Page 11 of 15
184
Fig. 5 SC voltammograms and corresponding calibration graph (inset a) obtained for different concentrations of CIT in
0.1 mol L−1 PBS (pH 7) (black line), n = 3; inset b: magnified
voltammograms for 0.05–
1.0 μmol L−1 CIT). SCV param- eters: Es =8 mV, ts = 80 ms,
Eacc =0 mV, tacc = 180 s
purpose, a variety of possible interfering organic and inorgan- ic compounds (in concentration between 0.1–100 μmol L−1), such as ascorbic acid, glucose, lactose, titanium dioxide, talc,
starch, magnesium stearate, Triton X-100 (non-ionic surfac- tant), sodium dodecyl sulfate (SDS; anionic surfactant), and hexadecyltrimethylammonium bromide (CTAB; cationic sur- factant), were tested. The tolerance limit was set as a percent- age signal change less than ± 5% corresponding to the peak
current recorded for 10 μmol L−1 of CIT without and in the presence of an interferent.
The majority of tested interferents, regardless of its concen- tration, did not affect visibly the recorded signal for CIT (for more details see Table S2). Only the presence of magnesium
stearate, Triton X-100, and CTAB resulted in a decline of the measured peak current (Fig. S6a), without change of its shape and position. The suppressive influence of magnesium stea- rate on the recorded signal was observed for its concentration
higher than 10 μmol L−1. This slight signal drop (less than
12%) is most likely related to the presence of a solid phase in
the supporting electrolyte, which inhibits the course of the electrode reaction. This unfavorable effect, which can occur in pharmaceuticals analysis, can be easily eliminated by sub- jecting samples to filtration. In contrast, both Triton X-100 and CTAB strongly affect the recorded peak current, even at low concentrations of these surfactants. In the case of 10-fold excess of CTAB or Triton X-100 in comparison to CIT
Table 2 Characteristics of various modified electrodes implemented for CIT determination
Electrode
Technique
Linearity range/
μmol L−1
LOD/
μmol L−1
Reference
Fe3O4@[(EtO)3Si–L]/MWCNTs/GCE1
AuPdNPs-GE/AuE2
MWCNTs-pABSA-β-CD/GCE3
DPV
SWV DPV
0.3–100.0
0.5–50.0
0.09–1.0
1.0–11.0
11.0–100.0
0.012–1.54
0.05–1.0
1.0–10.0
15.0–115.0
0.053
0.049
0.044
[12]
[13]
[11]
ZnO-NPs /MWCNTs/CPE4
JUK-2-MWCNTs-AuNPs/GCE
SWV
SCV
0.005
0.011
[10]
This
work
1 Glassy carbon electrode chemically modified with multi-walled carbon nanotubes and core-shell functionalized magnetic nanoparticles of iron(II,III) oxide; 2 gold electrode modified with graphene decorated by gold-palladium nanoparticles; 3 glassy carbon electrode modified with multi-walled carbon nanotubes coated with poly p- aminobenzene sulfonic acid/β-cyclodextrin film; 4 carbon paste electrode modified with zinc oxide nanoparticles and multi-walled carbon nanotubes

184 Page 12 of 15
Microchim Acta (2021) 188: 184
concentration, the signal decreases even by 70%. Based on these observations, it can be concluded that the occurrence of cationic or non-ionic surfactants in highly polluted water samples could make the CIT determination difficult; however these surfactants usually occur in environmental waters at a nanomolar level [34, 41]; therefore the risk is relatively low. The influence of synthetic biological matrices on the re- corded signal was verified through testing of increasing per- centage content of the Blood Serum or Urine Certified Reference Material (CRM) in PBS volume (Fig. S6b). The obtained results clearly indicate the presence of strong inter- ferences even for small content of biological matrices, which at values of 5% (for urine) and 2% (for serum) lead to com- plete disappearance of the recorded peak for CIT. However, at low content of biological matrix (not exceeding 0.5% for urine, and 0.1% for serum), the decrease of the signal can be negligible; thus under such conditions, it should be possible to determine CIT, but at relatively high concentrations only (at
least 100 μmol L−1). In the case of samples with lower CIT concentration, an employment of extraction at a sample prep-
aration stage may be a good alternative for reducing the matrix impact.
The conducted studies indicate that the developed sensor is characterized by good selectivity and can be used for CIT determination in commercial medicaments and water samples without special sample pretreatement. The possible interfer- ence risk can be related to the presence of high quantities of non-ionic and cationic surfactants or components of biological matrices. However, these interferences should be compensat- ed by employment of adequate calibration approach, like the standard addition method or more complex approaches, e.g., generalized calibration strategy [42].
analyzed three times, and then the relative error (RE) with confidence intervals (CI) and RSD was calculated (Table 3). For both medicaments the RE value did not exceed 2.5%, and the RSD values were less than 0.5%, indicating good compli- ance between the obtained value and the amount declared by manufacturer. Thus, it can be concluded that the developed sensor provides a good alternative for real sample analysis of pharmaceutical tablets.
For Vistula water, Surface water, and Wastewater CRM samples, no target analyte was found; therefore a recovery test was conducted to assess sensor applicability. The obtained results are summarized in Table 4. The reasonable recovery (between 98.6 to 103.8%) with satisfactory RSD (below 1.0%) values suggest that the application of JUK-2-based sen- sor for CIT determination in environmental samples is char- acterized by good accuracy and excellent precision.
The JUK-2-MWCNTs-AuNPs/GCE was also successfully applied for synthetic urine and serum sample analysis. The determination of CIT in biological samples was performed in the same manner as for water analysis, and based on ob- tained results, the recovery and RSD values were calculated (Table 4). Using the previous interference tests performed for biological matrices (Fig. S6b), it was possible to estimate sample dilution rates, which allowed for interference effect minimization. Thus, the performed analyses are characterized by acceptable precision and accuracy, with recovery rates of
100.2 and 104.8%, with RSD values lower than 1.05%, indi- cating that the proposed method could be efficiently used for CIT analysis in samples with complex matrix.
Conclusions
Practical applications of JUK-2-MWCNTs-AuNPs/GCE
This work describes a fabrication of novel MOF-based sensor (JUK-2-MWCNTs-AuNPs/GCE) prepared by a simple one- step drop casting of a multicomponent suspension on the sur- face of a GCE. Our investigations proved that the JUK-2 MOF plays a crucial role in the unique performance of the devel- oped sensor. It perfectly binds MWCNTs and AuNPs, ensures a very high reproducibility and accuracy of performed
To evaluate the analytical applicability and reliability of the JUK-2-MWCNTs-AuNPs/GCE sensor was applied for CIT determination in tablets, surface and waste water CRMs, and synthetic biological samples. Two commercially available pharmaceuticals containing citalopram hydrobromide were
Table 3 Determination of CIT in pharmaceutical pills
Sample
Amount declared by manufacturer/mg
Amount determined/mg
Average amount*
±CI/mg
RE*
/%
RSD
/%
−2.2
Oropram
20.00
19.57
19.66
19.48
19.56
19.63
19.46
0.45
19.57 ± 0.15
−2.3
Cital
20.00
0.45
19.55 ± 0.14
*mean value (n = 3); CI confidence interval calculated for α = 0.05

Microchim Acta (2021) 188: 184
Page 13 of 15 184
Table 4 Determination of CIT in spiked water and biological samples
Sample
Amount added/μmol L−1
Amount determined/μmol L−1
Recovery*
/%
RSD
/%
Vistula water
5.00
4.97
4.93
4.89
5.15
5.19
5.15
5.22
5.17
5.19
99.8
99.4
101.4
417.5
418.2
421.4
98.6 ± 1.4
0.84
Surface water CRM
5.00
103.3 ± 0.7
0.42
Wastewater CRM
5.00
103.8 ± 0.8
0.47
Synthetic urine
100.0
100.2 ± 1.8
1.05
Synthetic serum
400.0
104.8 ± 0.9
0.49
*mean value (n = 3)
Klaudia Skaźnik—Methodology, investigation, validation, formal analysis, data curation, visualization
Radosław Porada—Methodology, investigation, validation, formal analysis, writing (original draft)
Bogusław Baś—Conceptualization, supervision, and writing (reviewing and editing)
Jolanta Kochana—Conceptualization, methodology, data curation, supervision, and writing (reviewing and editing)
measurements, and guarantees a favorable synergistic effect of the nanomaterials used. The coexistence of proton and electron conductivity in the tested nanocomposite, which acts as a mixed ion-electron conductor (MIEC), is most likely responsible for an excellent electrocatalytic activity of the proposed sensor towards CIT oxidation with a wide linear range, low detection limit, high selectivity and good reproducibility. It is noteworthy that at all stages of the developed method, including JUK-2 synthesis and sensor fabrication, only environmental friendly reagents and sol- vents were used. Moreover, the analysis was carried out in a neutral environment, which makes the proposed method fully compliant with the principles of GAC. Finally, the satisfactory results obtained for CIT determination in all tested samples (tablets, water samples, biological materi- al) clearly demonstrate the potential of the fabricated JUK-2-MWCNTs-AuNPs/GCE sensor for practical applications.
Funding M.M. and R.P. have been partly supported by The National Centre for Research and Development in Poland (Narodowe Centrum Badań i Rozwoju; NCBiR) - the EU Project POWR.03.02.00–00-I004/
16. D.M. gratefully acknowledges the National Science Centre (NCN, Poland) for the financial support (Grant no. 2019/35/B/ST5/01067).
Availability of data and material Not applicable.
Code availability Not applicable.
Declarations
Competing interests The authors have no conflicts of interest to declare that are relevant to the content of this article.
Supplementary Information The online version contains supplementary material available at https://doi.org/10.1007/s00604-021-04835-9.
Conflict of interest The authors declare that they have no com- peting of interests.
Acknowledgements M.M. and R.P. have been partly supported by The National Centre for Research and Development in Poland (Narodowe Centrum Badań i Rozwoju; NCBiR) - the EU Project POWR.03.02.00- 00-I004/16. D.M. gratefully acknowledges the National Science Centre (NCN, Poland) for the financial support (Grant no. 2019/35/B/ST5/ 01067).
Open Access This article is licensed under a Creative Commons Attribution 4.0 International License, which permits use, sharing, adap- tation, distribution and reproduction in any medium or format, as long as you give appropriate credit to the original author(s) and the source, pro- vide a link to the Creative Commons licence, and indicate if changes were made. The images or other third party material in this article are included in the article's Creative Commons licence, unless indicated otherwise in a credit line to the material. If material is not included in the article's Creative Commons licence and your intended use is not permitted by statutory regulation or exceeds the permitted use, you will need to obtain
Authors’ contributions Maria Madej—Conceptualization, methodolo- gy, investigation, validation, formal analysis, resources, data curation, visualization, writing (original draft), and writing (review and editing)
Dariusz Matoga—Conceptualization, methodology, investigation, formal analysis, resources, writing (original draft), supervision, and writ- ing (reviewing and editing)

184 Page 14 of 15
Microchim Acta (2021) 188: 184
permission directly from the copyright holder. To view a copy of this licence, visit http://creativecommons.org/licenses/by/4.0/.
15.
Hönicke IM, Senkovska I, Bon V, Baburin IA, Bönisch N, Raschke S, Evans JD, Kaskel S (2018) Balancing mechanical stability and ultrahigh porosity in crystalline framework materials. Angew Chem Int Ed 57:13780–13783. https://doi.org/10.1002/anie.201808240 Farha OK, Eryazici I, Jeong NC, Hauser BG, Wilmer CE, Sarjeant AA, Snurr RQ, Nguyen SBT, Yazaydın AÖ, Hupp JT (2012) Metal–organic framework materials with ultrahigh surface areas: is the sky the limit? J Am Chem Soc 134:15016–15021. https:// doi.org/10.1021/ja3055639
Allendorf MD, Dong R, Feng X, Kaskel S, Matoga D, Stavila V (2020) Electronic devices using open framework materials. Chem Rev 120:8581–8640. https://doi.org/10.1021/acs.chemrev.0c00033 Baig N, Sajid M, Saleh TA (2019) Recent trends in nanomaterial- modified electrodes for electroanalytical applications. TrAC Trend Anal Chem 111:47–61. https://doi.org/10.1016/j.trac.2018.11.044 Bhardwaj SK, Bhardwaj N, Kaur R, Mehta J, Sharma AL, Kim KH, Deep A (2018) An overview of different strategies to introduce conductivity in metal–organic frameworks and miscellaneous ap- plications thereof. J Mater Chem A 6:14992–15009. https://doi.org/ 10.1039/c8ta04220a
Lim D-W, Kitagawa H (2020) Proton transport in metal-organic frameworks. Chem Rev 120:8416–8467. https://doi.org/10.1021/ acs.chemrev.9b00842
Onorato JW, Luscombe CK (2019) Morphological effects on poly- meric mixed ionic/electronic conductors. Mol Syst Des Eng 4:310– 324. https://doi.org/10.1039/C8ME00093J
Papa AM, Curto V, Braendlein M et al (2016) Organic transistor arrays integrated with finger-powered microfluidics for Multianalyte saliva testing. Adv Healthc Mater 5:2295–2302. https://doi.org/10.1002/adhm.201600494
Matoga D, Oszajca M, Molenda M (2015) Ground to conduct: mechanochemical synthesis of a metal–organic framework with high proton conductivity. Chem Commun 51:7637–7640. https:// doi.org/10.1039/c5cc01789k
Stolar T, Užarević K (2020) Mechanochemistry: an efficient and versatile toolbox for synthesis, transformation, and functionalization of porous metal-organic frameworks. CrystEngComm 22:4511–4525. https://doi.org/10.1039/ D0CE00091D
Matoga D, Gil B, Nitek W, Todd AD, Bielawski CW (2014) Dynamic 2D manganese(II) isonicotinate framework with revers- ible crystal-to-amorphous transformation and selective guest ad- sorption. CrystEngComm 16:4959–4962. https://doi.org/10.1039/ C4CE00647J
Matoga D, Roztocki K, Wilke M, Emmerling F, Oszajca M, Fitta M, Bałanda M (2017) Crystalline bilayers unzipped and rezipped: solid-state reaction cycle of a metal–organic framework with triple rearrangement of intralayer bonds. CrystEngComm 19:2987–2995. https://doi.org/10.1039/C7CE00655A
Pollap A, Knihnicki P, Kuśtrowski P, Kozak J, Gołda-Cępa M, Kotarba A, Kochana J (2018) Sensitive voltammetric amoxicillin sensor based on TiO2 sol modified by CMK-3-type mesoporous carbon and gold nanoparticles. Electroanalysis 30:2386–2396. https://doi.org/10.1002/elan.201800203
Zajt T (2001) Metody woltamperometryczne i elektrochemiczna spektroskopia impedancyjna (Voltammetric methods and electro- chemical impedance spectroscopy). Wydawnictwo Gdańskie, Gdańsk
Sipa K, Brycht M, Leniart A, Urbaniak P, Nosal-Wiercińska A, Pałecz B, Skrzypek S (2018) β-Cyclodextrins incorporated multi- walled carbon nanotubes modified electrode for the voltammetric determination of the pesticide dichlorophen. Talanta 176:625–634. https://doi.org/10.1016/j.talanta.2017.07.084
Asif M, Aziz A, Ashraf G, Wang Z, Wang J, Azeem M, Chen X, Xiao F, Liu H (2018) Facet-inspired Core−Shell gold Nanoislands on metal oxide octadecahedral heterostructures: high sensing
16.
References
1.
World Health Organization (2017) Depression and other common mental disorders: Global Health estimates. WHO, Geneva
Joubert AF, Sanchez C, Larsen F (2000) Citalopram – review. Hum Psychopharmacol Clin Exp 15:439–451. https://doi.org/10.1002/ 1099-1077(200008)15:6%3C439::AID-HUP222%3E3.0.CO;2-H
Baumann P, Eap CB (2001) Enantiomeric antidepressant drugs should be considered on individual merit. Hum Psychopharmacol Clin Exp 16:85–92. https://doi.org/10.1002/hup.336
Lajeunesse A, Smyth SA, Barclay K, Sauvé S, Gagnon C (2012) Distribution of antidepressant residues in wastewater and biosolids following different treatment processes by municipal wastewater treatment plants in Canada. Water Res 46:5600–5612. https://doi. org/10.1016/j.watres.2012.07.042
Yuan S, Jiang X, Xia X, Zhang H, Zheng S (2013) Detection, occurrence and fate of 22 psychiatric pharmaceuticals in psychiatric hospital and municipal wastewater treatment plants in Beijing, China. Chemosphere 90:2520–2525. https://doi.org/10.1016/j. chemosphere.2012.10.089
Puckowski A, Mioduszewska K, Łukaszewicz P, Borecka M, Caban M, Maszkowska J, Stepnowski P (2016) Bioaccumulation and analytics of pharmaceutical residues in the environment: a re- view. J Pharm Biomed 127:232–255. https://doi.org/10.1016/j. jpba.2016.02.049
Jiménez-Holgado C, Chrimatopoulos C, Stathopoulos V, Sakkas V (2020) Investigating the utility of fabric phase sorptive extraction and hplc-uv-Vis/dad to determine antidepressant drugs in environ- mental aqueous samples. Separations 7:39. https://doi.org/10.3390/ separations7030039
Świądro M, Stelmaszczyk P, Wietecha-Posłuszny R, Dudek D (2020) Development of a new method for drug detection based on a combination of the dried blood spot method and capillary electrophoresis. J Chromatogr B 1157:122339. https://doi.org/10. 1016/j.jchromb.2020.122339
Madej M, Kochana J, Baś B (2019) Selective and highly sensitive determination of citalopram with glassy carbon electrode. J Electrochem Soc 166:H359–H369. https://doi.org/10.1149/2. 1291908jes
Ghaedi H, Afkhami A, Madrakian T, Soltani-Felehgari F (2016) Construction of novel sensitive electrochemical sensor for electro- oxidation and determination of citalopram based on zinc oxide nanoparticles and multi-walled carbon nanotubes. Mater Sci Eng C 59:847–854. https://doi.org/10.1016/j.msec.2015.10.088 Gholivand MB, Akbari A (2016) A novel voltammetric sensor for citalopram based on multiwall carbon nanotube/(poly(p- aminobenzene sulfonic acid)/β-cyclodextrin). Mater Sci Eng C 62:0480–0488. https://doi.org/10.1016/j.msec.2016.01.066
Keypour H, Saremi SG, Veisi H, Noroozi M (2016) Electrochemical determination of citalopram on new Schiff base functionalized magnetic Fe3O4 nanoparticle/MWCNTs modified glassy carbon electrode. J Electroanal Chem 780:160–168. https://doi.org/10.1016/j.jelechem.2016.08.022
Daneshvar L, Rounaghi GH, Es’haghi Z et al (2016) Fabrication a new modified electrochemical sensor based on au–Pd bimetallic nanoparticle decorated graphene for citalopram determination. Mater Sci Eng C 69:653–660. https://doi.org/10.1016/j.msec. 2016.07.025
Kaskel S (2016) The chemistry of metal-organic frameworks: syn- thesis, characterization and applications. Wiley-VCH, Weinheim
2.
17.
18.
3.
19.
4.
20.
5.
21.
6.
22.
23.
7.
24.
8.
25.
9.
26.
10.
27.
11.
12.
28.
29.
13.
30.
14.

Microchim Acta (2021) 188: 184
Page 15 of 15 184
performance toward sulfide in biotic fluids. ACS Appl Mater Interfaces 10:36675–36685. https://doi.org/10.1021/acsami. 8b12186
Asif M, Aziz A, Wang H, Wang Z, Wang W, Ajmal M, Xiao F, Chen X, Liu H (2019) Superlattice stacking by hybridizing layered double hydroxide nanosheets with layers of reduced graphene oxide for electrochemical simultaneous determination of dopamine, uric acid and ascorbic acid. Microchim Acta 186:61. https://doi.org/10. 1007/s00604-018-3158-y
Riess I (2003) Mixed ionic-electronic conductors – material prop- erties and applications. Solid State Ionics 157:1–17. https://doi.org/ 10.1016/S0167-2738(02)00182-0
Gosser DK (1993) Cyclic voltammetry: simulation and analysis of reaction mechanisms. VCH, New York
Olkowska E, Ruman M, Kowalska A, Polkowska Ż (2013)
Determination of surfactants in environmental samples part I. cat- ionic compounds. Ecol Chem Eng S 20:69–77. https://doi.org/10. 2478/eces-2013-0005
Laviron E (1979) General expression of the linear potential sweep voltammogram in the case of diffusionless electrochemical sys- tems. J Electroanal Chem Interfacial Electrochem 101:19–28 Wang J (2001) Analytical electrochemistry, 2nd edn. Wiley-VCH, New York
Masaui M, Sayo H, Tsuda Y (1968) Anodic oxidation of amines. Part I Cyclic voltammetry of aliphatic amines at a stationary glassy- carbon electrode. J Chem Soc B 9:973–976
38.
Adenier A, Chehimi MM, Gallado I et al (2004) Electrochemical oxidation of aliphatic amines and their attachment to carbon and metal surfaces. Langmuir 20:8243–8253. https://doi.org/10.1021/ la049194c
Shrivastava A, Gupta VP (2011) Methods for the determination of limit of detection and limit of quantification of the analytical methods. Chron Young Sci 2:21–25. https://doi.org/10.4103/ 2229-5186.79345
Gałuszka A, Migaszewski Z, Namieślnik J (2013) The 12 principles of green analytical chemistry and the significance mnemonic of green analytical practices. TrAC Trend Anal Chem 50:78–84. https://doi.org/10.1016/j.trac.2013.04.010
Olkowska E, Ruman M, Kowalska A, Polkowska Ż (2013) Determination of surfactants in environmental samples part III. Non-ionic compounds. Ecol Chem Eng S 20:449–461. https:// doi.org/10.2478/eces-2013-0033
Kościelniak P, Wieczorek M, Kozak J, Herman M (2011) Generalized calibration strategy in analytical chemistry. Anal Lett 44:411–430. https://doi.org/10.1080/00032719.2010.500782
31.
39.
40.
32.
33.
41.
34.
42.
35.
Publisher’s note Springer Nature remains neutral with regard to jurisdic- tional claims in published maps and institutional affiliations.
36.
37.