# Article
The Effect of Surface Modification of Ti13Zr13Nb
Alloy on Adhesion of Antibiotic and
Nanosilver-Loaded Bone Cement Coatings Dedicated
for Application as Spacers
Magda Dziaduszewska 1'*, Marcin Wekwejt1 , Michał Bartmanski 1 , Anna Pałubicka 2'3,
Grzegorz Gajowiec 1, Tomasz Seramak 1, Anna M. Osyczka 4 and Andrzej Zielinski1
1. Biomaterials Division, Department of Materials Engineering and Bonding, Gdańsk University of Technology,
   80-233 Gdansk, Poland
2. Department of Laboratory Diagnostics and Microbiology with Blood Bank, Specialist Hospital in
   Koscierzyna, 83-400 Koscierzyna, Poland
3. Department of Surgical Oncology, Medicial University of Gdansk, 80-210 Gdansk, Poland
4. Department of Biology and Cell Imaging, Institute of Zoology and Biomedical Research, Faculty of Biology,
   Jagiellonian University, 30-387 Kraków, Poland
* Correspondence: magda.dziaduszewska@pg.edu.pl
Received: 21 August 2019; Accepted: 10 September 2019; Published: 12 September 2019
Abstract: Spacers, in terms of instruments used in revision surgery for the local treatment of
postoperative infection, are usually made of metal rod covered by antibiotic-loaded bone cement.
One of the main limitations of this temporary implant is the debonding effect of metal-bone cement
interface, leading to aseptic loosening. Material selection, as well as surface treatment, should
be evaluated in order to minimize the risk of fraction and improve the implant-cement fixation
the appropriate manufacturing. In this study, Ti13Zr13Nb alloys that were prepared by Selective
Laser Melting and surface treated were coated with bone cement loaded with either gentamicin or
nanosilver, and the effects of such alloy modifications were investigated. The SLM-made specimens of
Ti13Zr13Nb were surface treated by sandblasting, etching, or grounding. For each treatment, Scanning
Electron Microscope (SEM), contact profilometer, optical tensiometer, and nano-test technique carried
out microstructure characterization and surface analysis. The three types of bone cement i.e., pure,
containing gentamicin and doped with nanosilver were applied to alloy surfaces and assessed for
cement cohesion and its adhesion to the surface by nanoscratch test and pull-off. Next, the inhibition
of bacterial growth and cytocompatibility of specimens were investigated by the Bauer-Kirby test and
MTS assay respectively. The results of each test were compared to the two control groups, consisting
of commercially available Ti13Zr13Nb and untreated SLM-made specimens. The highest adhesion
bone cement to the titanium alloy was obtained for specimens with high nanohardness and roughness.
However, no explicit relation of adhesion strength with wettability and surface energy of alloy was
observed. Sandblasting or etching were the best alloys treatments in terms of the adhesion of either
pure or modified bone cements. Antibacterial additives for bone cement affected its properties.
Gentamicin and nanosilver allowed for adequate anti-bacterial protection while maintaining the
overall biocompatibility of obtained spacers. However, they had different effects on the cement's
adhesive capacity or its own cohesion. Furthermore, the addition of silver nanoparticles improved
the nanomechanical properties of bone cements. Surface treatment and method of fabrication of
titanium affected surface parameters that had a significant impact on cement-titanium fixation.
Keywords:	spacer;	coating adhesion; Ti13Zr13Nb alloy; selective laser melting; bone
cements modification
Materials 2019,12,2964; doi:10.3390/ma12182964
www.mdpi.com/j ournal/materials

Materials 2019,12,2964
2 of 21
### 1. Introduction
Spacers are instruments used in revision surgery for local control of postoperative infection.
Usually, this is a temporary implant made of antibiotic-loaded bone cement or it consists of a metal rod
covered with bone cement containing an antibiotic [1]. These procedures have recently become the
gold standard for the treatment of bone infection with a high, near 90% effectiveness [2,3]. The need of
using spacers results from the lack of systematic antibiotherapy efficacy and the necessity to ensure
mobility of the patient [4]. Spacers can bought commercially bought or manually prepared by applying
bone cement containing antibiotic on a metal rod, sometimes using a special mold. The following main
types of spacers can be distinguished: revision hip or knee endoprosthesis spacers or intramedullary
nail spacers [1-5]. Literature reports that the overall complication rate of using spacers is about 26%.
The most common complications after spacer implantation include the dislocation of a spacer, fracture
of implant or surrounding bones, erosion in bones, and systemic antibiotic toxicity. Their main risk
factors are: loss of bone support as a results of improper spacer design, incorrectly prepared cements
and overall patient's state of health and functioning [5,6]. Moreover, one of the primary problems
with using spacers is related to the debonding effect of the metal-bone cement interface, where
micro-sliding movements between the components lead to aseptic loosening [7,8]. The prevention
of aseptic loosening requires the appropriate manufacturing, material selection, as well as surface
treatment of spacers.
In numerous studies [9-11] it has been suggested that implant-cement fixation properties might
be improved by surface roughness augmentation of the metal rod, which exposes sharp-edges texture,
enlarges effective surface, and chemically activates it. The ideal metal and cement joining interface
should rely on the association of micromechanical and chemical bonding [12]. A wide variety of rod
surface modification techniques have been developed regarding the adhesion strength, including:
blasting, acid-etching, grinding, oxidizing, plasma-spraying, as well as coating application [13,14].
Mostly, blasting, acid-etching, and grinding are used, due to their convenient and widespread
commercial implication [15-17]. Moreover, it is assumed that the bactericidal modifications of bone
cement can also affect the adhesion of the cement. Antibiotics, such as gentamicin or ciprofloxacin, are
widely used, however they affect the mechanical properties of bone cement [18]. Hence, nanosilver
recently is an intensively tested alternative. Its main advantage is high bactericidal effectiveness,
resistance to polymerization temperature, as well as no adverse effect on the mechanical properties of
the cement.
For the longevity of spacers, the appropriate material selection is another important issue. The use
of a metal rod can significantly increase the strength and stability of the spacer [1,4]. So far, the
primary metal used for that kind of medical solution was corrosion-resistant steel. However, due to
low biocompatibility and high density negatively affecting its weight, titanium alloys have commonly
substituted corrosion-resistant steel. Among others, titanium with zirconium and niobium deserves
special attention, due to their excellent biocompatibility, corrosion resistance, high strength/weight
ratio, good fatigue resistance, and lower Young's modulus as compared to most of the metals, thus
better matching the values of those existing in human bones [17,19].
Last but not least, the ability to create personalized spacer might result in their longevity. Hence,
selective laser melting (SLM), which shows several advantages over conventional manufacturing
techniques, including highly accurate and predictable structure, customized mechanical properties,
personal-customized architectures, might be an interesting option [20]. Giving the importance of
surface topography, the SLM process seems to have the prominent ability to obtain satisfied parameters
of roughness. However, it may also produce unmelted particles after processing [21]. Those particles
might have a negative effect on mechanical properties of the rod, or eventually they may loosen, which
can have a significant impact on metal-bone cement interface. For this reason, post-processing, such
as sandblasting or chemical etching, is usually implemented [22]. Although some authors provide
the evaluation of adhesion strength of two dental adhesives in the laboratory under shear or tensile
modes [23], the problem of adhesion strength between the metal/bone cement still seems to be poorly

Materials 2019,12,2964
3 of 21
explored. Other studies evaluate shear adhesive strength [16], or tensile strength of different materials,
mostly metal/ceramic [24]. To our knowledge, there are no reports on the adhesive joining between
SLM-made specimens and cement.
The main aim of this research was to investigate the effect of surface modifications of Ti13Zr13Nb
alloy on the adhesion of pure and modified bone cement. We evaluated the titanium surface after
different surface treatment, assessed cement adhesion to this surface, and checked the biological
properties of modified bone cements. A commercially available titanium solid rod and SLM-produced
were used as controls. Subsequently, the printed specimens were subjected to the following surface
treatments: sandblasting, grinding, and etching. Three types of bone cement were applied to
the prepared titanium substrates, i.e., pure, containing gentamicin and doped with nanosilver.
Designing the proper shape of the titanium rod obtained by SLM method, application of surface
treatment, and selection of an effective bactericidal bone cement coating were the key approaches for
eliminating potential complications during implantation as well as ensuring the success of two-stage
revision treatment.
### 2. Materials and Methods
  1. Specimens Preparation
    1. Titanium Specimens Fabrication
The cylindrical specimens (discs) were manufactured by SLM machine while using Ti13Zr13Nb
spherical powder (TLS Technik GmbH & Co. Spezialpulver KG, Bitterfeld-Wolfen, Germany) with
particle size ranging from 20 to 70 pm. Two sets of alloy samples were designed while using Materialise
Magics (Materialise NV, Ghent, Belgium) software in the form of specimens suitable for surface
evaluation (d = 20 mm, h = 3 mm) and biological tests (d = 10 mm, h = 3 mm). The cylinders were
oriented in the following combination: the smaller ones in parallel position, while the bigger ones
perpendicularly to the building platform. The alloy specimens were manufactured by the Realizer
SLM 100 (Realizer GmbH, Borchen, Germany) machine that was equipped with the ytterbium one
mode fiber laser CW YLR-100-SM (IPG Laser GmbH, Burbach, Germany) using 1070 nm wavelength.
The process parameters (i.e., spot diameter, laser power and scanning speed, and layer thickness)
are patent pending. Additionally, the laser melting process was carried under a protective argon
atmosphere (grade 4.8). As a reference material, a hot-rolled bar of Ti13Zr13Nb alloy provided by a
commercial supplier was used (Xi'an SAITE Metal Materials Development Co., Ltd., Xi'an, China).
The rod specimens (d = 20 mm, h = 3 mm) were divided into five groups according to the surface
treatment that was applied (Figure 1). The first group, which was made of the commercial solid
bar (Solid bar), was used as a control A, while the second group were surface untreated SLM-made
specimens (Untreated SLM), designated as a control B. Group 3 consisted of the SLM-made specimens
that were treated by sandblasting (Sandblasted SLM) with aluminum oxide particles (50 pm) for 20 s
and a distance of 10 mm from the sandblaster nozzle tip (d = 2 mm) with the working pressure 6 bars.
Group 4 consisted of discs acid-etched for 15 min. in solution of nitric acid (65%) and hydrofluoric acid
(40%) at ratio 3:1 (Etched SLM). The discs were then rinsed in distilled water for 10 min. in an ultrasonic
bath and left to dry for 24 h. Group 5 consisted of ground SLM-made cylinders (Ground SLM). Manual
grinding was based on the wet process with the use of the 220 pm (P80) grade silicon carbide paper
that was mounted on a lapping machine Saphir 330 (ATM GmbH, Mammelzen, Germany) at a feeding
rate of 400 rpm. Each sample was ground until the SLM porous structure was removed. Specimens of
each group were finally cleaned in 95% ethanol while using an ultrasonic cleaner for 10 min. and left
to dry in room temperature for 24 h.

Materials 2019,12,2964
4 of 21
Figure 1. Images of the alloy specimens a fter different surface treatments (a) Solid bar; (b) Untreated
selective la ser melting (SLM); (c) Sandblasted SLM; (d) Etched SLM; and, ( e) Ground SLM.
CementCoatingP reparation
The commercially available PMMA bone cement (Tecres, Verona, ItaCy) was used to create
the coating on the titanium alloy specimens. Cement was prepared following the manufacturer's
instructions and according to tine international standard ISO 5833:2002 [25]. The procedure consisted
of manually combining the liquid component with the powdea in a bowl and then mixing to a paste
with a spetula for about onc minute. A fter mixing, the paste was applie d on titanium alloy specimens
that wece placed in speccally prepared silicone; moulds. Afterwards, it was allowed to cuee for 1 Itl in
ambient conditions. Three different cements were investigated: (1) unmodified bone cement (BC), (2)
antibiotic-loaded bone cement (gentamicin sulfate, Sigma Aldrich, Steinheim, Germany) (A-BC), and
(3) bone cement modified with nanosilver (MkNano, Mississauga, ON, Canada—average particles size
50 nrtt purity 99.9%) (NpAg-BC). For the modified bone cement, the prooedure included the addition
of modifiers to the cement powder and then mixing tor one min. The proposed modifications and
their concentradions were selected based on the previous studies [26-28]. Table 1 presents the final
compositions of the tested bone cements.
Table 1. Chemical composition of bone cements used for titanium specimens coating.
Polymethyl metha crylate	Unmodified Bone Cement/BC/ / 84.30	Antibiotic-Loaded Bone Cement/A-BC/ / Powder component [% w/w / 83.05	Bone Cement Modified with / Nanosilver/NpAg-BC/ / ] / 83.05
Barium sulphate	13.00	12.80	12.80
Benzoyl peroxide	2.70	2.65	2.65
Gentamicin sulphate	X	1.5	X
Nanosilver	X	X	1.5
Liquid component [% w/w]
Methyl Methacrylate	99.10	99.10	99.10
N, N-dimethyl-p-toluidine	0.90	0.90	0.90
Hydroquinon e	75	n5	75
Methyl Methacrylate	99.10	99.10	99.10

## 2.2. Surface Evaluation of Titanium Specimens
    1. Surface Topography
The surface topography of titanium alloy specimens after surface treatments was observed using
a high-resolution Scanning Electron Microscope SEM (JEOL JSM-7800F, Tokyo, Japan). Moreover,
the X-ray Energy Dispersive Spectrometer (EDS) investigated its chemical composition (Edax Inc.,
Mahwah, NJ, U SA)

Materials 2019,12,2964
5 of 21
    1. Surface Roughness
The surface roughness of the titanium alloy specimens was assessed by using; a contact profilometer
with EVOVIS software (Hommel Etamic Waveline, Jenoptik, Jena, Germany). The test was conducted
using cut-off wavelength Ac =2.5 pm, eveluatioo length ln = 12.5 mm and scanning speed of 0.50 mm/s,
according to ISO 4287-1997 [2g]. Ths average v alues of roughness (Ra), the peak-to-valley roughness
(Rz) and maximum peak-to-mean height (Rp) were obtained from this test.
    1. °2r(oce Wettability oed °2r(oce Energy
The surfoce wettabdity and surface energy; (Ys) were determined by the contact angle
measurements. The studies were carried out by fhe falling drop method with two liquids: distilled
water and diiodomethane (Sigma Aldrich, Steinheim, Germany) while using an optical Censiomefer
(Attention Theta Life, Biolin Scie ntific, Espoo , Finlend). Tire volu me of the li quids was about 1 pL/sample.
Each measurement was carried out five times, immediately after deposition of the drop of the liqrnd.
dhp analysis was c arried out while using tho OneA/tensmn program ( Biilin Scientific Espoo, Fmland)
pnd mathematical calculations that are lbased on the OWRK method [30-32].
    1. Surface Nanomechanical Properties
Nanomechanical properties of titanium alloy specimens were determined while using
nanoindentation technique performed with the NanoTest™ Vantage (Micro Materials, Wrexham,
UK) equipment. Three-sided diamond, pyramidal Berkovich indenter was used. Twenty-five (5 X 5)
inde pendent nan oiidentation measurements on each of tested surface were performed. The mvxi mum
fotce od indentation was 50 mN, the loading and unloading time were set up» to 20 s and holding time
with muximum force was 5 s. The distance be tween the indentations was 20 pm for both axes (x and y).
During a single me asurement, the load-dis placement oufve was determ ined based on the Oliver and
Pharr method [33]. The surface haedness (IT) and reduced Young's modulus (Er) were calculated while
using; the integrated software.
  1. Assessment of Cement Cohesion and its Adhesion to the Surface of Titanium Specimens
    1. Cohesion of Cement
The cohesion of cements was determined while using nanoscratch-test performed with the
NanoTest™ Vantage (Micro Materials, Wrexham, UK) equipment. The Berkovich indenter was usid
for nanomechanical properties investigation. Ten measurements of tested bone cement coating were
performed. The applied lead range of a sinfle test was 2rom zero to 200 mN with force rate 1.3 mN/s
at a distance of 500 pm. During; the measurements, the appleed torce in functihnol of friction force
ou rve was determined. IB ased on the ebrupt change in frictional fotce, the cohesion o f tested cements
was d etermined.
    1. Adhesion of Cements to the Titanium Surface
The adhesive strength ol bone cement coutings to the surface-modified titanium alloy specimens
was tested according to the p ull- off teet pxocebure thaf was p roposed by Ozturk & Tannant [34]. The
purpose of the test was to determine the interface bond strength i.e., between cement coating and
titanium specimens. For this reason, the greatest tensile stress that the surface area can bear before
being detached was investigatad. dhe bone cement coating with a thickness or 2 mm was applied
on the titanium specimens in disc form (d = 120 mm, h = 3 mm). Next, the; specimens weee glue(2
with e two componebf epoxy resin Metacryl (Technicqll, Trzebinia, Polandl to the flaf threaded screws
(M12) set in a linear manner. The screws and specimen had both been ground and cleaned while using
acetone. After the epoxy was cured (24 h), the specimens were installed in the universal testing machine
INSTRON 1195 (INSTRON, Norwood, MA, USA). The special holders contained cardan joint and a set

Materials 2019,12,2964
6 of 21
of hooks was used in order to minimize the bending moments (Figure 2). The measurements were
carried out with lo ad at a constant rate of 1 mm/min. The maximum force was recorded upon failure
and assumed as the adhesive strength, while the failed surface was visually examined and documented.
Figure 2. The specimen holder mounted in universal tensile machine.
  1. Assessment of Biological Properties of Bone Cement Coatings
    1. Inhibition of Bacterial Growth
The inhibition of bacterial growth by bone cement coatings was evaluated by test of
Bauer-Kirby et al. [35]. Two bacterial reference strains were used: S. aureus (ATCC 2921., Microbiologics
Inc., St. Cloud, MN, USA) and E. coH (ATCC 25922, Microbiologics Inc., St. Cloud, MN, USA).
Each bacteria strain was incubated in a sterile 0.9% NaCl solution for 10 min. Subsequently, 100 pL
bacteria suspensions containing 1.5 X 108 CFU mL-1 were taken and seeded on Mueller-Hinton
Agar plates (BecSon Dickinson GmbH, Heidelberg, Germany). The inhibioion zone tests involved the
incubation at 36 °C of the titanium speuimens thot were coveered with bone cement coatings with the
resulting; bacterial suspension. The experiment wns performod while using three specimens (n = 3)
for each type of bone cement coating applied un tiUanium specimens (Untreated SLM) in disk form
(d = 10 mm, h = 3 mm). Before testing, tire bone cement spocimeno were sterilized in 70% EtOH,
followed by 30 min. exjeosure to UV light. The measuremunts of the inhibition zone: were carried out
after: 24, 72, and 168 ho. The bacterial growOh inhibition zone was determined as a diameter withuut
bacteeial growth in mitt.
    1. Cytocompatibility
All of the reagents that were used in this study were purchased from Thermofisher (Thermofisher
Scientific, Waltham, MA, UUA), unless stated otherwise. Ths MTS assay was used in order to aseess the
cytocompatibility of tested bone cement coatings. Fot the assay, the mouse osteoblast precursor cell line
MC3T3-E1 (ATCC CRL-2593, Manassas, VA, USA) was uted. The cells were expanded in culture while
using alpha-MEM supplemented with 15% fetal bovine serum (FBS), 0.1 mM l-ascorbic acid phosphate
(Sigma Aldrich, Steinheim, Germany), and 1% antibiotics (penicilin-streptomycin—10,000 U/mL,

Materials 2019,12,2964
7 of 21
Thermo Fisher Scientific, Waltham, MA, USA). When the cells reached a confluent monolayer, they
were lifted from culture flasks while using 0.25% Trypsin-EDTA, counted and seeded directly on the
surface of the cements, as described below. Briefly, three specimens for each type of bone cement
(n = 3) in disk form (d = 20 mm, h = 2 mm) were used. Before testing, the bone cement specimens
were sterilized in 70% EtOH followed by 30 min. exposure to UV light. Afterwards, the specimens
were placed in separate wells of 24-well culture plates and were covered by 2 X 104 MC3T3-E1 cells
suspended in 1 mL of standard culture medium. The cells were incubated at 37 °C in air containing
5% CO2. The medium was aspirated and replaced with a fresh one every 48 h. Cell viability was
evaluated while using CellTiter96 Aqueous One Solution Cell Proliferation Assay (MTS, Promega,
Kraków, Poland). After three and seven days of culture, culture media were withdrawn and each
material sample was covered with 400 pL of MTS solution diluted 10X in phenol-red free alpha-MEM.
The plates were then incubated at 37 °C in a culture incubator until the development of a brownish
color of the MTS solution. The intensity of the developed color is proportional to the number of
actively metabolizing live cells. Approximately after 30 min., 200 pL aliquots of the MTS solutions
from individual wells were transferred to separate wells in 96-well plates and their absorbance at
492 nm was colorimetrically measured while using a plate reader. The results were expressed as %
change in MTS absorbance values for modified bone cements compared to 100% value assumed for
pure bone cement.
  1. Statistical Method
Statistical analysis of the data was performed while using commercial software (SigmaPlot 14.0,
Systat Software, San Jose, CA, USA). The Shapiro-Wilk test was used to assess the normal distribution of
the data. All of the results were presented as mean ± standard deviation (SD) and they were statistically
analyzed while using one-way analysis of variance (one-way ANOVA). Multiple comparisons versus
the control group between means were performed while using the Bonferroni t-test with the statistical
significance set at p < 0.05. Two controls of Ti13Zr13Nb were adopted in the statistical analysis: (A)
specimens from commercially purchased bar (Solid bar) and (B) specimens that were produced by
SLM (Untreated SLM).
### 3. Results and Discussion
  1. Surface Evaluation of Titanium Alloy Specimens
    1. Surface Topography Evaluation
In the control group, a commercial bar appeared to have the smoothest surface compared with
other groups. Small pits and shallow grooves generated in the grinding procedures could be observed
(Figure 3e). In opposite, SLM-made specimens exhibited the surface covered by numerous unmelted
spherical particles of initial powder of various sizes. Most of the bigger grains consisted of the solid
linchpin on the surface, while the smaller were grouped or loosely covered the specimen's surface or
other non-melted particles (Figure 3a). In the case of sandblasting treatment, the stream of accelerated
particles removed most of the spherical particles and significantly expanded the surface with many
irregular micro-cracks (Figure 3b). However, it could be noticed that some bigger particles were still
mounted to the ground. Moreover, the EDS analysis of Sandblasted SLM specimens revealed the
presence of alumina, likely from the abrasives after the sandblasting process (Figure 4). The etching
treatments resulted in the formation of a more uniform fine-textured titanium alloy surface with most
of the spherical grains being dissolved. This surface was relatively smoother and irregularities with
well define depression areas were more evenly distributed (Figure 3c). The ground SLM specimens
had relatively flat surface with clearly-visible shallow grooves generated by grinding and characteristic
for small grade carborundum paper. However, the morphology revealed protruding elevations and
depressions that were composed of several layers characteristic for SLM-made specimens (Figure 3d).

Materials 2019,12, 2964
8 of 21
As often reported by others, the SLM process causes irregularity due to internal porosity. For example,
Fousova M. et al. [36] observed the 0.78 ± 0.11% internal porosity of fully dense samples. Such porosity
of studied samples might be a result of processing; technology, and/or different sizes of titanium
alloy powder.
Figure 3. SEM images of the titanium alloy specimen surfaces after different surface treatments and at
two different mabnifications (X500 and X2000): (a) Untreated SLM; (lb) Sandblasted SLM; (c) Etched
SLM; (dL Ground SLM; and, (e) Solid bar (the presented pictures are representative Sor five specimens).

Materials 2019,12, 2964
9 of 21
Figure 4. Results of Energy Dispersive Spectrometer (EDS) analysis after different surface treatments:
(a) Solid bar, Untreated SLM, Etched SLM, Ground SLM; and, (b) Sandblasted SLM (the presented
pictures are rep re sentative for three an alyses of each surfa ce treatment specimens).
3.1.2. Surface Roughness Evaluation
A comparison of the mean roughness parameters: the average roughness (Ra), the peak-to-valley
roughness (Rz) and maximum peak-to-mean height (Rp) values among the five experimental groups
of titanium specimens that were obtained by surface profilometry testing are given in Table 2.
Table 2. Surface roughness parameters: Ra, Rz, Rp of surface-modified titanium specimens measured
by profilometer (n = 5; data are the means ± SD; * significantly different from control A (p < 0.05);
# significantly different from control B (p < 0.05)).
Surface Roughness Parameters (pm)
	Solid Bar/Control A/	Untreated SLMlControl Bl	Sandblasted / SLM	Etched SLM	Ground SLM
Ra	0.14 ± 0.02#	10.13 ± 0.20*	10.79 ± 0.36 *	8.82 ± 0.99 *'#	1.38 ± 0.24 *'#
Rz	1.71 ±0.28#	55.95 ± 3.49 *	54.64 ± 0.79 *	45.5)2 ± 5 . 09 *'#	7.32 ± 6.90 #
Rp	1.30 ±0.20 #	28.78 ± 1.89*	30.52 ± 0.26 *	26.00 ± 3.09 *'#	3.68 ± 0.57#

The lowest me? an surface roughness value was obtained for control Solid bar (0.14 pm]) and Ground
SLM (1.38 pm), and it was significantly lower than for till either specimens tested. Statistically significant
(p < 0.05) higher roughness was detected for all types of SLM-made specimens: without surface
treatment (10.13 pm), sandblasted (10.79 pm), etched (8.82 pm), and ground (1.38 .m). Sandblasted and
etched modifications had significantly inrreased the Ra parameter when compared to control specimenr
(Solid bar and Untreated SLM). However, the? highest value of Ra was observed for sandblasting, while
etching decreased the mean roughness parameters as compared to non-modified SLM specimens.
Furthermore, the peak-to-valley roughness and maximum peak-to-mean height were in agreement
with Ra values, mrrely the Rz and Rp of Ground SLM are not statistically significant when compared
to Solid bar.
There?; obtaineL results are consistent with SEM imaoes (Figure 3), in which better surface
development was observed due to various modifications. In the case of unmodifird SLM, roughness
was mainly created by unm elted powder particles. The obtained iesultr are io line? with the literature,
where the valoes of Ra for SLM-made specimens are between 5 pm < Ra < 40 pm [37-39] and they
mainly depend on the SLM processing and the properties of the powder [38]. Increasing values of
roughness parameters for the specimens that were subiecteh to sane-blasting are osoociated 'with the
existence of the additional craters, ridges, and sharp) edoes of alumina particles [40]. Those plastic
defrrmations accounted, among others, for the highese values of Rp (30.52 pm) [41]. Similarly,
Aparicio C. et al. [42] reported that this sandblasted specimens were found to possesr a rougher ruriace

Materials 2019,1 2,2964
10 of 21
(4.74 pm) when compared to 6hose after etching (1.69 pm). This might be explained by the action
of tho ac4d t;loci^ removes most of the blasted material. However, Ro values were lower lhan rhose
foisnd by us, lfc> iL.i lo this tould Ite attributed to the various particles tize, 4he shepe or hardnpss of ohe
blasting maOerlal [40]. Ovorall, it is assumed that the surface roughnass Increases w4th increasing
ceramio hartlctee size [43]. The turface roughneso can be dividhd into shree lovels: macra-roughness
(Re ~ 10 pm), micro-roughness (Ra ~ 1 pmf and oieasc-i'ojughness (Ita < 200 nm) [44,45]. Thus, it may be
concladed that sandblystmg and etchmg results in macro-roughness of specimens, whereas grinding
resu Its in specimen mlc ro-roughnes s.
cl.1.3^ Surface Wettobitity end Surface Energy Evaluation
Table o illustrates the suioace wettability results o1 imv.oii^t^jrifti^cf modifications of titanium
alloy specimens.
Table 3. Surface wettability of the investigated surface-modified titanium alloy specimens determined
by the measurements of the contact angle of distilled water (n = 5; data are the means ± SD; * significantly
different from control A (p< 0.05); # significantly different from nontrol B (p < 0.05)).
	Surface Wettability-the Vafue of Contact Angle (°)
Solid Bar/Control A/	Unareated SLM/Control B/	Sandblasted SLM	Etched SLM	Ground SLM
76.4 ± 7.1 #	P20.3 ± 23 *	68.3 ± 5.U3 #	8:1.2 =t 6.6 #	50.9± 4.8 *'#

The tltemum alio. specimens thai were produced by SLM were characterized by hydrophobic
surhrce (120.3° > 90°f. Applied surface modifications improved the surface wettability and resafted in
the hydrophilic properties. The commerccal titanium rod alro had e hydrophilic surlace (74.4f. Overall,
io is tisumed that materials hotter adhere to hydrophilic surfaces [46]. Moreover, lower lurface energies
of adhesives should allow for their behter edhetirn and good bonding [40]. Gfvcnz a consteut surtece
energr of cementdt their adhesion wild depend oro the surface anergy of titanium elloy spedmdos.
Table 4 showe the rurface energy of mvestigoted specimsnr.
Table 4. The rurface energy surface-modified titanium adoy specimens. The results are dioplayed as
mean valoes of coot tart anzies; * significantly different from control A (p < 0.05); S significently different
from oonfrol B ]p < 0.05)o
Surfece Energy (mN/m)
	Staled BartControl A/	Untreated SLM(Control B/	Sandblasted / SLM	Etched SLM	Grtund SLM
Ys	-41.11	48.69	5 1.39#,*	385.40 #,*	6 a. 2 5 #,*
YsD	33.84	45.07	47.27	34.71	856.94
Ysp	t.27	3.62	4.12	3.69	30.31

Based on the surface energy analysis, Etched SLM should be the best for adhesion and the Ground
SLM should be dhe worft For the most specimens, the adhe sion wilt be matnly ba wed on dispers [ve
interactiono. The only exception is Ground SLM, where disparsive and polar mteractions occurred
reasonably equally.
e.1.4. Su4(ace Nanomechanical Properties Evaluation
Table 5 shows the results of nanomechanical tests (nanohardness, Reduced Young's modulus
H/E and H3/E2 ratio). The nanohardness and Reduced Young's modulus were shown in Figure 5
as three-dimensional (3D) maps. The highest nanomechanical values were observed for control
specimens (6.77 GPa and 5.15 GPa for Untreated SLM and Solid bar, respectively). This may be

Materials 2019,12,2964
11 of 21
due to stresses in the surface layer of these specimens that occur in the material after the grinding
process. The appearance of the stress in a layer of material after grinding was proven by Tao et al. [48].
The smallest values of nanohardness and Reduced Young's modulus were obtained for Etched SLM
(0.41 GPa). The etching involves concentrated acids that change the surface topography and deteriorate
the mechanical properties of specimens. Along with the increase in surface roughness values, the
increasing standard deviation values are observed, which confirms the reliability of the tests that
were carried out [49]. Rough surfaces are characterized by a non-uniform distribution of mechanical
properties. Differences in roughness values and their impact on the standard deviation values of
nanomechanical measurements correlate well with the graphic presentation of the results in Figure 5.
The control surface (Solid bar) with the lowest roughness parameter Ra (0.14 pm) is characterized by
the smallest relative standard deviation (~20% of value) and the most homogeneous distribution of
results that are presented in Figure 5e. According to nanoindentation test, Etched SLM is the most
suitable for implanting to the human body, because of the value of Young's modulus (24 GPa) closest to
that of a human bone (10-30 GPa) [50,51]. It is highly recommended that Young's modulus of implants
should be similar to the human bone [52]. Low mechanical properties of a surface layer might have
a negative impact on the implantation process, because such a layer might be damaged during the
implantation procedure. The nanoindentation method is also used to determine the resistance to failure
of tested surface, which is defined by H/Er ratio [53]. A higher ratio reflects better resistance to elastic
stress [54]. In addition, the H3/Er2 ratio indicates resistance to plastic deformation [55]. The results in
the Table 5 clearly show that the Untreated SLM and Ground SLM specimens display the best wear
resistance (H/E 0.066 and 0.060, H3/E2 0.027 GPa and 0.020 GPa, respectively). The smallest value of
the H/Er and H3/Er2 ratio were obtained for Etched SLM (0.018 and 0.0002 GPa, respectively), which
was also characterized by the lowest values of the nanomechanical properties.
Table 5. Nanomechanical properties (nanohardness and reduced Young's modulus) of investigated
surface-modified titanium alloy specimens (n = 25; * significantly different from control A (p < 0.05);
# significantly different from control B (p < 0.05)).
	Solid Bar/Control A/	Untreated SLM/Control B/	Sandblasted / SLM	Etched SLM	Ground SLM
Nanohardness / (GPa)	5.15 d 1.12#	Э.94 d 1.76 *	2.4Э d 1.46 *'#	0.41 d 0.19 *'#	6.77 d 2.14 *'#
Reduced
Young's	1ЭЭ.50 d 17.Э6 #	62.9Э d 26.Э0 *	5Э.15 d 24.15 *	24.00 d 9.92 *'#	110.Э9 d 20.26 *'#
modulus(GPa)
H/Er	0.0Э8 d 0.004 #	0.066 d 0.025 *	0.052 d 0.0Э8	0.018 d 0.006 *'#	0.060 d 0.009 *
H3/Er2 (GPa)	0.008 d 0.004 #	0.020 d 0.017 *	0.016 d 0.00Э	0.0002 d 0.0015 *'#	0.027 d 0.015 *

Materials 2019,12,2964
12 of 21
Figure 5. 3D maps of nanomechanical (nanohardness—left and Reduced Young's modulus—right)
properties of studied specimens after different surface treatments: (a) Untreated SLM; (b) Sandbla sted
SLM; (c) Etched SLIM; (d) Ground SLM; and, (e) Solid bar.
Table; 6 presents the cohesion of bone cement coatings. The antibiotics did cot significantly increase
cohesion of the coatings (from 82.1 GPa to 84.9 GPa) opposite to the addition of nanosilver (from
81.2 mN for BC cement to 122 mN for NpAg-BC). Nanosilver doping into cements (NpAg-BC) had a
significanC and positive effect on their nan omec hanical properties. This may be d ue to strengthen mg

Materials 2019,12,2964
13 of 21
effect of metallic nanoparticles or the creation of a composite cement matrix. The improvement
of bone cement mechanical properties by silver nanoparticles was previously reported by us [27].
Russo et al. also obtained similar effects [56] for gold nanoparticles. However, the presence of metallic
nanoparticles in bone cements could deteriorate the mechanical properties of such composites due to
stress concentration at the matrix/filler interface [57,58]. Besides, silver nanoparticles are probably able
to form permanent bonds with matrix that improve the resistance of cements.
Table 6. Cohesion of bone cement coating (n = 10; * significantly different from BC—p < 0.05;
# significantly different between group—p < 0.05).
BC (mN)	A-BC (mN)	NpAg-BC (mN)
81.2 ± 19.4	84.9 ± 13.8#	122.0 ± 11.5 *,#

  1. Creation of Bone Cement Coating on Titanium Specimens
Pure and modified cements were both successfully applied to all tested titanium alloy specimens.
We noticed that the type of titanium alloy affects the cement adhesion. The most problematic in terms
of their application were solid bar specimens (control A). Furthermore, there was no difference in the
process of application of the coating for pure cement and doped with modifications.
  1. Assessment of Cement Adhesion to the Surface of Titanium Specimens
The results of our study indicate the phenomena that the increase of the surface roughness
increases the effective area that is available for the better adhesion of cement. The SLM-made
specimens: untreated, sandblasted and etched demonstrated significantly higher adhesion of either
BC, A-BC, or NpAg-BC cements than solid titanium (control A). Similar results were obtained for the
ground group covered by BC and NpAg-BC. Only SLM ground specimen with pure bone cement (BC)
showed no significant difference when compared to the control A group (p < 0.05). All of the studied
cement types displayed significantly lower adhesion to solid and SLM ground specimens as compared
to untreated SLM specimens (control B), while sandblasting and etching of titanium alloys significantly
increased the adhesion of BC (p < 0.05) and the modifications of BC by antibiotic or nanosilver did
not affect this process. The addition of gentamicin and silver nanoparticles to BC also resulted in
their better adhesion to untreated SLM, but they adhered poorer to the ground titanium alloys when
compared to the non-modified BC. Furthermore, cement with antibiotic increased the pull-off-strength
adhesion for solid titanium specimens, while the addition of nanosilver decreased. It would be ideal
if the failure only appeared at the cement-titanium interface, but a very strong adhesion between
BC and sandblasted and etched titanium alloys displayed higher than glued fixation. Overall, the
debonding force for studied specimens was high enough to treat the titanium/cement adhesion as
sufficient (Table 7).
Table 7. Adhesion of bone cements to surface-modified titanium alloy specimens (n = 5; data are the
means ± SD; * significantly different from control A (p < 0.05); # significantly different from control B
(p < 0.05)).
Adhesion of Bone Cements to Surface-Modified Titanium (N)
	Solid Bar/Control A/	Untreated SLM/Control B/	Sandblasted / SLM	Etched SLM	Ground SLM
BC	86.8 ± 53.1 #	1432.7 ± 79.5 *	>2000 x'#'*	>2000 x,#,*	827.9 ± 276.6 *'#
A-BC	168.3 ± 92.6 #	1929.1 ± 18.8 *	>2000 x'*	>2000 x,*	61.6 ± 46.7#
NpAg-BC	10.0 ± 0.5 y,#	1980.3 ± 20.3 *	>2000 x,*	>2000 x,*	69.1 ± 25.6 *,#

x The rupture of adhesive-specimens connection with an average strength about 2000 ± 234.5 N. y The rupture of
cement-specimens connection when the preload was applied.

Materials 2019,12,2964
14 of 21
The high standard deviationof the results can be explained by the methodology of the applied
epoxy test, which is highly randomized and it depends on surface roughening, cleaning, drying,
fixation time and glue mixing; speed, epoxy-ingredients proportions, time of glue drying, position of
drying;, etc.
The highest adhesion of BCs to untreated SLM, as well as sandblasted and etched specimens, is
associated with their roughness. Topographical find chemical properties of the coatings both might
have contributed to the pull-off-strength results. Appropriate roughness enlarges the actual interface
betwehn the metal surface and cement, and thus increases the adhesive strength [5h,60]. The alumina
particles remaining after sandblasting provide mechamoal interlocking, which improves the bonding
strength with BC [61]. The oxide layer that formed after the etching of titanium alloys aleo improved
the adhesion to BC. The slightly lower value of BC adhesion for untreated SLM might result frtm
unmated titanium patticles loosely covering the surface and thus decreasing tire titanium/cement
bonding. Weak adhesion between the bone cements and ground titanium alloy; correlated wi th the
lack of mechanical interlocking and their low roughness. Some authors [62-64] showed the correlation
between coating adhesion to SLM-made specimens as compared to soiid mateoial. Different types of
fracture images (adheuive, cohesive, and mixed) were obsetved in the samples after the deeachment of
the coating by puil-off-test (Figure 61
Figure 6. Images of sefected specimens after detachment of bune cement coatings by pull-off-test
(a) Solid bar; (b) Ground SLM; (c) Untreated SLM; (d) Etched SLM; andi (e) Sandblasted SLM- (the
presented pictures are representative for five experimente).
The solid and ground specimens showed some dotachment of bone cement. The cement
delamination occurred, followed by the failure of the coating adhesion tc tha substrate. This suggested
uneven adhesion of the oements to the metal. Cememt islande of various size, numbers, and shape
characterized the modified titanium alloy surfaces. After the pull-off test, the titanium alloy surfaces
were enricher mostly with BC tslands, less A-BC, and ihe smallest amounis NpAg-UC was found on
titanium alloy surfaces. This confirms the results of the putl-off strength test, where the adhesive
strength is lower. There were also morphological differences in fracture mode between untreated,
sandblasted and etohed SLM specimens as compared to -he ground ond solid ones1 The untreated,
sandblasted, and etchud SLM disks were regularly covered by cgment, not in the form of cement island,
and thfs could Ire correlated with higher roughness and better adhesion strength of these titanium
alloy specimens (Figure 6).

Materials 2019,12,2964
15 of 21
# 3.4. Assessment of Biological Properties of B one Cement Coating
I nhibition o f Bacterial Growth
The occurrence of the inhibition zones indicated that the modified cements effectively inhibited
bacterial growth (Figure 7).
Figure 7. Bacterial growth inhibition on untreated SLM titanium alloy specimens coated with
bone cements: unmodified ( BC), antibiotic-loaded iA-BC) and modified with nanosilver (NpAg-ESC)
after 7 days of bacteria culture (the presented pictures a re representative tor three experiments):
(a) Staphylococsus aureus and (b) Escherichia coli.
Pure cement had no antibacterial properties. The addition of gentamicin or silver nanoparticles to
BC inhibited local bacterial growth. (Figure 7, Table 8).
Table 8. Bacterial growth inhibition for untreated SLM titanium alloy specimens coated with bone
cements (n = 3; data are the means ± SD; * significantly different from control (p < 0.050; # significantly
different between modifications (p < 0.05)).
Tha Bacterial Growth Inhibition Zone (mm)
Time (h)	BC	A-BC	NpAg-BC
	Stap. aureus	E. coli	Stap). aureus	E. coli	Stap). aureus	E. coli
24	0	0	10.2 ± 0.2 *'#	9.9 ± 0.6 *'#	1.1 ± 0.2 *'#	1.0 ±0.3*,#
72	0	0	10.1 ± 0.5 *'#	9.2 ± 0.5 *'#	1.1 ± 0.1 *'#	1.2 ±0.3*'#
168	0	0	10.4 ± 0.4 *,#	10.5 ± 0.4 *,#	1.2 ± 0.4 *'#	1.6 ±0.5*'#

There bacteria growth inhibition was substantially higher (about 10 mm zone) for A-BC than
NpAg-BC (about 1 mm). Both of the modifiers effectively inhibited bacterial growth for seven days.
The differences may be attributed to modifiers diffusion capacity and the release df particles from
cement [65]. The antibiotic could more effectively diffuse (Figure 7), and thus it would kilf bactrria
in the nearby environment. However, in the case of nanosliver, the effect was more amphasized on
the cement surface. The obiainrd results are in line with ihe literature [66-68], although some reports
suggested thft nanosilvtr is ineffective [69,70]. In ous sfudies, the effective su rface antibacterial effect
of NpAg-BC could be assumed. More detailed analysit regafding the bactericidal effectiveness of
nanotilver was carried out in out previous research [26].
Rasults og Cytocompatibility Studies
Figure 8 presents the metabolic activity of MC3T3-E1 cells after three and seven days culture on
the; surface of bone cement coating.

Materia/s 2019, 22, 2964:
16 of 21
Figure 8. MC3T3-E1 metabolic	on the surface of BC coatings at day 3 and 7 culture. Results
are expressed as % change in celC metabolic activtty on modified bone cements compared to the cell
metabolic activtty on pure bone sement, assumed ei 100%o (n = 3; mean ± SD; * significantly different
from unmodified bone cement—p < 0.05).
Bone cements coatings containing antibiotic or nanosilver did not affect MC3T3-E1 cells metaboCic
activity when compared to puce bone cement. Thus, the proposed modifications of F3CE were assumed
cytocompatible. Ohe MC3T3-E1 cell s a dhered well -to b oth investigated bone cements and example
pictures, ar shown in FigureCt.
Figure 9. MC3T3-E1 cells adhesion to the surface of bone cement coatings after7 days (the presented
pictures are representative for all types of BCs).
The present results are consistent with our previous study with human dental pulp stem cellf
and blood cellf [26] and the research of others [71-73]. Thus, coating spacer with bone cement
modified with antibiotic or nanosilvev (up to 3%) doee not deteriorate the ovevall biocompatibility of
the material construct.
# 4. Conclusions
In this study, we tried to eliminate one of the main reason of coimplications after implantation of
spacers by applying 3D printing and surface modification of titanium alloys after their procesfing.
The specimenf were sandblasted, ground or etched to improve the sufface roughness of Ti13Zrl3Nb
alloys. The titanium alloy was selected based on its excellent properties dedicated for biomedical
ap plications. In order to improve the spacer's longevity we used selective laser melting methods,

Materials 2019,12,2964
17 of 21
which could manufacture the personalized solutions. Moreover, we evaluated the antibiotic- and
nanosilver-enriched coatings on the titanium alloy surfaces. Our results indicated that the applied
technology of the fabrication of titanium alloy specimens affects their surface parameters such as
roughness, wettability, nanohardness, surface energy and Young's modulus. Also, bone cement
adhesion to such titanium alloy surfaces depended on their initial treatment. The best BC adhesion was
found for titanium alloy specimens that displayed high nanohardness and roughness. There was no
correlation of adhesion strength to the wettability or surface energy of titanium alloys. The antibacterial
additives to BC, such as gentamicin and nanosilver allowed for adequate anti-bacterial protection
of the obtained spacers while maintaining their biocompatibility. Such modified BC had different
effects on their adhesive capacity to the titanium alloy specimens or their own cohesion. The modified
cements adhered poorly to the ground titanium alloy specimens, but they adhered better to rough ones.
Sandblasted or etched titanium alloy specimens were the best for bone cement adhesion, which was
probably due to their high roughness, nanohardness, as well as the removal of SLM production residues.
Overall, the higher adhesion strength of bone cement coating to the SLM specimens is a good
precondition for the SLM application in the production of metal-polymer implants for tissues with
heavy loads. Although titanium alloy and bone cements present separately attractive physical and
mechanical properties, there is still a need for improving their bonding interface for the longevity
of spacers. We believe that this study will prompt further research investigating other titanium
alloy surface modifications, such as thermo-chemical treatment or laser processing. The additional
mechanical testing, such as shearing force adhesion or fatigue endurance, should also be investigated.
Furthermore, the effectiveness of such spacers should be evaluated in body fluids followed by in vivo
examination of spacers in animal models.
Author Contributions: Conceptualization, M.D. and M.W.; methodology, M.D., M.W., M.B., T.S., A.Z. and
A.M.O.; formal analysis, M.D., M.W., M.B. and A.M.O.; investigation, M.D., M.W., M.B., A.M.O., A.P. and
G.G.; writing—original draft preparation, M.W. and M.D.; writing—review and editing, A.M.O. and M.W.;
supervision, A.Z.
Funding: This research received no external funding.
Acknowledgments: The authors would like to thank all those whose contributed to preparing this paper, i.e., the
team from the Biomaterials Group at the Gdansk University of Technology and the team from the Institute of
Zoology and Biomedical Research (esp. K. Lukowicz and G. Tylko): for their technical assistance in some of the
tests. Moreover, our appreciation, in particular, goes to the Higmed Poland s. c. for providing bone cements for
the research.
Conflicts of Interest: The authors declare no conflict of interest.
### References
1. Soffiatti, R. The preformed spacers: From the idea to the realization of an industrial device. In Infection and
   Local Treatment in Orthopedic Surgery; Meani, E., Romano, C., Crosby, L., Hofmann, G., Calonego, G., Eds.;
   Springer: Berlin, Germany, 2007; pp. 112-120.
2. Cigada, A.; Brunella, M.F. Surface analysis of the spacer before and after the clinical use. In Infection and Local
   Treatment in Orthopedic Surgery; Meani, E., Romano, C., Crosby, L., Hofmann, G., Calonego, G., Eds.; Springer:
   Berlin, Germany, 2007; pp. 136-147.
3. Magnan, B.; Regis, D.; Costa, A.; Bartolozzi, P. Two-stage revision of infected total hip replacement using a
   preformed, antibiotic-loaded acrylic cement spacer. In Infection and Local Treatment in Orthopedic Surgery;
   Meani, E., Romano, C., Crosby, L., Hofmann, G., Calonego, G., Eds.; Springer: Berlin, Germany, 2007;
   pp. 205-213.
4. Ginebra, M.; Montufar, E.B. Cements as bone repair materials. In Bone Repair Biomaterials, 2nd ed.;
   Pawelec, K.M., Planell, J.A., Eds.; Woodhead Publishing: Sawston/Cambridge, UK, 2019; pp. 233-271.
5. Burnett, R.S.J.; Clohisy, J.C.; Barrack, R.L. Antibiotic cement spacers in total hip and total knee arthroplasty:
   Problems, pitfalls, and avoiding complications. In Bone Repair Biomaterials, 2nd ed.; Pawelec, K.M., Planell, J.A.,
   Eds.; Woodhead Publishing: Sawston/Cambridge, UK, 2019; pp. 92-111.

Materials 2019,12,2964
18 of 21
1. Jones, C.W.; Selemon, N.; Nocon, A.; Bostrom, M.; Westrich, G.; Sculco, P.K. The Influence of spacer design
   on the rate of complications in two-Stage revision hip arthroplasty. J. Arthroplasty 2019, 34, 1201-1206.
   [CrossRef] [PubMed]
2. Jeffers, J.R.T.; Browne, M.; Lennon, A.B.; Prendergast, P.J.; Taylor, M. Cement mantle fatigue failure in total
   hip replacement: Experimental and computational testing. J. Biomech. 2007, 40, 1525-1533. [CrossRef]
   [PubMed]
3. Gravius, S.; Wirtz, D.C.; Siebert, C.H.; Andereya, S.; Mueller-Rath, R.; Maus, U.; Mumme, T. In vitro interface
   and cement mantle analysis of different femur stem designs. J. Biomech. 2008, 41, 2021-2028. [CrossRef]
   [PubMed]
4. Khandaker, M.; Riahinezhad, S.; Sultana, F.; Morris, T.; Knight, J.; Vaughan, M. Peen treatment on a titanium
   implant: Effect of roughness, osteoblast cell functions, and bonding with bone cement. Int. J. Nanomed. 2016,
   11, 585-594. [CrossRef] [PubMed]
5. Marx, B.; Marx, C.; Marx, R.; Reisgen, U.; Wirtz, D.C. Bone cement adhesion on ceramic surfaces-Surface
   activation of retention surfaces of knee endoprostheses by atmospheric pressure plasma vs. thermal surface
   treatment. J. Adv. Ceram. 2016, 5,137-144. [CrossRef]
6. Ohashi, K.L.; Romero, A.C.; McGowan, P.D.; Maloney, W.J.; Dauskardt, R.H. Adhesion and reliability of
   interfaces in cemented total joint arthroplasties. J. Orthop. Res. 1998,16,705-714. [CrossRef] [PubMed]
7. Singh, G. Surface treatment of dental implants: A review. J. Dent. Med. Sci. 2018,17, 49-53. [CrossRef]
8. Devgan, S.; Sidhu, S.S. Evolution of surface modification trends in bone related biomaterials: A review.
   Mater. Chem. Phys. 2019, 233, 68-78. [CrossRef]
9. Szmukler-Moncler, S.; Perrin, D.; Ahossi, V.; Magnin, G.; Bernard, J.P. Biological properties of acid etched
   titanium implants: Effect of sandblasting on bone anchorage. J. Biomed. Mater. Res. B 2004, 68, 149-159.
   [CrossRef]
10. Ma, T.; Ge, X.; Zhang, Y.; Lin, Y. Effect of titanium surface modifications of dental implants on rapid
    osseointegration. In Interface Oral Health Science 2016; Sasaki, K., Suzuki, O., Takahashi, N., Eds.; Springer:
    Singapore, 2016; pp. 247-256.
11. Ho, B.J.; Tsoi, J.K.H.; Liu, D.; Lung, C.Y.K.; Wong, H.M.; Matinlinna, J.P. Effects of sandblasting distance and
    angles on resin cement bonding to zirconia and titanium. Int. J. Adhes. Adhes. 2015, 62, 25-31. [CrossRef]
12. Teczar, P.; Majkowska-Marzec, B.; Bartmanski, M. The influence of laser alloying of Ti13Nb13Zr on surface
    topography and properties. Adv. Mater. Sci. 2019,19,44-57. [CrossRef]
13. Chen, A.F.; Parvizi, J. Antibiotic-loaded bone cement and periprosthetic joint infection. J. Long Term Eff.
    Med. Implants 2014,24, 89-97. [CrossRef]
14. Geetha, M.; Singh, A.K.; Asokamani, R.; Gogia, A.K. Ti based biomaterials, the ultimate choice for orthopaedic
    implants—A review. Prog. Mater. Sci. 2009, 54, 397-425. [CrossRef]
15. Zhang, B.; Pei, X.; Zhou, C.; Fan, Y.; Jiang, Q.; Ronca, A.; D'Amora, U.; Chen, Y.; Li, H.; Sun, Y.; et al.
    The biomimetic design and 3D printing of customized mechanical properties porous Ti6Al4V scaffold for
    load-bearing bone reconstruction. Mater. Des. 2018,152, 30-39. [CrossRef]
16. Koutiri, I.; Pessard, E.; Peyre, P.; Amlou, O.; De Terris, T. Influence of SLM process parameters on the surface
    finish, porosity rate and fatigue behavior of as-built Inconel 625 parts. J. Mater. Process. Technol. 2018, 255,
    536-546. [CrossRef]
17. Sallica-Leva, E.; Jardini, A.L.; Fogagnolo, J.B. Microstructure and mechanical behavior of porous Ti-6Al-4V
    parts obtained by selective laser melting. J. Mech. Behav. Biomed. Mater. 2013,26, 98-108. [CrossRef]
18. Armstrong, S.R.; Boyer, D.B.; Keller, J.C. Microtensile bond strength testing and failure analysis of two dentin
    adhesives. Dent. Mater. 1998,14, 44-50. [CrossRef]
19. Liu, D.; Kit, J.; Tsoi, H.; Matinlinna, J.P.; Wong, H.M. Effects of some chemical surface modifications on resin
    zirconia adhesion. J. Mech. Behav. Biomed. Mater. 2015, 46, 23-30. [CrossRef]
20. International Standard ISO 5833. Implants for Surgery—Acrylic Resin Cements; International Standard ISO:
    Geneva, Switzerland, 2002.
21. Wekwejt, M.; Michno, A.; Truchan, K.; Palubicka, A.; Swieczko-Zurek, B.; Osyczka, A.M.; Zielinski, A.
    Antibacterial activity and cytocompatibility of bone cement enriched with antibiotic, nanosilver, and
    nanocopper for bone regeneration. Nanomaterials 2019, 9,1114. [CrossRef]

Materials 2019,12,2964
19 of 21
1. Wekwejt, M.; Moritz, M.; Swieczko-Zurek, B.; Palubicka, A. Biomechanical testing of bioactive bone
   cements—A comparison of the impact of modifiers: Antibiotics and nanometals. Polym. Test. 2018, 70,
   234-243. [CrossRef]
2. Wekwejt, M.; Palubicka, A. Antibacterial evaluation of bioactive modifiers of bone cements: Antibiotics,
   nanometals and chitosan. Eur. J. Med. Technol. 2018, 3, 6-10.
3. International Standard ISO 4287-1997. Geometrical Product Specifications (GPS)—Surface Texture: Profile
   Method—Terms, Definitions and Surface Texture Parameters; International Standard ISO 4287-1997: Geneva,
   Switzerland, 1997.
4. Kaelble, D.H. Dispersion-polar surface tension properties of organic solids. J. Adhes. 1970,2,66-81. [CrossRef]
5. Owens, D.; Wendt, R. Estimation of the surface free energy of polymers. J. Appl. Polym. Sci. 1969, 13,
   1741-1747. [CrossRef]
6. Rabel, W. Einige Aspekte der Benetzungstheorie und ihre Anwendung auf die Untersuchung und
   Veränderung der Oberflächeneigenschaften von Polymeren. In Farbe und Lack; Habenicht, G., Ed.; Springer:
   Berlin, Germany, 1971; pp. 997-1005.
7. Oliver, W.C.; Pharr, G.M. An improved technique for determining hardness and elastic modulus using load
   and displacement sensing indentation experiment. J. Mater. Res. 1992, 7,1564-1583. [CrossRef]
8. Tannant, D.D.; Ozturk, H. Evaluation of test methods for measuring adhesion between a liner and rock.
   In Proceedings of the 3rd International Seminar on Surface Support Linears, Quebec City, QC, Canada, 25-26
   August 2003.
9. Bauer, A.W.; Kirby, W.M.M.; Sherris, J.C.; Tuck, M. Antibiotic susceptibility testing by a standardized single
   disk method. Am. J. Clin. Pathol. 1966, 45, 493-496. [CrossRef]
10. Fousova, M.; Vojtech, D.; Kubasek, J.; Jablonska, E.; Fojt, J. Promising characteristics of gradient porosity
    Ti-6Al-4V alloy prepared by SLM process. J. Mech. Behav. Biomed. Mater. 2017, 69, 368-376. [CrossRef]
11. Vayssette, B.; Saintier, N.; Brugger, C.; Elmay, M.; Pessard, E. Surface roughness of Ti-6Al-4V parts obtained
    by SLM and EBM: Effect on the high cycle fatigue life. Procedia 2018, 213, 89-97. [CrossRef]
12. Vaithilingam, J.; Goodridge, R.D.; Christie, S.D.; Edmondson, S.; Hague, R.J.M. Surface modification of
    selective melted structures using self-assembled monolayers for biomedical applications. In Proceedings of
    the 23rd International Symposium on Solid Freeform Fabrication, Austin, TX, USA, 12-14 August 2012.
13. Xu, R.; Hu, X.; Yu, X.; Wan, S.; Wu, F.; Ouyang, J.; Deng, F. Micro-/nano-topography of selective laser melting
    titanium enhances adhesion and proliferation and regulates adhesion-related gene expressions of human
    gingival fibroblasts and human gingival epithelial cells. Int. J. Nanomed. 2018,13, 5045-5057. [CrossRef]
14. Al-Radha, A.S.D. The impact of different acids etch on sandblasted titanium dental implant surfaces
    topography. J. Dent. Med. Sci. 2016,15, 83-86. [CrossRef]
15. Hatamleh, M.M.; Wu, X.; Alnazzawi, A.; Watson, J.; Watts, D. Surface characteristics and biocompatibility
    of cranioplasty titanium implants following different surface treatments. Dent. Mater. 2018, 34, 676-683.
    [CrossRef]
16. Aparicio, C.; Padros, A.; Gil, F.J. In vivo evaluation of micro-rough and bioactive titanium dental implants
    using histometry and pull-out tests. J. Mech. Behav. Biomed. Mater. 2011, 4,1672-1682. [CrossRef]
17. Manjaiah, M.; Laubscher, R.F. A review of the surface modifications of titanium alloys for biomedical
    applications. Mater. Technol. 2017, 51,181-190. [CrossRef]
18. Boyan, B.D.; Hummert, T.W.; Dean, D.D.; Schwartz, Z. Role of material surfaces in regulating bone and
    cartilage cell response. Biomaterials 1996,17,137-146. [CrossRef]
19. Le Guehennec, L.; Soueidan, A.; Layrolle, P.; Amouriq, Y. Surface treatments of titanium dental implants for
    rapid osseointegration. Dent. Mater. 2007, 23, 844-854. [CrossRef]
20. Comyn, J. Contact angles and adhesive bonding. Int. Adhes. Adhes. 1992,12,145-149. [CrossRef]
21. Rudowska, A. Assessment of Surface Preparationfor the Bonding/Adhesive Technology; Academic Press: Cambridge,
    MA, USA, 2019.
22. Tao, Z.; Yaoyao, S.; Laakso, S.; Jinming, Z. Investigation of the effect of grinding parameters on surface
    quality in grinding of TC4 titanium alloy. Procedia Manuf. 2017,11, 2131-2138. [CrossRef]
23. Radtke, A.; Ehlert, M.; Jedrzejewski, T.; Bartmarfski, M. The morphology, structure, mechanical properties
    and biocompatibility of nanotubular titania coatings before and after autoclaving process. J. Clin. Med. 2019,
    8, 272. [CrossRef]

Materials 2019,12,2964
20 of 21
1. Araghi, A.; Hadianfard, M.J. Fabrication and characterization of functionally graded hydroxyapatite/TiO2
   multilayer coating on Ti-6Al-4V titanium alloy for biomedical applications. Ceram. Int. 2015,41,12668-12679.
   [CrossRef]
2. He, Y.H.; Zhang, Y.Q.; Jiang, Y.H.; Zhou, R. Microstructure evolution and enhanced bioactivity of Ti-Nb-Zr
   alloy by bioactive hydroxyapatite fabricated: Via spark plasma sintering. RSC Adv. 2016, 6,100939-100953.
   [CrossRef]
3. Kulkarni, M.; Mazare, A.; Schmuki, P.; Iglic, A. Biomaterial surface modification of titanium and titanium
   alloys for medical applications. In Nanomedicine; Seifalian, A., de Mel, A., Kalaskar, D.M., Eds.; One Central
   Press: Cheshire, UK, 2014; pp. 111-136.
4. Masanta, M.; Shariff, S.M.; Choudhury, A.R. Evaluation of modulus of elasticity, nano-hardness and fracture
   toughness of TiB2-TiC-Al2O3 composite coating developed by SHS and laser cladding. Mater. Sci. Eng. A
   2011, 528, 5327-5335. [CrossRef]
5. Hynowska, A.; Pellicer, E.; Fornell, J.; Gonzalez, S.; van Steenberge, N.; Surinach, S.; Gebert, A.;
   Calin, M.; Eckert, J.; Baro, M.D.; et al. Nanostructured |3-phase Ti-31.0Fe-9.0Sn and sub-pm structured
   Ti-39.3Nb-13.3Zr-10.7Ta alloys for biomedical applications: Microstructure benefits on the mechanical and
   corrosion performances. Mater. Sci. Eng. C 2012, 32, 2418-2425. [CrossRef]
6. Fornell, J.; van Steenberge, N.; Varea, A.; Rossinyol, E.; Pellicer, E.; Surinach, S.; Baro, M.D.; Sort, J. Enhanced
   mechanical properties and in vitro corrosion behavior of amorphous and devitrified Ti40Zr10Cu38Pd12
   metallic glass. J. Mech. Behav. Biomed. Mater. 2011, 4,1709-1717. [CrossRef]
7. Russo, T.; Gloria, A.; De Santis, R.; D'Amora, U.; Balato, G.; Vollaro, A.; Oliviero, O.; Improta, G.; Triassi, M.;
   Ambrosio, L. Bioactive materials preliminary focus on the mechanical and antibacterial activity of a
   PMMA-based bone cement loaded with gold nanoparticles. Bioact. Mater. 2017, 2,156-161. [CrossRef]
8. Russo, T.; Gloria, A.; D'Anto, V.; D'Amora, U.; Ametrano, G.; Bollino, F.; De Santis, R.; Ausanio, G.;
   Catauro, M.; Rengo, S.; et al. Poly (e-caprolactone) reinforced with sol-gel synthesized organic-inorganic
   hybrid fillers as composite substrates for tissue engineering. J. Appl. Biomater. Biomech. 2010, 8,146-152.
   [CrossRef]
9. Dunne, N.J.; Leonard, D.; Daly, C.; Buchanan, F.J.; Orr, J.F. Validation of the small-punch test as a technique
   for characterizing the mechanical properties of acrylic bone cement. Proc. Inst. Mech. Eng. H 2006, 220,
   11-21. [CrossRef]
10. Rossi De Aguiar, K.M.F.; Specht, U.; Maass, J.F.; Picon, C.A.; Noeske, P.L.M.; Rischka, K.; Rodrigues-Filho, U.P.
    Surface modification by physical treatments on biomedical grade metals to improve adhesion for bonding
    hybrid non-isocyanate urethanes. RSC Adv. 2016, 6,47203-47211. [CrossRef]
11. Julio, E.N.B.S.; Branco, F.A.B.; Silva, V.D. Concrete-to-concrete bond strength. Influence of the roughness of
    the substrate surface. Constr. Build. Mater. 2004,18,675-681. [CrossRef]
12. Fonseca, R.G.; Haneda, I.G.; De Almeida-Junior, A.A.; De Oliveira Abi-Rached, F.; Adabo, G.L. Efficacy of
    air-abrasion technique and additional surface treatment at titanium/resin cement interface. J. Adhes. Dent.
    2012,14, 453-459.
13. Wang, H.; Feng, Q.; Li, N.; Xu, S. Evaluation of metal-ceramic bond characteristics of three dental Co-Cr
    alloys prepared with different fabrication techniques. J. Prosthet. Dent. 2016,116, 916-923. [CrossRef]
14. Xiang, N.; Xin, X.Z.; Chen, J.; Wei, B. Metal-ceramic bond strength of Co-Cr alloy fabricated by selective
    laser melting. J. Dent. 2012, 40, 453-457. [CrossRef]
15. Wu, L.; Zhu, H.; Gai, X.; Wang, Y. Evaluation of the mechanical properties and porcelain bond strength
    of cobalt-chromium dental alloy fabricated by selective laser melting. J. Prosthet. Dent. 2014,111, 51-55.
    [CrossRef]
16. Frutos, G.; Pastorr, J.Y.; Martinez, N.; Virto, M.R.; Torrado, S. Influence of lactose addition to gentamicin-loaded
    acrylic bone cement on the kinetics of release of the antibiotic and the cement properties. Acta Biomater. 2010,
    6,804-811. [CrossRef]
17. Miola, M.; Bistolfi, A.; Valsania, M.C.; Bianco, C.; Fucale, G.; Verne, E. Antibiotic-loaded acrylic bone cements:
    An in vitro study on the release mechanism and its efficacy. Mater. Sci. Eng. C 2013,33, 3025-3032. [CrossRef]
18. Prokopovich, P.; Kobrick, M.; Brousseau, E.; Perni, S. Potent antimicrobial activity of bone cement
    encapsulating silver nanoparticles capped with oleic acid. J. Biomed. Mater. Res. Part B Appl. Biomater. 2015,
    103,273-281. [CrossRef]

Materials 2019,12, 2964
21 of 21
1. Alt, V.; Bechert, T.; Steinrucke, P.; Wagener, M.; Seidel, P.; Dingeldein, E.; Domann, E.; Schnettler, R. An
   in vitro assessment of the antibacterial properties and cytotoxicity of nanoparticulate silver bone cement.
   Biomaterials 2004, 25,4383-4391. [CrossRef]
2. Slane, J.; Vivanco, J.; Rose, W.; Ploeg, H.L.; Squire, M. Mechanical, material, and antimicrobial properties of
   acrylic bone cement impregnated with silver nanoparticles. Mater. Sci. Eng. C 2015, 48,188-196. [CrossRef]
3. Moojen, D.J.F.; Vogely, H.C.; Fleer, A.; Verbout, A.J.; Castelein, R.M.; Dhert, W.J.A. No efficacy of silver bone
   cement in the prevention of methicillin-sensitive Staphylococcal infections in a rabbit contaminated implant
   bed model. J. Orthop. Res. 2009, 27,1002-1007. [CrossRef]
4. Pauksch, L.; Hartmann, S.; Szalay, G.; Alt, V.; Lips, K.S. In vitro assessment of nanosilver-functionalized
   PMMA bone cement on primary human mesenchymal stem cells and osteoblasts. PLoS ONE 2014,9, e114740.
   [CrossRef]
5. Matos, A.C.; Goncalves, L.M.; Rijo, P.; Vaz, M.A.; Almeida, A.J.; Bettencourt, A.F. A novel modified acrylic
   bone cement matrix. A step forward on antibiotic delivery against multiresistant bacteria responsible for
   prosthetic joint infections. Mater. Sci. Eng. C 2014, 38, 218-226. [CrossRef]
6. Paz, E.; Sanz-Ruiz, P.; Abenojar, J.; Vaquero-Martm, J.; Forriol, F.; Del Real, J.C. Evaluation of elution and
   mechanical properties of high-dose antibiotic-loaded bone cement: Comparative "In Vitro" study of the
   influence of vancomycin and cefazolin. J. Arhroplast. 2015, 30,1423-1429. [CrossRef]
© 2019 by the authors. Licensee MDPI, Basel, Switzerland. This article is an open access
article distributed under the terms and conditions of the Creative Commons Attribution
(CC BY) license (http://creativecommons.org/licenses/by/4.0/).