Signature:©PolJRadiol,2010;75(1):61-66
ORIGINALARTICLE
TheeffectsoftargetmotioninkV-CBCTimaging
Received:2009.10.21
Accepted:2009.10.28
SriramPadmanaban,RaghavendiranBoopathy,BhuvanaKunjithapatham,PrabakarSukumar,VivekanandanNagarajan
CancerInstitution(WIA),AdyarChennaiTamilnadu,India
Author’saddress:SriramPadmanaban,CancerInstitution(WIA),AdyarChennaiTamilnaduIndia,e-mail:srirampadm@gmail.com
Summary
TostudytheimpactoftargetmotioninkV-CBCTimaging.
Tosimulatetherespiratorymovement,dynamicphantomwasprogrammedtomoveinthree-dimensionwithaperiodoffoursecondsandoftwodifferentamplitudes(PA1andPA2).Thetargetsofwelldefinedgeometries(cylinder,sphere,solidtriangle,U-shapedanddumbbell)weremadeusingwax.ThestatictargetswereCTimaged(referenceimage).UsingCBCT,thetargetsinstaticanddynamicmodeswereimagedunderfull-fanbeam.Thelineprofilesalongcranio-caudaldirection,influenceoftarget'sinitialmovingphaseandvolumeestimationusingauto-contouringtoolwereusedtoanalyze theeffectsoftarget motiononCBCTimaging.
ComparingthelineprofilesoftargetsinCBCTwithCT,thelengthofaverageHUspreadwasreducedby42.54±1.85%,exceptthecylindricaltargetwhichisby19.35%forPA1.Thepercentagedifference inreconstructedvolumeofstatictargetsimagedusing CBCTandCT(HUWW–500to0)rangesfrom–1.32%to–5.94%.ThevolumelossesfortargetsimagedindynamicmodePA1rangesfrom14.35%to30.95%andforPA2itwas21.29%to43.80%.Thesolidtriangleandcylindricaltargetssufferedthemaximumandminimumvolumelossesrespectively.
Non-gatedCBCTimagingofthemovingtargetsencounterssignificantlossofvolumetricinformation,duetoscatterartifacts.Thesemayresultinasystematicerrorinre-contouringwhenCBCTimagesareusedforthere-planningwork.
Background:
Material/Methods:
Results:
Conclusions:
Keywords:
kV-CBCT•targetmotion•re-contouring•systematicerror
PDFfile:
http://www.polradiol.com/fulltxt.php?ICID=878435
Background
DifferentfromconventionalCT,kVCBCTcoversamuchlargerfieldofview(FOV)inthelongitudinaldirectionandscatterposesamoresevereproblemintheresultantimage.Alsothegantryrotationspeedislimitedto~1minbyIECregulation,whichmakestheCBCTmorepronetomotionartifacts[6].Thedeterioratedimagequalityraisesseriousconcernsaboutthepreciseestimationofreconstructedvolumeofmovingtargetswhichinfluenceontheaccuratedosimetricre-planningpurposes.Thisstudyevaluatedthechangesinvolumetricinformationofmovingtargetsusingin-housebuiltdynamicphantom.
Kilo-voltageconebeamcomputedtomography(kV-CBCT)integratedintothelinearacceleratorasatoolforimageguidedradiationtherapy,wasdevelopedtoacquireon-linevolumetricandanatomicalimages[1].TheCBCTimagesarecomparedwiththereconstructedradiographfromorig-inaltreatmentplanforthemostaccuratepositionsetup[2–5].Inadditiontoguidethepatientsetupprocess,CBCTdataacquiredpriortothetreatmentbeusedtorecalculateorverifythetreatmentplanbasedonthepatientanatomyofthetreatmentday.Thisaffordsaneffectivemeanstoadaptivelymodifythepatient'streatmentplanduringthecourseofaradiationtherapybasedonthedosethathasalreadybeendelivered.Butthisapplicationinimageguid-ancedependsontheimagequalityofCBCT[6].
MaterialandMethods
Anin-housebuiltdynamicphantomcapableofmovinginthree-dimensionswasconstructed(Figure1).Eachaxiswasbuildindependentlyusingsteppermotorandlead
61

OriginalArticle
©PolJRadiol,2010;75(1):61-66
Figure1.In-housebuiltdynamicphantomhavingcylindricaltargetswasimagedinkV-CBCTintegratedwithVarianCL-2100C/D.
Figure3.Thedifferentshapedwaxtargets:Sphere,Cylinder,Solidtriangle,DumbbellandU-shaped.
Figure2.(A).Across-sectionalimageofCATPHAN600imagedusingCT.(B)Across-sectionalimageofCATPHAN600imagedusingfull-fanCBCT.
(C)A cross-sectionalimage ofCATPHAN600imaged usinghalf-fanCBCT.
screwassemblydrivenbyindividualsteppercard.Thepositionalaccuracyofthephantomwas5µm.Tosimulatetherespiratorymovement,in-housebuiltdynamicphan-tomwasmadetomovelinearlywithaperiodoffoursec-ondsandoftwodifferentprogrammedamplitudes(PA):
(i)amplitudeof18mmincranio-caudal,14mminante-rior-posteriorand4mminlateraldirections(PA1)and(ii)amplitudeof24mmincranio-caudal,16mminanterior-posteriorand4mminlateraldirections(PA2)[7].
Thetargetsofwelldefinedgeometriesandknownvol-umenamelycylinder(42.3cc),sphere(21.1cc),solidtrian-gle(27.2cc),U-shaped(38.9cc)anddumbbell(44.3cc)weremadeusingwax.Thepurposeofdifferentshapedtargetsistounderstandthedependenceofthereconstructedvol-umeoftargetsasafunctionofshapeandamplitudeofitsmotion.Figure3showsthedifferentshapedwaxphan-toms:Sphere,Cylinder,Solidtriangle,DumbbellandU-shaped.
TheCBCTimagesinthisworkwereacquiredbytheon-board imager(OBI®,VarianMedicalSystems,PaloAlto,CA)integratedonaVarianCL-2100C/Dlinearaccelerator.InordertoobtainthecalibrationcurveofconventionalCTandCBCT,Catphan-600moduleCTP404wasused.TheCTP404hasadiameterof15cm(longitudinaldimensionis16cm)andcontainsdifferentsizesinsertswithsevendif-ferenttissuesubstitutematerials,air,PMP,LDPE,polysty-rene,acrylic,DelrinandTeflonrespectively.Theirelectrondensitiesrangedfrom0to2.16.ThecalibrationofaCTscannerinvolvesacquiringCTimages,obtainingaverageHUsforeachinsertingmaterials,andplottingtheHUasafunctionoftherelativeelectrondensity.Similarly,thecali-brationcurvewasobtainedforCBCTinfull-fanandhalf-fanmode.ThecalibrationcurvesofCT,full-fanCBCTandhalf-fanCBCTwerecompared.CrosssectionalimageofthephantomimagedusingCT,full-fanandhalf-fanCBCTwereshowninFigure2A–C.
Thesetargetswereimagedinstaticmodeusingconven-tionalCTwiththeimagingparametersof120kV,250mAs,andpitch 1.5 to yieldimages of 2 mmslice thickness. TheseCTimageswereconsideredasthereferenceimagesfortheanalysis.
DuringCBCTimageacquisitionthetargetswerefixedabovetheStyrofoamboardandimagedatbothstaticanddynamicmodesunderfull-fanacquisitionwithbowtiefil-terandimagingparameters:100kV,80mA,and512×512resolution.IndynamicmodetheimageswereacquiredforboththePA1andPA2.Thevolumetricdatawererecon-structedtoyieldimagesof2mmslicethickness;
InordertostudytheinfluenceofinitialmovingphaseduringCBCTimaging,thesphericaltargetisimagedeighttimesatrandominitialphasesindynamicmode(PA1).Itisfoundthat therewas nosignificant changein volume(stan-darddeviation=0.15).
62
A	B	C

©PolJRadiol,2010;75(1):61-66
PadmanabanSetal–TheEffectsoftargetmotioninkV-CBCTimaging
Figure4.Thecalibrationcurves(CTnumberversusRelativeelectrondensity)forCT,full-fanandhalf-fanCBCT.
Figure5.TheHUlineprofilesofthetargets(A)cylinder,(B)dumbbell
(C)sphere(D)U-shape(E)SolidtriangleimagedwithCTand CBCT instaticand dynamicmode along cranio-caudaldirection.
TheCTandCBCTimageswereexportedtoEclipsetreat-mentplanning system(Version8.6,VarianMedicalSystems).Thelineprofilesareobtainedalongcranio-caudaldirectionforstaticanddynamictargetsinCBCTimages.Thesepro-filesarecomparedwithprofilesofreferenceCTimages.
–112±6HU,hencetheupperwindowlevelwaschosenas0HU.ItisfoundthatinCTimagesthewindowwidthsettingof–500to0HUyieldthevolumeswithlessthanoneper-centvariationtotheactualvolumes.ThesevolumeswereconsideredasthereferencevolumesforcomparisonwithCBCTimages.InCBCTimagesthestaticandmoving(PA1
&PA2)target volumeswere delineatedfor variouswindowwidthsettingsandweretabulated.Thepercentagediffer-encesinvolumeweredeterminedandchangesinrecon-structedvolumeofmovingtargetswereanalyzed.
Sincethevolumepredictionusingautocontouringtech-niquehighlydependsonthewindowwidthlevels(WW),itismandatorytofindtheappropriateWWsetting.FromtheCTimagestheaverageHUofthetargetwasfoundtobe
63
A

B

C

D

E

OriginalArticle
©PolJRadiol,2010;75(1):61-66
Table1.Thereconstructedvolume(incc)ofthetargetsimagedwithCBCTinstaticanddynamicmodesobtainedusingauto-contouringforWW
-500to0andthepercentagedifferenceinvolumewithreferenceCTrespectively.
ReconstructedVolumeusingauto-contouringtechniqueforWW–500to0(incc)
Targets
ReferenceCTvolume
CBCT
(static)
Volumedifference(%)
CBCT
(dynamicPA1)
Volumedifference(%)
CBCT
(dynamicPA2)
Volumedifference(%)
Dumbbell
43.71
42.37
–3.07
31.94
–26.93
25.94
–40.65
Cylinder
42.17
41.61
–1.32
36.12
–14.35
33.19
–21.29
U-shape
38.19
35.92
–5.94
27.82
–27.15
22.02
–42.34
Solidtriangle
26.53
25.01
–5.73
18.32
–30.95
14.91
–43.80
Sphere
20.68
20.13
–2.66
15.65
–24.32
12.24
–40.81
lengthwasreducedby19.35%and25.81%forPA1andPA2respectively.ThesolidtriangleandU-shapetargetssufferedthemaximumreductioninlengthofaverageHU.Fromtheprofiles,itisalsofoundthat,thespreadofthereducedHUvaluesoutsidetheCTprofilewasfor6.4±0.9mmand8±1.4mmforthetargetsmovinginPA1andPA2respectively.
ThechangeinreconstructedvolumewereevaluatedoncomparingthevolumesofreferenceCTimagesandCBCTimages(bothstaticanddynamic)obtainedusingauto-contouringforWW–500to0andWW–600to0.Also,inCBCTimages,WWsettingswereadjustedforeachtargetshapessothatthevolumevariationfallslessthan1%whencomparedwiththereferencetargetvolumesandthevaluesweretabulated.
FortheHUWW–500to0,thepercentagedifferenceinreconstructedvolumeofstatictargetsimagedusingCBCTandreferenceCTrangesfrom–1.32%to–5.94%(Table1).TheU-shapedtargetsufferedthemaximumvolumeloss.ThevolumelossesfordifferenttargetsimagedindynamicmodeforPA1rangesfrom14.35%to30.95%andforPA2itwas21.29%to43.80%.Thesolidtriangleandcylindricaltargetssufferedthe maximumand minimumvolumelossesrespectively.InFigure6,theyellowlineindicatesthecon-touringoftargetsforHUWW–500to0.ItobviouslyshowsthevolumelossforthetargetsimagedindynamicmodesusingCBCToncomparisonwithreferenceCT.
Figure6.ThereconstructedvolumelossforthetargetsimagedinstaticanddynamicmodesusingCBCToncomparisonwithreferenceCT.TheyellowlineindicatesthecontouringoftargetsforHUWW–500to0.
Results
ThecalibrationcurvesofconventionalCT,full-fanandhalf-fanmodeCBCTareshowninFigure4.Infull-fanmode,themaximumdifferencesinHUvaluecomparedwithreferenceCTwerelessthan20HU.Forhalf-fanmode,bonematerialshowsthemaximumHUdifferencefromCTof40HU.ThisstudyconfirmstheCBCThastheabilitytogenerateimageswithHUvaluescomparabletothoseinCT.
SimilarlyfortheHUWW–600to0(Table2),thevolumelossesfortargetsimagedindynamicmodeforPA1werereducedanditrangesfrom2.51%to11.38%andforPA2itwas5.76%to21.07%.Thesolidtriangleandcylindricaltargetssufferedthe maximumand minimumvolumelossesrespectively.
ThelineprofilecomparisonofreferenceCTwithstaticanddynamicCBCTimagesalongcranial-caudaldirectionforvariousshapedtargetsareshowninFigure5.ItisfoundthattheCTandstaticCBCTprofilesoftargetsagreetowithin5%.Ontheotherhand,targetsimagedindynam-icmodeusingCBCTshowsamuchgreatervariation.Theactuallengthofthetargetsalongcranio-caudaldirectionisshownindash-dottedlinefortheCTprofiles.Withintheregionofactuallength,thelengthofaverageHU(–112±6HU)spreadinthetargetswasreducedby42.54±1.85%and55.79±1.80%exceptthecylindricaltargetwherethe
Table3indicatesthatHUWWbetween–650/–620to0andHUWW–675/–635to0resultedinlessthan1%volumelossesforallthetargetsimagedusingCBCTindynamicmodePA1andPA2respectively.
Discussion
TherehavebeenmanystudiesandeffortstoquantifyandfixthemotionartifactsonCBCT.Guckenbergeretal.
64

©PolJRadiol,2010;75(1):61-66
PadmanabanSetal–TheEffectsoftargetmotioninkV-CBCTimaging
Table2.Thereconstructedvolume(incc)ofthetargetswithCBCTindynamicmodesobtainedusingauto-contouringforWW–600to0andthepercentagedifferenceinvolumewithreferenceCTrespectively.
ReconstructedVolumeusingauto-contouringtechniqueforWW–600to0(incc)
Targets
ReferenceCTvolume
CBCT
(dynamicPA1)
Volumedifference(%)
CBCT
(dynamicPA2)
Volumedifference(%)
Dumbbell
43.71
39.24
–10.23
35.25
–19.35
Cylinder
42.17
41.11
–2.51
39.74
–5.76
U-shape
38.19
35.91
–5.97
30.68
–19.66
Solidtriangle
26.53
23.51
–11.38
20.94
–21.07
Sphere
20.68
19.34
–6.48
16.65
–19.49
Table3.TheWWsettingsforwhichthevolumeofthetargetsislessthanonepercentvariationwithreferenceCTvolumes.
CBCT(PA1)
CBCT(PA2)
Ref.
CTvolume(cc)
Targets
Volume(cc)
%difference
HUWW
Volume(cc)
%difference
HUWW
Dumbbell
43.71
43.62
–0.21
–645to0
43.61
–0.23
–665to0
cylinder
42.17
42.03
–0.33
–620to0
42.11
–0.14
–635to0
U-shape
38.19
38.02
–0.45
–650to0
38.15
–0.10
–675to0
Solidtriangle
26.53
26.30
–0.87
–650to0
26.33
–0.75
–675to0
Sphere
20.68
20.49
–0.92
–635to0
20.62
–0.29
–660to0
appliedaprobabilitydensityfunction(PDF)toidentifythepositionsandmovingextentsofthediaphragmaticdomeonfanbeamCTastheindicatorsoflivermotion.Theunad-justedCBCTimageswerematchedinrangeaccordingly[8].Snokeetal.acquiredfour-dimensionalCBCT(4D-CBCT)andsortedtheimagesatthesamerespiratoryphasesforrecon-struction[9].Thismethodaidedinestimationofaveragemovingtrack,positionandshapeofthetargetandtogetherreducesthegeometricaluncertaintiesfromrespiration,andissimilartotheclinicallyusedgatedsystemforradiothera-py.Thistechniqueisstilllimitedinthesmallernumberofsortedimagesandthereducedreconstructionquality.
wherethelengthwasreducedby19.35%.ThespreadofthereducedHUvaluesoutsidetheCTprofilewasforapproxi-mately6to8mmforthemovingtargets.ThisisbecauseofpartialvolumeaveragingattheedgesofthetargetssimilartofindingsbyJamesATanyietal.inCTimaging[12].
Instaticcondition,targetsimagedinCBCTsufferrecon-structedvolumelossesupto6%oncomparingwithref-erenceCTvolume.Incaseofmovingtargets,therecon-structedvolumelossesshowasignificantvariation.Inourstudy,usingtwodifferentamplitudesinthreedimensionmovingwithsamefrequency(15cycles/min)foracquiringmovingtargets,thereconstructedvolumelossesobtainedusingauto-contouringforWW–500to0rangesfrom14.35%to30.95%forPA1andfrom21.29%to43.80%forPA2.Whenincreasingthewindowwidthforauto-contour-ingtechnique(–600to0),thevolumelosseswerereducedanditrangesfrom2.51%to11.38%forPA1andfrom5.76%to21.07%forPA2.Thisindicatesthatthevolumelossesdependontheamplitudeofmotionandshapeofthetar-gets.Largertheamplitudeofmotionresultsingreatervol-umelosses.Thecylindricaltargetbeingmoreregularalongthecranio-caudal directionsuffers theminimum loss.SincetheU-shapeand solidtriangletargets hasnarrow andsharpedgegeometryalongthecranio-caudaldirection,thereconstructedvolumelossesweremorepronounced.
Whenusingnon-gatedCBCTforthepotentialre-contour-ingandre-planningwork,itisessentialtounderstandthemotioninducedinaccuracy.Fromthestudy,onverifyingtheimage quality ofCBCT,the differenceinHUvaluesbetweenCTandCBCTfortheCatphanwerelessthan40HU,Teflonshowsthemaximumvariationinhalf-fanmode.ThusthisstudyconfirmstheCBCThastheabilitytogener-ate imageswith HUvalues comparabletothose inCT.
InCBCT,duetoconebeamgeometry,theinfluenceofscat-terradiationismorethatoffanbeamgeometry[10,11].Ingeneral,X-rayscatterreducesimagecontrast,increasesimagenoiseandmayintroducereconstructionerrorintoCBCT.Suchartifactsaremuchworseformovingobjects.ComparingthelineprofilesofstatictargetsimagedinCBCTandCTimagesalongcranio-caudaldirectionsshowslesservariationwithintheregionofactuallength.Butincaseofmovingtargets,thelengthofaverageHUspreadwasreducedby42.54±1.85%exceptthecylindricaltarget
Incomputedtomography(CT),theexperimentsandsim-ulationsshowssignificantdistortionsoftrueshapeofanobjectwhenimagedatdifferentinitialphasesofmotion.[13]Inourstudy,theCBCTimagingofsphericaltargetsindifferentinitialphasesdoesnothavesignificantimpacton
65

OriginalArticle
©PolJRadiol,2010;75(1):61-66
Conclusions
theestimatedvolume(standarddeviation:0.15).UnlikeCT,theimagingofmovingtargetsinCBCTyieldslesserdistor-tionofshape.ButtheCTnumberdensitywassmearedatthe edgeswhich causesblurring ofthe targets.Hence whileauto-contouring,thetargetssufferlossofvolumetricinfor-mation.OntheotherhandtherelaxedWWsettingresultsinvolumegainbecauseofthescatteredHU'sbeyondtheactualvolume.Thisresultsofteninlessaccuratethree-dimensionalcontouringofgrosstargetvolume.
Accuratedelineationoftargetvolumeinadaptiveradio-therapyusingkV-CBCTimagingisnotonlycrucialfordesigninganappropriateandclinicallyeffectivetreatmentplan,butalsonecessaryforaccuratedosecalculation.Thestudysuggestsaconclusionduringnon-gatedCBCTacqui-sition:non-gatedCBCTimagingofthemovingtargetscanleadtosignificant lossof volumetricinformation andlesserdistortionofshape.Thesemayresultinasystematicerrorinre-contouringwhenCBCTimagesareusedforthere-planningwork.ThusafailureinaccuratedelineationoftargetintheprocessofadaptiveradiotherapyusingCBCTimagingimposesalimitationonthetherapeuticgainsofconformallyimplementedRTtechniques.
Thisstudywaslimitedtowaxtargets.Furtherinvestiga-tionsarerequiredforvariousdensitytargetsunderslowscanCBCT.
References:
1. XingL,ThorndykeB,Schreibmannetal:Overviewofimageguidedradiationtherapy.MedDosim,2006;91–112
2. JaffrayDA,SiewerdsenJH:Cone-beamcomputedtomographywithaflat-panelimager:Initialperformancecharacterization.MedPhys,2000;27:1311–23
3. JaffrayDASiewerdsenJH,WongJWetal:Flatpanelcone-beamcomputedtomographyforimage-guidedradiationtherapy.IntJRadiatOncolBiolPhys,2002;53:1337–49
4. YanD,LockmanD,BrabbinsDetal:Anoff-linestrategyforconstructingapatient-specificplanningtargetvolumeinadaptivetreatmentprocessforprostatecancer.IntJRadiatOncolBiolPhys,2000;48:289–302
5. LetourneauD,WongR,MoseleyDetal:Onlineplanninganddeliverytechniqueforradiotherapyofspinemetastasesusingcone-beamCT:imagequalityandsystemperformance.IntJRadiatOncolBiolPhys,2001;67:1229–37
6. YangY,SchreibmannE,LiTetal:Evaluationofon-boardkVcone-beamCT(CBCT)-baseddosecalculation.PhysMedBiol,2007;52:685–705
1. StevensCW,MundenRF,ForsterKMetal:Respiratory-drivenlungtumormotionisindependentoftumorsize,tumorlocation,andpulmonaryfunction.IntJRadiatOncolBiolPhys,2001;51(1):62–68
2. GuckenbergerM,SweeneyRA,WilbertJetal:Image-guidedradiotherapyforlivercancerusingrespiratory-correlatedcomputedtomographyandcone-beamcomputedtomography.IntJRadiatOncol BiolPhys,2008; 71:297–304
3. SnokeJJ,ZijpL,RemeijerPetal:RepiratorycorrelatedconebeamCT.MedPhys,2005;32:1176–86
4. SiewerdsenJH,JafferyDA:Cone-beamcomputedtomographywithaflat-panelimager:magnitudeandeffectsofx-rayscatter.MedPhys,2001;28:220–31
5. YooS,YingFF:Dosimetricfeasibilityofcone-beamCT-basedtreatmentplanningcomparedtoCT-basedtreatmentplanning.IntJRadiatOncolBiolPhys,2006;66:1553–61
6. TanyiJA,FussM,VarchenaV et al: Phantom investigation of3Dmotion-dependentvolumealiasingduringCTsimulationforradiationtherapyplanning.RadiatOncol,2007;2:10
7. ChenGTY,KungJH,BeaudetteKP:Artifactsincomputedtomographyscanningofmovingobjects.SeminRadiatOncol,2004;14:19–26
66