https://doi.org/10.1007/s10967-018-6107-7

Medical activated charcoal tablets as a cheap tool for passive monitoring of gaseous 131I activity in air of nuclear medicine departments

Tomasz Mróz1,3 • Kamil Brudecki2 • Jerzy W. Mietelski2 • Mirosław Bartyzel2 • Ryszard Misiak2 •
Andrzej Kornas3

Received: 3 June 2018 / Published online: 14 August 2018
© The Author(s) 2018

Abstract
It is well known that monitoring of radioactivity released from nuclear medicine departments is necessary to ensure the radiological safety of patients and personnel. Unfortunately, equipment for air sampling is often expensive, loud and is not suitable to use in hospitals. Our goal was to find cheap and simple system for passive monitoring of 131I activity concentration in the air of nuclear medicine departments. Medical activated charcoal tablets were used, because charcoal is excellent material for 131I trapping and tablets are readily available. Our proposed sampling protocol contains tablets preparation, exposure and measurements using HPGe detector. Different methods of tablets preparation (drying, impregnation with KI or NaOH) were tested while an experimental chamber was prepared for estimating 131I (released from Na131I, similar to that used in therapy) trapping efficiency of tablets in different conditions. Finally, tablets were placed in plastic holders and tested in nuclear medicine facilities.

Keywords Passive monitoring · Nuclear medicine · Radioiodine · Gamma-ray spectrometry

# Introduction

In modern nuclear medicine, 131I is still one of the most widely used radionuclides as a short-lived (T1/2 = 8.03 d, where it’s effective half-life time in organisms is nearly equivalent to the physical half-live [1]) beta emitter for therapy and diagnostics of the thyroid gland. It can be used in many chemical forms, of which Na131I is the most popular, but other compounds such as 131I-metaiodoben- zoguanidine and 131I-norcholesterol are also used [2]. Because activities of 131I used in medicine are very high

(up to several dozen GBq), there is an enhanced risk of releasing 131I into the air and also to septic tanks systems [3]. There are several ways wich 131I can be released, e. g. opening packages with Na131I (used as a standard chemical form of 131I wich can be administrated for patients) or emitting gaseous 131I by patient (by breathing or perspi- ration). The gaseous fraction can be inhaled, and 131I can be deposited in thyroids of medical staff (technicians, nurses and doctors). 131I has been detected in medical staff using whole body spectrometer (WBS) and HPGe detec- tors. Activities of 131I measured in WBS were in the range from 5 ± 2 to 217 ± 56 Bq corresponding annual thyroid
 equivalent dose ranging from 0.4 mSv (for a female nurse)
& Tomasz Mro´z tomasz.mroz@uj.edu.pl

1	Institute of Physics, Jagiellonian University, Łojasiewicza 11 st., 30-348 Cracow, Poland
2	The Henryk Niewodniczanski Institute of Nuclear Physics Polish Academy of Sciences, Radzikowskiego 152 st.,
31-342 Cracow, Poland
3	Institute of Biology, Pedagogical University of Cracow, Podchorazych 2 st., 30-084 Cracow, Poland
to 15.5 mSv (for a male technician) [4]. Due to these results, measurements of airborne 131I activity were per- formed at the Department of Endocrynology and Nuclear Medicine Holycross Cancer Centre (E&M HCC Kielce, Poland) where the previously measured medical staff is worked. Results showed, that the activity of 131I in the work space ranged from 28 ± 1 to 492 ± 4 Bq m-3 depending on location (nurses station or hot room) which was similar to results obtained by other authors [2, 5]. With

this levels of 131I concentration in the workplace the annual inhalation effective dose can range from 0.47 mSv (female nurse) to 1.3 mSv (male technician) [6]. 131I activity was measured with a charcoal trap covered with a Petryanov filter cloth and was connected to a mobile HVS 30 aerosol sampler. This combination, allowed us to measure both, aerosols (trapped on filter) and the gaseous fraction of 131I with high efficiency. Unfortunately, this equipment is not a good choice to be routinely used in hospitals due to the size and noise of the unit. Therefore, we developed a passive 131I monitoring technique, similar to 222Rn monitoring with a PicoRad system. In order to introduce a cheap and simple system, charcoal tablets, which are readily available and are commonly used in medicine, were used. Because charcoals are often impregnated, (e.g. with KI or Ag salts) some modification of tablets to compare them with non- treated tablets were performed. This study is similar to that presented by Jime`nez et al. [2] but the activity of 131I is measured directly, without the chemical treatment, that is necessary for LSC measurements. It is a practical approach, because nuclear medicine departments are equipped with counting systems, and therefore facilities can make measurements independently.

# Materials and methods

## Charcoal tablets preparation

During the in situ experiment three methods of tablet preparation for 131I trapping were tested. First, tablets were impregnated in 2% solution of KI in water. The second was impregnation of tablets with 2% NaOH solution. Both the KI and NaOH tablets were baked at 105 °C and impreg- nated by dropping of solution onto surface of the tablets using a plastic Pasteur pipette. After impregnation, tablets were held at room temperature by one hour and then were baked again for 24 h at 105 °C to remove water. The third method of tablet preparation used only baking, without any wet impregnation.

## Containers

Empty PicoRad containers were filled with dry silica gel to reduce moisture exposure of tablets during transport and to stabilize moisture variations during sampling. We have used silica gel with a color indicator, which slowly becomes greenish during exposure of the tablets at the hospital, indicates that the silica gel still had some moisture capacity. Dry charcoal tablets were placed into small plastic holders with Petryanov filter cloth on the top and on the bottom side. For each PicoRad container we have
placed two tablets. Then, containers were closed, sealed with Parafilm and sent to the hospital for 131I collection.

## Sampling and activity measurement

Patient rooms and hot rooms were chosen as a sampling sites. In both sites we placed a set of three containers and left them open for 1 week. After exposure, containers were sent back to our laboratory for gamma-ray spectrometry measurements, performed using HPGe detector with 9% relative efficiency constructed at the Institute of Nuclear Physics Polish Academy of Sciences. Separate measure- ments of the tablets, silica gel and filter cloths were performed.

## Preparation of calibration chamber

The sampling system was also tested using a calibration chamber and a system for producing tracer isotopes. Pro- duction of tracer (126I) is possible by irradiation of non- enriched TeO2 with a proton beam at 30 MeV (corre- sponding to the highest cross-section for 126I [7]), using
AIC-144 cyclotron (Institute of Nuclear Physics PAS). Activated targets are heated in the furnace up to 750 °C and gaseous iodine is trapped in 0.01 N NaOH solution. From the basic solution, iodine can be released in gaseous 126I2 by adding concentrated HNO3. The test chamber was prepared from a glovebox sealed with chemical-resistant silicone and equipped with a Geiger based counting system for detecting eventual leaks form the chamber.

# Results and discussion

Results of the gamma spectrometry measurements are presented in Table 1. There was no 131I detected on sil- icagel or the filter cloths. The mean trapping efficiency calculated by the simple activity ratio method ranges from 0.13 to 2.9%.
The highest activity in both sampling stations was measured in tablet, which were baked at 105 °C without any wet treatment. The total adsorption of 131I on activated charcoal with chemical modification is the sum of the adsorption process on charcoal itself and chemical trapping
by the impregnating agent. In the case of KI and NaOH impregnation these processes can be described as follows:
KI þ 131I2 ! K131I3	ð1Þ
131I2 þ 2NaOH ! Na131I þ Na131IO þ H2O	ð2Þ
However, impregnation can cause a reduction of active surface and with higher concentrations of impregnating solution absorption will be reduced due to occupation of

Table 1 131I activities measured in charcoal tablets after

Sampling location	Tablets type	Gaseous

131I activity (Bq)
exposure
Hot room	Baked only	4.71 ± 0.82 2% KI and baked	3.33 ± 0.74
2% NaOH and baked	0.83 ± 0.48
Nurses station	Baked only	0.59 ± 0.36 2% KI and baked	0.19 ± 0.16
2% NaOH and baked	0.23 ± 0.18

the active surface [8]. Also, the impregnation technique is likely important. In first approach we used an impregnation method described by Gourani et al. [9], but commercially made medical tablets degraded in impregnating solution. Later, we attempted to use Pasteur pipettes to gently, drop by drop, impregnate form both sides of a tablet. However, the results showed that ‘‘stock’’, baked-only tablets should also be a good choice as a passive monitoring absorption material. Further studies will show if impregnation with a reduced concentration of impregnating solution is viable. Another issue in the case of passive monitoring is the calibration of the monitoring system, where it is the goal to find the relationship between activity absorbed in the activated charcoal and the activity concentration in air. In a first approximation for our pilot study we have calculated the trapping efficiency as a ratio between the activity trapped in tablets and the activity in the air of the sampling site as determined by conventional sampling with HVS-30 station. However, this approach is only a rough estimation,
and tablets must be tested in controlled conditions. For
where k is the decay constant of 131I or 126I (depending on the isotope used for calibration). The issue with this is that there may be only small amounts of charcoal used in a single container, which may only provide a weak signal in small HPGe detectors. However, the application of ger- manium detectors with high efficiency or with a well geometry should give quite good results.

# Conclusions

This article describes results of pilot studies about the application of medical activated charcoal for passive monitoring of gaseous 131I activity concentration in nuclear medicine departments, where activity concentration of 131I can reach levels of few hundred Bq. Conventional air sampling methods are based on large and noisy air sam- plers, which are not suitable for routine measurements in hospitals. Our method is based on direct gamma-ray
131
future efficiency tests, we have decided to use 126I as a
measurements of
I trapped in charcoal tablets prepared
tracer, because it has a half-life almost two times longer than 131I (13.11 d) and is easy to produce on site with a cyclotron from non-enriched targets (TeO2) using the reaction natTe(p,xn)124,126I. Therefore transport of highly radioactive Na131I (commercially available) is not neces- sary. 222Rn trapping efficiency can also be used for cali- bration, where the standard equation used for calibration of activated charcoal detectors in a chamber is [2]:
CaV ¼ N=ðfafct kÞ	ð3Þ
where Ca (Bq m-3) is concentration inside glovebox dur- ing calibration, V (m3) is the chamber volume, N is the
count rate of the calibrated sample (cps), fa is the trapping efficiency (Bq m-3 cps-1), fc is the germanium spec- trometer efficiency for the gamma line (364 keV for 131I or 666 keV for 126I), t is the charcoal exposure time in the chamber (s) and k is the correction factor for radioactive decay during calibration and counting (s-1) calculated as:
(
in different ways using a HPGe detector, requiring no
chemical treatment or mixing with scintillating cocktail. Results showed, that 131I can be trapped in both, impreg- nated and non-impregnated tablets. Since measurements of 131I activity in air are necessary for the radioprotection of patients and personnel after precise calibration this method will become useful.

Acknowledgements This study was funded by the National Science Centre, Poland (Grant Number 2014/15/B/NZ7/00925).

Open Access This article is distributed under the terms of the Creative Commons Attribution 4.0 International License (http://creative commons.org/licenses/by/4.0/), which permits unrestricted use, dis- tribution, and reproduction in any medium, provided you give appropriate credit to the original author(s) and the source, provide a link to the Creative Commons license, and indicate if changes were made.

# References
k ¼ kt= 1 - eð-k tÞ
ð4Þ
North DJ, Shearer DR, Hennessey JV, Donovan GL (2001) Effective half-life of 131I in thyroid cancer patients. Health Phys 81:325–329

1. Jime`nez F, Deban L, Pardo R, Garcia-Talavera P (2012) Combi- nation of liquid scintillation counting and passive sampling strategy for the determination of 131I in air and application to estimate the inhalation dose to the staff of a nuclear medicine service. J Environ Sci Health 47:1843–1848
2. Mietelski JW, Grabowska S, Nowak T, Bogacz J, Gaca P, Bartyzel M, Budzanowski M (2005) Inhalation dose due to presence of 131I in air above septic tank system of an endocrinology hospital. Radiat Prot Dosim 117:395–401
3. Brudecki K, Kowalska A, Zagrodzki P, Szczodry A, Mro´z T, Janowski P, Mietelski JW (2017) Measurement of 131I in thyroid of nuclear medical Staff and internal dose assessment in a Polish nuclear medical hospital. Radiat Environ Biophys 56:19–26
4. Gru¨ndel M, Kopka B, Schulz R (2007) 131I exhalation by patients undergoing therapy of thyroid diseases. Radiat Prot Dosim 129:435–438
5. Brudecki K, Szczodry A, Mro´z T, Kowalska A, Mietelski JW (2018) Measurement of 131I activity in air indoor Polish nuclear medical hospital as a tool for an internal dose assessment. Radiat Environ Biophys 57:77–82
6. El-Azony KM, Suzuki K, Fukumura T, Szelecsenyi F, Kovacs Z (2008) Proton induced reactions on natural tellurium up to 63 MeV: data validation and investigation of possibility of 124I production. Radiochim Acta 96:736–769
7. Sing KS (1982) Reporting physiosorption data for gas/solid systems with special reference to the determination of surface area and porosity. Pure Appl Chem 54:2201–2218
8. Gourani M, Sadighzadeh A, Mizani F (2014) Effect of impreg- nating materials in activated carbon on iodine-131 (131I) removal efficiency. Radiat Prot Environ 37:179–183