# ANALYSIS OF PROMAZINE IN BIOLOGICAL FLUIDS BY HPLC AND MECC WITH SPECTROPHOTOMETRIC DETECTION IN TWO CASES OF FATAL POISONING
Katarzyna MADEJ1, Maria KAŁA2, Michał WOŹNIAKIEWICZ1
1. Faculty of Chemistry, Jagiellonian University, Cracow
2. Institute of Forensic Research, Cracow
ABSTRACT: Results of quantitative analyses of promazine in biological fluids
(blood and urine) by high performance liquid chromatography (HPLC) and micellar
electrokinetic capillary chromatography (MECC) were compared. Spectrophotomet-
ric detection was used for both methods. The examined samples originated from two
forensic cases of fatal poisoning with promazine. The obtained results were pre-
sented and summarised. The basic validation parameters of the two applied methods
were determined and compared.
KEY WORDS: Promazine; High performance liquid chromatography; Micellar
electrokinetic capillary chromatography; Biological fluids.
Problems of Forensic Sciences, vol. LVI, 2003, 15-23
Received 14 November 2003; accepted 18 December 2003
INTRODUCTION
Promazine (Figure 1) is a tricyclic antipsychotic compound which belongs
to the most commonly used phenothiazine drugs. Promazine in its hydro-
chloride form is prescribed mainly for the treatment of psychosis in older
people and is also used as an antiemetic medicine [1]. Therapeutic and toxic
(or lethal, L) concentrations of this drug in blood are 0.1-0.4 and above 1
Promazine has been determined alone
in various forms of pharmaceutical prepa-
rations or simultaneously with other
structurally related drugs by various ana-
lytical techniques such as: chromato-
graphic [2], electrochemical [3], spectro-
photometric [4] and capillary electropho-
retic [5] methods, and in biological fluids
as well.
(L > 5) Rg/ml, respectively.
CH3
'ch3
Fig. 1. Structure of promazine.

16
K. Madej, M. Kala, M. Wozniakiewicz
Capillary electrophoresis (CE) is a drugs analysis technique that has de-
veloped rapidly since the 1990’s. Two separation modes of CE: capillary zone
electrophoresis (CZE) and micellar electrokinetic capillary chromatography
(MECC) play a special role in this field. A variety of detection techniques
such as: spectrophotometric, amperometric, thermo-optical and mass spec-
trometry have been coupled with the above mentioned separation tech-
niques. However, to the knowledge of the authors, CE techniques have not
yet been applied to the analysis of blood samples and are not routinely used
in toxicological analysis.
The purpose of this work was to show that MECC may be used as a com-
plementary or alternative method to HPLC in the toxicological analysis of
promazine present in biological fluids. For this purpose, the MECC method
based on the use of sodium taurodexycholate as a surfactant [6] was applied
to the analysis of three samples of biological fluids originating from two fo-
rensic cases of fatal poisoning with promazine. Blood and urine samples
were simultaneously examined using HPLC and MECC methods, both with
spectrophotometric detection.
EXPERIMENTAL
# Materials
Drug standards of promazine and imipramine (used as internal stan-
dard, IS) were obtained from Sigma-Aldrich (St. Louis, USA). They were
added to drug free blood (control blood) purchased from the local blood bank
in Cracow and to drug free urine (control urine) obtained from a healthy per-
son with no history of medicine-taking.
Forensic post-mortem blood and urine samples were sent to the Institute
of Forensic Research in Cracow for toxicological examinations.
The mobile phase for HPLC used in gradient mode consisted of aceto-
nitrile (B) and water with phosphoric acid (pH = 2.3) (A). The separation
buffer for electrophoresis was prepared by the addition of 10 mM sodium salt
of taurodeoxycholic acid (STDC) (Sigma-Aldrich, St. Louis, USA) to 40 mM
sodium borate, pH = 9.5.
# Methods
HPLC and MECC were used as the analytical methods.
A series D-7000 liquid chromatograph equipped with an L-7450 diode ar-
ray detector (DAD), LaChrom System (Merck-Hitachi, Darmstadt, Ger-
many) and a Prince 550 air thermostated capillary electrophoresis system

Analysis of promazine in biological fluids
17
(Prince Technologies, Emmen, Holland) with a Lambda 1010 spectrophoto-
meter (Bischoff, Leonberg, Germany) as UV-VIS detector were used.
Samples were injected into the separation column of a liquid chromato-
graph using a Rheodyne injection valve (fitted with a 20 Rl loop) and into the
capillary using the hydrodynamic method (100 mbar pressure for 6 s).
Separation of analytes was conducted in a LiChroCART (125 x 4 mm)
column packed with octasilica LichroSpher RP Select B (5 Rm) (Merck,
Darmstadt, Germany) and in a bare fused silica capillary of 50 rm ID and
100/66 cm length at 30 kV, using HPLC and CE methods respectively. The
flow rate of the mobile phase for HPLC was 1 ml/min. The gradient was pro-
grammed as follows: 80% A and 20% B linearly decreased to 70% A and in-
creased to 30% B for 10 min, then returned to the initial conditions,
e. 80% A and 20% B for 5 min.
Measurements of analytes were performed at 254 nm wavelength.
# Extraction procedure
1 ml of control blood (or control urine) was spiked with standard solutions
of promazine and imipramine (IS) at appropriate concentrations. The sam-
ples were adjusted to pH = 12 using 0.6 M NaOH (3 ml) and extracted with
5 ml hexane-isoamyl alcohol (99:1, v/v). Then 4 ml of organic layers were sep-
arated, evaporated to c.a. 0.5 ml under nitrogen and re-extracted to 100 Rl of
01% phosphoric acid. The aqueous layer was divided into two equal parts.
One part was analysed by HPLC and the second one was evaporated to dry-
ness, dissolved in 50 Rl of 0.01 M STDC solution and examined by MECC.
The autopsy samples of blood and urine were spiked with 5 Rg of
imipramine and treated in the above mentioned way.
RESULTS
The forensic samples from two cases of fatal poisonings by promazine
were analysed simultaneously by two methods: HPLC and MECC, in the
conditions summarised in Table I. The quantitative results ofpromazine de-
terminations achieved by these two methods are presented in Table II.
The MECC method was developed and validated for determination of
promazine in relation to imipramine as the internal standard, using control
blood (or control urine) spiked with the standard drug solutions. The follow-
ing concentrations of promazine were obtained: 0.313; 0.625; 1.25; 2.5;
5.0 Rg/ml for blood and 1.25; 2.5; 5.0; 7.5 Rg/ml for urine. Imipramine was
added to the analysed samples at a concentration of 5 Rg/ml.

Analysis of promazine in biological fluids
18
TABLE I. ANALYTICAL CONDITIONS FOR HPLC AND MECC METHODS
Conditions	HPLC	MECC
Column/ / capillary	RP-Select B 125 x 4 mm (5 Rm)	Bare fused silica capillary L = 100/66 cm, ID = 50 Rm
Mobile phase/ separation buffer	Acetonitrile + water with phosphoric acid (pH = 2.3) Gradient mode Flow = 1ml/min	40 mM sodium borate + 10 mM STDC, pH = 9.5
Detection / method	Spectrophotometric X = 254 nm	Spectrophotometric X = 254 nm
Temperature	25oC	25oC
Voltage	-	30 kV

TABLE II. CONCENTRATIONS OF PROMAZINE IN BLOOD AND URINE SAMPLES
DETERMINED BY HPLC AND MECC METHODS
Material	Case / number	Concentration of promazine [rg/ml]
		HPLC	MECC
Blood	I	4.59	4.72
	II	1.31	—
Urine	II	5.81	6.09

TABLE III. VALIDATION PARAMETERS OF HPLC AND MECC METHODS
Parameter	HPLC	MECC
Precision of identification pa-	0.56	0.43
rameter, Rt/Mt [RSD%]	(n* =11)	(n = 9)
Precision of quantification	10.00	9.88
parameter, Ap/Ai [RSD%]	(n = 4)	(n = 9)
Limit of quantification (LOQ) [r g/ml]	0.125	0.63
Limit of linearity (LOL)	0.125-2.0	1.25-5.0
[r g/ml]	r2 > 0.998	r2 > 0.996

*Number of repeated measurements.

Analysis of promazine in biological fluids
19
The validation parameters, such as precision of identification parameter
(repeatability of relative times of retention, Rt and migration, Mt), precision
of quantification parameter (repeatability of the peak area ratios of proma-
zine and imipramine, Ap/Ai), limit of quantification (LOQ), limit of linearity
(LOL) and correlation coefficient (r2) were determined and compared for both
methods (Table III). The amount of blood sample from case II was not suffi-
cient for promazine to be successfully determined by the MECC method.
The obtained chromatograms and electrophoregrams were compared and
presented in Figures 2 A and B.
Retention time [min]
Fig. 2. Chromatograms and electrophoregrams of control blood sample spiked with
standard drug mixtures of promazine (added at concentrations of: 0; 0.625; 2.5 and
5.0 pg/ml) and imipramine (added at concentration of 5.0 pg/ml), examined by HPLC
(A) and MECC (B) methods. EOF - electroosmotic flow, IS - imipramine, Pro -
promazine.

20
K. Madej, M. Kala, M. Wozniakiewicz
SUMMARY AND CONCLUSIONS
The concentrations of promazine in blood and urine samples determined
by two methods: HPLC and MECC, were comparable.
The LOQ for promazine in blood achieved by the MECC method was
higher than that obtained by the HPLC method. However, the sensitivity of
the MECC method allows this method to be used for biological fluids exami-
nation in cases of poisoning with promazine because the toxic concentration
of promazine in blood is above 1 |4g/ml. The repeatability of identification
and quantitative parameters of both methods was comparable (Table I).
The linearity of both methods was examined in the concentration range
of 0.125-2.0 Rg/mlfor HPLC and 1.25-5 Rg/ml for MECC, and the obtained
correlation coefficients (r2) were greater than 0.996.
Taking into account the above mentioned facts we can conclude that the
MECC technique may be used in forensic toxicological analysis of pro-
mazine, and can be considered a complementary method to HPLC. It is quite
probable that other drugs that are structurally-related to promazine can
also be determined by the elaborated MECC method.
References:
1. Anxiolytic, sedatives, hypnotics and antipsychotics, [in:] Martindale. The com-
   plete drug reference, Parfitt K. [ed.], Pharmaceutical Press, London 1999.
2. Madej K., Parczewski A., Kala M., HPLC/DAD screening method for
   selected psychotropic drugs in blood, Toxicology Mechanisms and Methods
   2003, vol. 13, pp. 121-127.
3. Belal F., El-Ashry S. M., Shehata I. M. [et al.], Differential-pulse
   polarographic determination of some n-substituted phenothiazine derivatives
   in dosage forms and urine through treatment with nitrous acid, Mikrochimica
   Acta 2000, vol. 135, pp. 147-154.
4. El-MaaboudA. I. M., Utility of diphenylamine and N-bromosuccinimide
   for colorimetric determination of certain phenothiazine drugs, Talanta 1997,
   vol. 44, pp. 1173-1182.
5. Muijselaar P. G. H. M., Claessens H. A., Cramers C. A., Determi-
   nation of structurally related phenothiazines by capillary zone electrophoresis
   and micellar electrokinetic chromatography, Journal of Chromatography A,
   1996, vol. 735, pp. 395-402.
6. Aumatell A., Wells R. J., Determination of cardiac antiarrhythmic,
   tricyclic antipsychotics and antidepressants in human and animal urine by
   micellar electrokinetic capillary chromatography using a bile salt, Journal of
   Chromatography B, 1995, vol. 669, pp. 331-344.

ANALIZA PROMAZYNY W P£YNACH USTROJOWYCH
METODAMI HPLC I MECC ZE SPEKTROFOTOMETRYCZN¥
DETEKCJ¥ W DWÓCH PRZYPADKACH ŚMIERTELNEGO
ZATRUCIA
Katarzyna MADEJ, Maria KAŁA, Michał WO NIAKIEWICZ
WSTĘP
Promazyna (rycina 1) jest trójpierścieniowym związkiem przeciwpsychotycz-
nym, który należy do najczęściej stosowanych leków fenotiazynowych. Promazyna
w postaci jej chlorowodorku przepisywana jest głównie w terapii psychoz starszych
ludzi oraz jest stosowana również jako lek przeciwwymiotny [1]. Stężenia terapeu-
tyczne i toksyczne (lub spotykane w przypadkach śmiertelnych, L) tego leku we krwi
wynoszą odpowiednio 0,1-0,4 i powyżej 1(L > 5) pg/ml.
Promazyna była oznaczana, osobno lub jednocześnie, z innymi podobnymi struk-
turalnie lekami różnymi technikami analitycznymi, takimi jak metody chromato-
graficzne [2], elektrochemiczne [3], spektrofotometryczne [4] i elektroforezy kapilar-
nej [5] w różnych postaciach leku, jak również w płynach biologicznych.
Elektroforeza kapilarna (CE) jest - od ostatniej dekady ubiegłego wieku - dyna-
micznie rozwijającą się techniką analizy leków. W tej dziedzinie specjalnąrolę odgry-
wają dwa jej separacyjne warianty: strefowa elektroforeza kapilarna (CZE) i mice-
larna elektrokinetyczna chromatografia kapilarna (MECC). Wiele różnorodnych
technik detekcji, takich jak spektrofotometryczna, amperometryczna, termooptyczna
i spektrometrii mas, było łączonych z wyżej wspomnianymi technikami separacyjny-
mi. Aczkolwiek, zgodnie z wiedzą autorów, metody CE nie były jeszcze stosowane
w analizie prób krwi i nie sąrutynowo wykorzystywane w analizie toksykologicznej.
Celem pracy było wykazanie, że MECC może być stosowana jako uzupełniająca
lub konkurencyjna metoda w stosunku do HPLC w analizie toksykologicznej proma-
zyny obecnej w płynach ustrojowych. W tym celu metoda MECC, oparta na zastoso-
waniu soli sodowej kwasu taurodeoksycholinowego jako substancji powierzchniowo
czynnej [6], została zastosowana w analizie trzech prób płynów biologicznych po-
chodzących z dwóch ekspertyz, dotyczących zatrucia śmiertelnego promazyną. Próby
krwi i moczu poddano jednoczesnemu badaniu, stosując metody HPLC i MECC, obie
z detekcją spektrofotometryczną.
CZĘŚĆ EKSPERYMENTALNA
## Materiały
Substancje wzorcowe leków: promazyny i imipraminy (stosowane jako standard
wewnętrzny, IS) otrzymano z firmy Sigma-Aldrich (St. Louis, Stany Zjednoczone).
Były one dodawane do krwi wolnej od leków (krew kontrolna) otrzymanej ze stacji

22
K. Madej, M. Kała, M. Woźniakiewicz
krwiodawstwa w Krakowie oraz do moczu wolnego od leków (mocz kontrolny) pobra-
nego od zdrowej osoby nie przyjmującej leków.
Próby sekcyjnej krwi i moczu zostały przysłane do Instytutu Ekspertyz Sądowych
w Krakowie celem przeprowadzenia badań toksykologicznych.
Faza ruchoma do HPLC, stosowana w trybie gradientu składu, składała się z ace-
tonitrylu (B) oraz wody z kwasem fosforowym (pH = 2,3) (A). Bufor separacyjny do
elektroforezy był przygotowany przez dodanie 10 mM soli sodowej kwasu taurodeok-
sycholinowego (STDC) (producent: firma Sigma-Aldrich, St Louis, Stany Zjedno-
czone) do 40 mM boranu sodowego, pH = 9,5.
## Metody
Techniki HPLC i MECC zostały użyte jako metody analityczne. Zatosowano chro-
matograf cieczowy LaChrom System, model D-7000, wyposażony w detektor szeregu
diod (DAD) L-7450 (Merck-Hitachi, Darmstadt, Niemcy) oraz zestaw do elektrofore-
zy kapilarnej Prince 550 termostatowany powietrzem (Prince Technologies, Emmen,
Holandia) ze spektrofotometrem Lambda 1010 (Bischoff, Leonberg, Niemcy) jako de-
tektorem UV-VIS.
Próbki wstrzykiwano na kolumnę separacyjną chromatografu cieczowego, uży-
wając zaworu wstrzykowego Rheodyne z pętlą 20 |il, a do kapilary wykorzystano me-
todę hydrodynamiczną (ciśnienie 100 mbar przez 6 s).
Rozdział analitów prowadzono w kolumnie LiChroCART (125 x 4 mm) z wy-
pełnieniem typu oktasilica LichroSpher RP Select B (5 |im) (Merck, Darmstadt,
Niemcy) i w kapilarze kwarcowej o średnicy wewnętrznej 50 m i długości 100/66 cm
z przyłożonym napięciem 30 kV, stosując - odpowiednio - metody HPLC i CE.
Przepływ fazy ruchomej w metodzie HPLC wynosił 1 ml/min. Program gradientu
składu fazy był następujący: 80%Ai20%B, liniowo obniżał się do 70% A i wzrastał do
30% B przez 10 min, a następnie wracał do warunków wyjściowych, tj. 80% A i 20% B
przez 5 min. Pomiary analitów przeprowadzano przy długości fali 254 nm.
## Procedura ekstrakcji
Do 1 ml krwi kontrolnej (lub moczu kontrolnego) dodawano roztwory wzorcowe
promazyny i imipraminy (IS) w odpowiednich stężeniach. Próby były doprowadzane
do pH = 12 za pomocą0,6 M NaOH (3 ml) i ekstrahowane 5 ml mieszaniny heksan-al-
kohol izoamylowy (99:1, v/v). Następnie oddzielano 4 ml warstwy organicznej, odpa-
rowywano do ok. 0,5 ml pod azotem i reekstrahowano do 100 |il 0,01% kwasu fosforo-
wego. Warstwa wodna była dzielona na dwie części. Jedna część była analizowana
metodą HPLC, a druga odparowywana do sucha, rozpuszczona w 50 |il 0,01 M roz-
tworu STDC i badana metoda MECC.
Do próbek sekcyjnych krwi i moczu dodawano 5 g imipraminy i traktowano je
w wyżej podany sposób.
WYNIKI
Próbki przesłane do badań dla celów sądowych z dwóch przypadków śmiertelnego
zatrucia promazyną analizowano jednocześnie dwiema metodami: HPLC i MECC
w warunkach zestawionych w tabeli I.

Analiza promazyny w płynach ustrojowych
23
Wyniki oznaczeń promazyny uzyskane tymi metodami podano w tabeli II.
Opracowano i poddano walidacji metodę MECC do oznaczania promazyny wzglę-
dem imipraminy jako standardu wewnętrznego, stosując kontrolną krew (lub kon-
trolny mocz) z dodatkiem roztworów wzorcowych. Uzyskano następujące stężenia
promazyny we krwi: 0,313; 0,625; 1,25; 2,5; 5,0 |ig/ml i w moczu: 1,25; 2,5; 5,0;
7,5 g/ml. Imipramina byla dodawana do badanych próbek w stężeniu 5 |ig/ml.
Parametry walidacyjne, takie jak precyzja parametru identyfikacyjnego (pow-
tarzalność względnych czasów retencji, Rt i migracji, Mt), precyzja parametru ilościo-
wego (powtarzalność stosunków pól powierzchni pików promazyny i imipraminy,
Ap /Ai), granica oznaczalności (LOQ), zakres liniowości (LOL) i współczynnik kore-
lacji (r2) były wyznaczane i porównywane dla obu metod (tabela 3). Ilość próby krwi
(z przypadku II) była niewystarczająca, aby oznaczyć w niej promazynę metodą
MECC. Otrzymane chromatogramy i elektroforegramy zostały porównane i przed-
stawione na rycinach 2 i 3.
PODSUMOWANIE I WNIOSKI
Stężenia promazyny w próbach krwi i moczu oznaczone dwiema metodami:
HPLC i MECC były porównywalne.
LOQ dla promazyny we krwi uzyskana metodąMECC była wyższa niż LOQ uzys-
kana metodą HPLC, aczkolwiek czułość metody MECC pozwala zastosować jądo ba-
dania płynów biologicznych w przypadkach zatruć promazyną, ponieważ stężenie
toksyczne promazyny we krwi wynosi ponad 1 g/ml. Powtarzalność parametrów
identyfikacyjnego i ilościowego dla obu metod była porównywalna (tabela I). Zakresy
liniowości obu metod badania w przedziałach stężeń: 0,125-2,0 |ig/ml dla HPLC
i 1,25-5,0 |ig/ml dla MECC, a współczynniki korelacji (r2) były większe niż 0,996.
Biorąc pod uwagę wyżej wymienione fakty, można wyciągnąć wniosek, że metoda
MECC może być stosowana w analizie toksykologicznej promazyny i może być uzna-
na za uzupełniającą metodę dla HPLC. Jest całkiem prawdopodobne, że inne leki,
strukturalnie podobne do promazyny, mogąbyć również oznaczane opracowaną me-
todą MECC.