Gineko/ Pol. 2009, 80, 256-262
PRACE	ORYGINALNE
ginekologia
Rapid HPLC method for the determination
of vitamin A and E and cotinine concentration in human serum in women with CIN
and cervical cancer
Szybka metoda oceny stężenia witaminy A, witaminy E oraz kotyniny
w surowicy krwi kobiet z sródnabfonkową neoplazją (CIN)i rakiem szyjki macicy
i"5'
Bystrowska Beata1, Gomółka Ewa4, Szczudrawa Andrzej2, Brandys Jerzy1, Pawlik Michal2, Milewicz Tornasz2, Dulinska-Litewka Joanna3, Jach Robert2
1 Department of Toxicology, Collegium Medicum Jagiellonian University, Krakow, Poland
2 Department of Gynecology and Obstetrics, Collegium Medicum, Jagiellonian University, Krakow, Poland,
3 Department of Medical Biochemistry, Collegium Medicum, Jagiellonian University, Krakow, Poland
4 Department of Clinical and Environmental Toxicology, Collegium Medicum, Jagiellonian University, Krakow, Poland
Abstract
Objective: The aim of this study was to elaborate on the analytical method for quantitative determination of retinal and a-tocopherol in serum of women diagnosed with C/N and cervical cancer. The basic problem in the analysis of the vitamins content in biological material is their low physiological concentration level and instability. Liquid chro­ matography with diode array detector (DAD) wa s applied.
Material and methods: The material consisted of serum and urine collected from 12 women diagnosed with cervical intraepithelial neoplasia (CIN) and 16 diagnosed with cervical cancer.
The method was evaluated for the following parameters: linearity, recovery, sensitivity, precision, accuracy, selec­ tivity, stability, limit of quantification (LOQ) and limit of detection (LOO).
Results: Results showed good linearity (r '?.0, 99) in the range 0, 1 µg!ml-10mg!ml for retinal and 0,25µg!ml-15µg!ml for a-tocopherol. The Lower Limit of Detection was 0, 1 Sµg!ml for vitamin E and 0,05µg!ml for vitamin A. The with­ in-run R.5. Ds were below 5,2% at all concentration levels and the between-run R.5. Ds were below 10,0% at all con­ centration levels.
Conclusions: The advantage of this method is that it measures both compounds in a more rapid, reproducible and accurate manner when compared to the previous HPLC studies. The compounds (vitamin A and E and internal stan­ dards) are measured in the same sample at the same time.
Quantitative determination of cotinine may reveal active smokers and subjects exposed to environmental tobacco smoke, which is independent measurable carcinogenetic co-factor.
The following study is a part of a project determining non-viral causative agents in cervical carcinogenesis.
i"5'
Key words: vitamin A / vitamin E / retinol / alpha-Tocopherol /
/ high pressure liquid chromatography - HPLC / cotinine /
Correspondence to:
Robert Jach
Dept. of Gynecology, Obstetrics and Oncology, Jagiellonian University, Medical College; 23 Kopernika str., 31-501 Krakow, Poland,
Tel +48 12 4248560,
Fax +48 12 4248584,
e-mail: jach@cm-uj.krakow.pl
2
"'
1,c
0u
.c"'
f-
I
Otrzymano 30.12.2008
Zaakceptowano do druku 23.02.2009
256
Nr 4/2009

Gineko/ Pol. 2009, 80, 256-262
PRACE	ORYGINALNE
ginekologia
Bystrowska B, et aI.
Streszczenie
Cel pracy: Celem niniejszej pracy jest ocena przydatnosci analitycznej metody oznaczania ilosciowego retinolu i a-tokoferolu w surowicy kobiet z rozpoznanq sr6dnablonkowq neoplazjq (CIN) i rakiem szyjki macicy. Gl6wnym problemem w analizowaniu zawartosci witamin w materiale biologicznym jest ich niskie fizjologiczne st(;!zenie i niestabilnosc. W niniejszym badaniu zastosowano metodf! cieklej chromatografii z detektorem diodowym (DA D). Material i metodyka: Material badawczy stanowilo osocze i mocz, 12 kobiet z rozpoznanq sr6dnablonkowq neo­ plazjq szyjki macicy (CIN) i 16 z rakiem szyjki macicy.
Metodf! oceniono pod kqtem nast(;!pujqcych parametr6w: liniowosc, powtarzalnosc, czulosc, precyzja, dokladnosc, selektywnosc, stabilnosc, granica kwantyfikacji (limit of quantification - LOQ) i detekcji (limit of detection - LOO). Wyniki: Uzyskano wysokq liniowosc (r 0,99) w zasif;!gu 0, 1 µg!ml-10mg!ml dla retinolu i 0,25µg!ml-15µg!ml dla a-tokoferolu. Daina granica wykrywalnosci wyniosla 0, 1 Sµg!ml dla witaminy E i 0,05µg!ml dla witaminy A. Zasim metody R.5.Ds wyni6sl 5,2% we wszystkich poziomach st(;!zen a interwalowy zasi(;!g R.5.Ds ponizej 10,0% we wszystkich poziomach st(;!zen.
Wnioski: Zaletq zbadanej metody jest fakt, iz mozna przy jej pomocy uzyskac wyniki w szybszy, bardziej pow­ tarzalny i dokladniejszy spos6b, niz w opisywanych w literaturze badaniach z uzyciem HPLC. Dodatkowq zaletq stosowanej metody jest mozliwosc pomiaru z tej same} pr6bki witamin A i E.
llosciowa ocena st(;!zenia kotyniny maze wskazac pacjentki, bi;!dqce aktywnymi palaczkami lub narazonymi na pale­ nie bierne. W aspekcie karcinogenezy w obri;!bie szyjki macicy maze bye to cenna wiadomosc dotyczqca istotnego a cz(;!sto zatajanego kofaktora nowotworzenia
Niniejsze badanie jest CZf!SCiq projektu oceniajqcego pozawirusowe czynniki onkogenne w procesie karcinogenezy w obri;!bie szyjki macicy.
0u
.c""
f-
Sfowa kluczowe: witamina A / witamina E / retinol / alfa-tokoferol /
/ wysokosprawna chromatografia cieczowa - HPLC I kotynina /
Vitamin A
Vitamin A, the most important form of which is retinal, is produced in the intestinal mucosa from provitamin A (known as beta-carotene) through an oxidative cleavage reaction.
Retinal is oxidized in vivo to aldehyde retinal (the form used in visual pigment) and retinoid acid. Both of them are known as retinoids and are active forms of vitamin A. Vita­ min A, produced and absorbed in the proximal part of the intestine, is transported in chylomicrons via enterohepatic cir­ culation to the liver for esterification and storage. When retinal esters in the liver are mobilized, released retinal is then bound to a-1-globulin and albumin. Most of the vitamin A reserves in a body are stored in the liver. Important dietary sources of vitamin A are animal-derived (fish liver, cod-liver oil, milk, eggs, butter).
It should be underlined that human breast milk contains twice as much vitamin A as whole milk of a cow. In humans, beta-carotene remains the basic source of vitamin A. Yellow and leafy green vegetables such as carrots provide large amounts of carotenoids, many of which are provitamins that can be metabolized in the active form in vivo. Thus, the plants neither produce nor contain vitamin A, they merely supply provitamins for its synthesis. The rate of vitamin A produc­ tion from a-carotene is influenced by the actual demands of the body. Therefore, its excessive dietary intake can lead to vitamin A toxicity.
Molecular mechanism of retinoids' influence is based on their binding to intracellular receptor, impact on gene expres­ sion and, consequently, on protein synthesis. This mode of action is responsible for synthesis of a dozen proteins, amongst them: EGF receptor, kreatine, cell-growth control polyamide.
Introduction
Vitamins A and E, their functions and deficiency syndromes
Vitamins are small chemical molecules, essential for opti­ mal health. Their biological role, as coenzymes, is attributed to the control of catalytical reactions. They have influence on numerous hormone-dependent processes.
In small amounts, vitamins are essential for health. How­ ever, neither animals nor humans are able to produce enough of them and therefore vitamins must be supplemented. Thus, the main source of vitamins or their precursors is a balanced diet. It should be emphasized that a balanced diet consists of sufficient amounts of vitamins.
Deficiency of vitamins may be dietary in orgin, or may occur due to disturbances in intestinal absorption, presence of antivitamins or liver damage. In case of certain vitamins defi­ ciency is caused by their impaired synthesis by intestinal flora or prolonged treatment with antibiotics. Due to their solubili­ ty, vitamins are divided into fat- (A, D, E, K), and water-solu­ ble (group B, C, PP) which act as coenzymes. [l]
Monitoring of the status of vitamins A and Eis important to detect inadequate dietary intake of these vitamins or a pos­ sible deficiency of either retinal-binding protein, protecting retinal from degradation pathways or of a-tocopherol trans­ fer protein (a-TTP), which incorporates a-tocopherol into very-low-density lipoproteins in the liver. [2, 3, 4, 5, 6, 7, 8, 9,
10].
In this paper we discuss a procedure for simultaneous detection of vitamins A and E in serum using HPLC with DAD method.
2
0
u
."c '
f-
© 2009 Polskie Towarzystwo Ginekologiczne
257

PRACE	ORYGINALNE
ginekologia
Gineko/ Pol. 2009, 80, 256-262
Rapid HPLC method for the determination of vitamin A and E and cotinine concentration ..
The most important function in visual process involves predominantly one form of vitamin A, namely retinal. Retinal enters into the composition of rhodopsin, the most light-sen­ sitive pigment and thus highly important in perception of reduced light. Rhodopsin, under the influence of light, under­ goes a sequence of changes to ultimately disintegrate to trans­ retinal and opsin. In the process, a nerve impulse is generated which assures good vision. Vitamin A deficiency leads to dep­ rivation of rhodopsin, night blindness and xerophthalmia.
Vitamin A prevents the keratinization of mucous-secreting epithelium of gastrointestinal and respiratory tract. It acceler­ ates cell division and growth of fibroblasts, thus acting as the growth factor, predominantly in young organisms. Vitamin A plays an important role in spermatogenesis and embryogene­ sis. There is evidence that vitamin A may inhibit the growth of experimental breast cancer.
Recommended daily intake of vitamin A is 500IU. The demand for vitamin A increases significantly during pregnan­ cy and breast feeding.
This essential nutrient belongs to the group of antioxi­ dants whose function is scavenging free radicals formed in redox reactions throughout the body. It plays a role in termi­ nation of oxidative-generated lipid peroxidation chain reac­ tions, particularly in cellular membranes that are rich in polyunsaturated lipids. Vitamin E is a component of LDLs and protects them from oxidative damage.
Hypovitaminosis E is associated with inadequate bile release and small intestinal disease. Vitamin E deficiency con­ tributes to neurologic manifestations, peripheral and central neuropathies. In case of adults, unlike in children, neu­ ropathies are not a feature of the deficiency state and may be present after many years of inadequate intake. Vitamin E defi­ ciency may affect mature red cells by shortening their life span. Fertility disorders may be seen in women (impaired implanta­ tion), as well as men (reduced spermatogenesis). It has been suggested that vitamin E may have protective effects against atherosclerosis and cancer, the two most common causes of death in developed countries. It is prescribed in treatment of anemia, blood vessel spasm disorders, collagenoses and ather­ osclerosis. Vitamin E can be used in recurrent abortions, oligospermia and muscular atrophies. However, therapeutic effect of vitamin E has not yet been conclusively stated and defined.
Daily requirement of vitamin Eis 10-3-mg.
0u
.c""
f-
Vitamin E
Vitamin Eis widely distributed in nature, particularly large amounts may be found in vegetable oils (soy), vegetables, dairy products and nuts. A group of eight closely related com­ pounds, known as tocopherols, exhibit vitamin E activity, but a tocopherol is the most active one. It constitutes (up to) 90% of all body tocopherols.
As it is currently known, tolopherols are exclusively pro­ duced by plants. The absorption of vitamin E, as of all fat-sol­ uble vitamins, takes place in ileum, requires normal biliary tract and pancreatic function, and depends on diet. If the intestinal absorption is impaired, the intramuscular oil solu­ tions of vitamin E may be administered. Vitamin E is excret­
Cotinine
Nicotine is a pyridine alkaloid that constitutes 85 to 90% of all alkaloids in tobacco leaves. Its main metabolite (80%) is cotinine; the metabolism is shown in Figure 1.
The sources of nicotine and cotinine in biological fluids are tobacco smoking and medicines containing nicotine (chewing-gum, plasters, inhalers) [11, 12].
Cotinine - the main nicotine metabolite - is eliminated from the body for a longer period of time than nicotine.
ed with bile, as glucuronide, pheronic acid.
and by the kidneys, as toco­
2
8
."c '
Figure 1. Nicotine metabolism [12,14].
f-
Nr 4/2009
258

"0
t.lcolineN□--

D	D	o	J?; o OH
l J	c ,	lJ	O	CH,
N	o?N	,
COll	N
Tr	--OH-col.irme

d? 1	N'	N	c;<I"" l?"	0	□
H	cl?	()mineN- -□	¾.	I	di,	;	Of-I
HO
N	,:Y	N	"	17·	OH
Nornicotm	rnI ,	O	IN/ o	O	CO Hi
I	2
N	CH,
0	"
OH
H
0
OH	OH
H

OH

Gineko/ Pol. 2009, 80, 256-262
PRACE	ORYGINALNE
ginekologia
Bystrowska
B, et aI.
Its biological halftime is 17 hours, clearance 0,04 1/min, distribution volume 55 1 (0,9 1/kg) [13]. The main way of coti­ nine and its metabolites elimination is via urine. Cotinine is one of the best and most useful biomarkers of tobacco smoke exposition [14, 15, 16].
The urine cotinine concentration reflects exposition to nicotine and tobacco smoke in the last 2-3 days. Research showed that urine cotinine concentrations may be useful in distinguishing active smokers from the ones exposed to envi­ ronmental tobacco smoke [15]. The cotinine concentration in active smoker urine is higher than lO0ng/ml and can be high­ er than 3000ng/ml. The cotinine concentration in the urine of a no-smoker exposed to environmental tobacco smoke one ranges from 25 to lO0ng/ml [15, 16, 17].
heather Lab-Line, ultrasonic cleaner Sonic-2, vacuum appara­ tus Alltech, vortex TH-3S TechnoCartel, pH meter MAT1202-SM.
The resolution was performed on Supelcosil LC-8 column (25cm, 4.6mm ID). The mobile phase was deionized water 88%, acetonitrile 12 %, sodium octasulphoniane l.lg/1, dicali­ um phosphate 5.95 g/1; pH was adjusted to 4.7 with the use of phosphoric acid. The HPLC analysis was isocratic with con­ stant flow rate lml/min. The Rheodyne injector loop was 100µ1. The detector wavelength was 254nm.
2
""
0u
."c '
f-
Standards and samples preparation
The blood was extracted and left to spontaneous coagula­ tion. Then the samples were centrifuged at 300 r.p.m. for 15 minutes, becoming excluded haemolysed samples. The clean serum was frozen and stored at the temperature of about - 40°C. "Stock solutions" of retinal, retinal acetate, a-toco­ pherol and a-tocopherol acetate were prepared first in the mixture of hexane/dichloromethane (50/50 v/v) containing ascorbic acid (0,025%) to protect vitamins against oxidation.
After this, standard solution was diluted in a mixture of dichloromethane/methanol (50/50 v/v) and, finally, in methanol (100%) and stored at -22°C, without the access of light. The QC samples were prepared in human plasma and stored in the same conditions as the analytical samples. The calibration samples were prepared immediately before the analysis.
The preparation of samples was done following the proce­ dure described by M.A. Rodriguez-Delgado at al. [13]. 100ml retinal acetate (internal standard I) and 100ml a-tocopherol acetate (internal standard 11) and 200ml of ethanol were added at 200ml volume of serum.
The mixture was shaken for 5 minutes. Subsequently, 200ml water and 800ml of hexane was added to the mixture and after blending for 15 minutes it was centrifuged for 5 min­ utes at 4000 r.p.m., and next 600ml of the upper organic layer was extracted. The organic phase was evaporated to dryness
Material and methods
28 patients, aged 19 to 67 years, diagnosed with pre-inva­ sive or invasive cervical lesion were included in the study in accordance with ethical approval regulation by Local Research Ethic Committee of Jagiellonian University in Krakow. The basis for the diagnosis were: punch or cone biop­ sy specimens obtained from consecutive pre-invasive CINl (n=3), CIN2 (n=3), and CIN3 (n=3) lesions at the Depart­ ment of Gynecology and Obstetrics, Jagiellonian University Medical College, Krakow between 2004 and 2006. Addition­ ally, specimen of primary squamous cell carcinoma of the cervix CaCx (n=l 6) were included into the study from patients who had undergone the radical hysterectomy with pelvic
lymph node dissection. The material for this pilot included 20ml of blood and 100ml of morning urine.
study
Chemicals and reagents
Retinal, a-tocopherol, retinal acetate, a-tocopherol acetate and ascorbinic acid in substance were provided from Sigma-Aldrich. Methanol, acetonitryl, hexane, water and dichloromethane used in this determination were of analytical grade (Merck). Ethanol was from PoCH (Gliwice). Chemicals used for cotinine determinations were: methanol HPLC Grade (Merck), octanesulfonic acid sodium salt monohydrate (Fluka), dichloromethane HPLC Grade, acetonitrile HPLC Grade, phosphoric acid, K2HP04, 2-propanol, cotinine stan­ dard (Sigma), NaOH, HCl (POCh), deionized water.
and
the
residue
was
dissolved	in
100ml
of
methanol/ethanol/hexane (88:10:2, v/v/v).
All samples were run in duplicate, one with and the other without internal standards.
At 20ml volume of this solution was injected directly into the batcher. The chromatographic separation took 18 minutes.
Chromatographic conditions
The chromatographic separations and quantitative deter­ mination were performed on a high performance liquid chro­ matograph Agilent 1100 series equipped with a pump, degasser, autosampler and DAD detector. A Supelco analyti­ cal column (Discovery Cl 8 column, 5mm particle size, 250x4.6mm) as the stationary phase was used.
Methanol was used as the mobile phase. The flow rate was
1.0 ml/min. Detection was monitored at two different wave­ lengths: 1=292nm for tocopherols and 1=325nm for retinols determination.
Cotinine was determined by high performance liquid chro­ matography (HPLC) with UV detector (254nm).
Cotinine extraction procedure
The liquid-liquid extraction was used for the preparation of the samples. 0.5ml of lM NaOH and 50µ1 of norephedrine (internal standard) was added to 5 ml of urine and mixed. The samples were extracted with 5 ml mixture of dichloromethane: 2-propanol (9:1) (v/v) for 15 minutes. The organic phase was collected to cone tubes and evaporated in the air stream in 40°C. The residues after evaporation were dissolved in 400ml of mobile phase and injected to HPLC.
Calibration curves
For calibration curves the stock solutions were diluted with methanol to obtain a mixture of working standards. Water containing measured amounts of retinal and a-toco­ pherol was used for the analysis.
2
"'
u
The equipment included: HPLC Merck Hitachi
Pump
0
L6200A, Detector Spectra 100 Thermo Separation Products, 4880 integration system, centrifuge MPW 220, multi-block
."c '
f-
© 2009 Polskie Towarzystwo Ginekologiczne
259

PRACE	ORYGINALNE
ginekologia
Gineko/ Pol. 2009, 80, 256-262
Rapid HPLC method for the determination of vitamin A and E and cotinine concentration ..
The samples of water were spiked with a-tocopherol to the following concentration: blank, 0.15; 0.25; 0.75; 1.5; 5.0; 10.0; 15,.0; 20.0mg/ml and with retinal to the following concentra­ tion: blank, 0.05; 0.10; 0.25; 0.5; 1.0; 2.5; 5.0; 10.0mg/ml.
Retinal acetate and a-tocopherol acetate were used as the internal standards. The concentration of retinal was 10µg/ml and a-tocopherol acetate was 100µg/ml. The blank samples were water and serum. Calibration for cotinine was made using blank urine spiked with cotinine to obtain standards: 10, 50,100,200,500, l000ng/ml.
These calibration samples were analysed following the procedure described for sample preparation (2.3 section). The number of QC's per batch was six (three concentrations level in duplicate). They were assessed according to 4-6-20 rule.
0u
.c""
f-
Results
Specificity / selectivity
The specificity/selectivity of the method can be illustrated by comparing the chromatograms obtained after analysis of a test solution of the pure compound with the chromatograms of the independent blank serum. No interfering peaks can be seen if blank serum is used. The peaks of the retinal, retinal acetate (Figure 2), a-tocopherol and a-tocopherol acetate (Figure 3) were well-resolved and showed no interferences with endogenous or exogenous materials.
The retention time was: 4.54 minutes for retinal; 5.46 min­ utes for retinal acetate, 9.4 minutes for a-tocopherol and 11.35 minutes for a-tocopherol acetate.
Figure 2. Chromatogram from DAD detector obtained at 1=325nm. It shows the peaks of the retinal and retinal acetate.
Sensitivity
The lower detection limit depends on several parameters such as the mobile phase and the DAD monitor used. Under the experimental conditions described, the LOD (defined as three times the baseline noise) was approximately 0.15µg/ml for vitamin E and 0.05µg/ml for vitamin A. The lowest con­ centration of the calibration graph for vitamin E was 100ng/ml and for the vitamin A was 250ng/ml and it was the Lower Limit of Quantification.
For cotinine the limit of detection and the limit of quan­ tification were 5ng/ml and 1Ong/ml respectively.
Linearity
Detector response for vitamin A, E and internal standards was linear to at least 100mg/ml. The resulting data was plot­ ted as peak height and peak area versus concentration and studied by the linear regression analysis. Tables 1 and 2 pres­ ent the results for the assessment of the goodness of fit/lack of fit for both compounds (vitamin A and vitamin E) in serum. The goodness of fit was highly significant. A test for lack of fit indicated that the linear model is appropriate for establishing a relationship between the concentration and the response. In general, correlation coefficients above 0.98 were observed dur­ ing the validation experiments.
The linearity for cotinine was 10-1000ng/ml. The urine samples with cotinine levels higher than 1000ng/ml were dis­ solved.
Figure 3. Chromatogram from DAD detector obtained at 1=292nm. It shows the peaks of the a-tocopherol
and a-tocopherol acetate.
Precision and accuracy
A summary of the results on precision and accuracy as derived from the measured concentration for the validation samples is presented in tables I and II.
The within-run R.S.Ds. were below 9.44% for tocopherol and below 9.52% for retinal at all concentration levels in serum. Criteria of acceptance are ±15% at all level concentra­ tions only at the LLOQ are ±20% [13, 21, 22].
The coefficient of variance (CV) for series of cotinine 200ng/ml controls was 5.13%.
0u
."c '
f-
260
Nr 4/2009

,.."a'._ .	.,

..
,•o l	.	g
f	!

Gineko/ Pol. 2009, 80, 256-262
PRACE	ORYGINALNE
ginekologia
Bystrowska
B, et aI.
The advantage of this method is its high sensitivity and specificity. The preparation of samples was made following the proceeding described by M.A. Rodriguez-Delgado at al.
[18] This technique of extraction is very simply and allowed analysis both compounds vitamin E and vitamin A in the same biological sample. Detection both vitamins have been developed and documented by analyses of serum obtained from gynaecology patients enrolled in a cervical carcinoma maintenance treatment.
The proposed method was applied to 12 women diagnosed with cervical intraepithelial neoplasia (CIN) and 16 with cer­ vical cancer in order to determine serum concentrations of retinal and a-tocopherol in these patients. The mean age of patients was 39 (19-67).
The range of concentration of vitamin A in this group of 28 patients was 1.40mg/lserious one? -7.71mg/l and of vitamin E was 2.6lmg/l-17.93mg/1. The minimal and maximal concen­ trations for both vitamins was measured for every subject. The mean concentration was 8.53mg/l for vitamin E and 3.75mg/l for vitamin A.
Cotinine in urine samples in the studied group ranged from 0.0ng/ml to 5917,4ng/ml. The concentration of cotinine in urine was in range 34.3-73.9ng/ml in 7 persons what indi­ cates they were only exposed to environmental tobacco smoke. The concentration of cotinine in range 0.0-8.8ng/ml in 9 patients indicates a non-smoking person. The concentration of cotinine in range 274.0-5917.4ng/ml in 10 patients indicates active smokers.
The mean concentration of vitamin A and E were respec­ tively 3.35mg/l and 8.29mg/l in the group of active smokers, 4.22mg/l and 8.65mg/l in the group of patients exposed to environmental tobacco smoke, 4.68mg/l and 10.32mg/l in the group of non-smokers.
Table I. Goodness of fit and lack of fit for Vitamin E (n=12).
Table II. Goodness of fit and lack of fit for Vitamin A (n=12).
Recovery
For determination of recovery of the analysis, the mean peak heights obtained for triplicate measurements were com­ pared with the mean peak heights obtained from triplicate direct injections performed in the same run. The five recover­ ies thus obtained were used to calculate the mean recovery.
Discussion
Squamous cervical cancer arises from the metaplastic epithelium of the transformation zone (TZ) ( squamocolumnar junction) and develop slowly through progressive dysplastic changes to carcinoma in situ and invasive cancer.
Cervical intraepithelial neoplasia (CIN) is divided into three stages according to the degree of epithelial dysplasia and differentiation.
Over the past few decades significant advances have been made in understanding the molecular genetics underlying the development of human cancers. However, we are still far from constructing complete sequelae of events leading to the devel­ opment of invasive cervical cancer. There are many associated risk factors including number of sexual partners, parity, oral contraceptives use, smoking, immunological system alter­ ations and lack of antioxidants [23] but persistent human papillomavirs (HPV) 16 and 18 infection is the most signifi­ cant factor in aetiology of this disease [24].
The applied method to determinate retinal and a-toco­ pherol in serum uses an isocratic separation with only methanol as mobile phase at l.0ml/min as flow rate which it allows a fast elution of both compounds from the column. This method do not require complex mobile phase. Applied DAD detector allows to measuring both vitamins in the same chromatographic separation. As can be observed the matrix effect obtained at 292nm can be eliminated measuring only a-tocopherol and a-tocopherol acetate (IS) at 292nm and retinal acetate (IS) and retinal at 325nm.
Conclusions
Quantitative determination of cotinine may reveal which patients are in fact active smokers and which are simply exposed to environmental tobacco smoke.
The advantage of this method is that it measures both compounds in a more rapid, reproducible and accurate man­ ner than in previous HPLC studies [9, 12, 18, 19, 20].
The compounds (vitamin A and E and internal standards) are measured in the same sample at the same time.
Acknowledgments:
This work was supported by the grant from the State Committee for Scientific Research of Poland- KEN (No 2 P05C 018 27 and 2 P025C 012 27).
2
1,c
0u
.c"'
f-
261
© 2009 Polskie Towarzystwo Ginekologiczne
Concentration (µg/ml)

mean

SD
R.S.D.
(%)
Blas
(%)
0.25
0.26
0.02
9.52
4.82
1
1.06
0.05
5.18
6.15
5
5.22
0.06
1.06
4.47
Concentration (µg/ml)

mean

SD
R.S.D.
(%)
Bias
(%)
0.25
0.31
0.01
4.19
23.08
1.5
1.27
0.12
9.44
-15.01
10
9.55
0.89
9.37
-4.54

PRACE	ORYGINALNE
ginekologia
Gineko/ Pol. 2009, 80, 256-262
Rapid HPLC method for the determination of vitamin A and E and cotinine concentration ..
References
1. Dekkers J, van Doomen L, Kemper H. The Role of AntioxidantVrtamins and Enzymes in the Prevention of Exercise-Induced Muscle Damage. Sports Med 1996, 21, 213-238.
2. Fumich R, Essig G. Hypervitaminosis A. Case report in adolescent soccer player. Am J Sports Med 1983, 11, 34-37.
3. Groff J, Gropper S, Hunt S The Fat Soluble Vitamins. In Advanced Nutrition and Human Metabolism. Minneapolis West Publishing Company, 1995, 284-324.
4. Kanter M, Nolte L, Holloszy J. Effects of an antioxidant vrtamin mixture on lipid peroxi­ dation at rest and postexercise. I Appl Physiol 1993, 74, 965-969.
5. Kanter M. Nutitional Antiozidants and Physical Activity. In Nutrition in Exercise and Sport. 3rd Edition. Ed. Wolinsky I. New York CRC Press, 1998, 245-253.
C
0
Mutschler E, Geisslinger G, Kroemer H, [i wsp.]. Farmakologia i Toksykologia. Red. Danysza A. Wroc!aw Urban&Partner, 2001, 751-763.
National Academy of Sciences. Dietary Reference Intakes: Recommended Intakes for Individuals. Food and Nutrition Board, Institute of Medicine. 2000.
Stacewicz-Sapuntzakis M. Vitamin A and Caroteniods. In: Sports Nutrition Vitamins and Trace Minerals. Ed. Wolinsky I, Driskell J. New York CRC Press, 1997, 101-110.
Taibi G, Nicotra C. Development and validation of a fast and sensitive chromatograph­ ic assai for all-trans-retinol and tocopherols in human serum and plasma using liquid­ liquid extraction. J Chromatogr B Analyt Technol Biomed Life Sci 2002, 780, 261-267.
6.
7.
0u
."c"
8.
f-
9.
1. Viguie C, Packer L, Brooks G. Antioxidant supplementation affects indices of muscle trauma and oxidant stress in human blood during exercise. Med Sci Sports Exercive. 1989, 21, 16.
2. Pacifici R, Pichini S, Altieri I, [et al.] Determination of nicotine and two major metabo­ lites in serum by solid-phase extraction and high-performance liquid chromatography, and high-performance liquid chromatography-particle beam mass spectrometry. J Chromatogr. 1993, 612, 209-213.
3. Zuccaro P, Altieri I, Rosa M, [et al.]. Determination of nicotine and four metabolites in the serum of smokers by high-performance liquid chromatography wrth ultraviolet detection. J Chromatogr. 1993, 621, 257-261
4. Benowitz N, Jacob P. Metabolism of nicotine and cotinine studied by a dual stable iso­ tope method. Clin Pharmacol Ther. 1994, 56, 483-493. Benowitz N. Biomarkers of Environmental Tobacco Smoke Exposure. Environ. Health Perspect. 1999, 107, Suppl 2, 349-355.
5. Piekoszewski W, Florek E. Markery naraienia na dym tytoniowy Red. Katedra i Zak!ad Toksykologii Akademia Medyczna im. Karola Marcinkowskiego, Poznar\ 2001
6. Sobczak A, Wardas W, Szoltysek-Bo!dys I. Ocena narazenia student6w, mieszkar'lc6w domu studenckiego, na dym tytoniowy Przeg Lek. 2004, 61, 1094-1097.
7. Pach J, Piekoszewski W The preliminary study of the addictive substances used by high school students. Acta Pol Toxicol 2000, 8, 179-185.
8. Piekoszewski W, Florek E, Wrzosek J. Using urinary cotinine tests to check the validity of tobacco smoking and ETS exposure survey responses. 3rd European Conference on Tobacco or Health Closing the Gaps - Solidarity for Health. Abstract Ill A4, Warsaw, 2002.
9. Rodriguez-Delgado M, Diaz-Flores Estevez J, Diaz-Flores Estevez F, [et al.]. Fast deter­ mination of retinol and alpha-tocopherol in plasma by LC. J Pharm Biomed Anal 2002, 28, 991-997.
10. Ake M, Fabre H, Malan A, [et al.] Column liquid chromatography determination of vita­ mins A and E in powdered milk and local flour: a validation procedure. J Chromatogr A 1998, 826, 183-189.
11. Miyagi M, Yokoyama H, Shiraishi H, [et al.] Simultaneous quantification of retinol, reti­ nal, and retinoic acid isomers by high-performance liquid chromatography with a sim­ ple gradiation. J Chromatogr B Biomed Sci Appl 2001, 757, 365-368
12. Benowitz N. Biomarkers of Environmental Tobacco Smoke Exposure. Environ. Health Perspect. 1999, 107, Suppl 2, 349-355.
13. Kanter M, Nolte L, Holloszy J. Effects of an antioxidant vrtamin mixture on lipid peroxi­ dation at rest and postexercise. I Appl Physiol 1993, 74, 965-969.
14. Jach R, Basta A, Szczudrawa A. Role of immunomodulatory treatment with lscador QuS and lntron A of women with CIN1 with concurrent HPV infection, Ginekol Pol 2003, 74, 729-735.
15. zur Hausen H. Papillomaviruses and cancer: from basic studies to clinical application.
Nat Rev Cancer. 2002, 2, 342-350.
C
0
2
0u
."c'
f-
262
Nr 4/2009