IOP
Phys. Med. Biol. 64 (2019) 055017 (14pp)
https://doi.org/10.1088/1361-6560/aafe20
&
Physics in Medicine
 Biology
Phys. Med. Biol.
64
PAPER
OPEN ACCESS
2019
Feasibility study of the positronium imaging with the J-PET
tomograph
©
 2019 Institute of Physics and Engineering in Medicine
RECEIVED
31 May 2018
P Moskal1, D Kisielewska1, C Curceanu2, E Czerwiński1, K Dulski1, A Gajos1, M Gorgol3,
B Hiesmayr4, B Jasińska3, K Kacprzak1, Ł Kapłon1, G Korcyl1, P Kowalski5, W Krzemień6,
T Kozik1, E Kubicz1, M Mohammed1,7, Sz Niedźwiecki1, M Pałka1, M Pawlik-Niedźwiecka1,
L Raczyński5, J Raj1, S Sharma1, Shivani1, R Y Shopa5, M Silarski1, M Skurzok1, E Stępień1,
W Wiślicki5 and B Zgardzińska3
PHMBA7
REVISED
21 December 2018
ACCEPTED FOR PUBLICATION
14 January 2019
055017
PUBLISHED
Faculty of Physics, Astronomy and Applied Computer Science, Jagiellonian University, 30-348 Cracow, Poland
1
7 March 2019
INFN, Laboratori Nazionali di Frascati, 00044 Frascati, Italy
2
P Moskal et al
Institute of Physics, Maria Curie-Skłodowska University, 20-031 Lublin, Poland
3
Original content from
Faculty of Physics, University of Vienna, 1090 Vienna, Austria
4
this work may be used
Department of Complex Systems, National Centre for Nuclear Research, 05-400 Otwock-Świerk, Poland
5
under the terms of the
Feasibility study of the positronium imaging with the J-PET tomograph
Creative Commons
High Energy Physics Division, National Centre for Nuclear Research, 05-400 Otwock-Świerk, Poland
6
Attribution 3.0 licence.
Department of Physics, College of Education for Pure Sciences, University of Mosul, Mosul, Iraq
7
Any further distribution
E-mail: daria.kisielewska@uj.edu.pl
Printed in the UK
of this work must
maintain attribution
Keywords: positron emission tomography, positronium atom, J-PET
to the author(s) and the
title of the work, journal
PMB
citation and DOI.
Abstract
A detection system of the conventional PET tomograph is set-up to record data from
+
−
e
10.1088/1361-6560/aafe20
annihilation into two photons with energy of 511 keV, and it gives information on the density
distribution of a radiopharmaceutical in the body of the object. In this paper we explore the
5
possibility of performing the three gamma photons imaging based on ortho-positronium
annihilation, as well as the possibility of positronium mean lifetime imaging with the J-PET
1361-6560
tomograph constructed from plastic scintillators. For this purposes simulations of the ortho-
positronium formation and its annihilation into three photons were performed taking into account
distributions of photons’ momenta as predicted by the theory of quantum electrodynamics and
1
the response of the J-PET tomograph. In order to test the proposed ortho-positronium lifetime
image reconstruction method, we concentrate on the decay of the ortho-positronium into three
14
photons and applications of radiopharmaceuticals labeled with isotopes emitting a prompt gamma.
The proposed method of imaging is based on the determination of hit-times and hit-positions of
registered photons which enables the reconstruction of the time and position of the annihilation
point as well as the lifetime of the ortho-positronium on an event-by-event basis. We have simulated
the production of the positronium in point-like sources and in a cylindrical phantom composed of
a set of different materials in which the ortho-positronium lifetime varied from 2.0 ns to 3.0 ns, as
expected for ortho-positronium created in the human body. The presented reconstruction method
for total-body J-PET like detector allows to achieve a mean lifetime resolution of
40 ps. Recent
∼
positron annihilation lifetime spectroscopy measurements of cancerous and healthy uterine tissues
show that this sensitivity may allow to study the morphological changes in cell structures.
1.
Introduction
Positron emitted inside the human body can either annihilate directly with one of the electrons of the examined
organism or it creates the metastable state of electron and positron called positronium (Ps). Positronium may be
trapped inside free volumes between and within molecules of the examined patient. Currently, in the commercial
PET technique, the phenomenon of positronium production is neither recorded nor used for imaging.
Imaging of the properties of positronium inside the body may deliver new diagnostic information. The aver-
age lifetime of positronium depends on the size of the free volumes between atoms (Tao 1972, Eldrup et al 1981,
7
Schrader and Jean 1988, Coleman 2000) and there are indications (Liu et al 2007, Pietrzak et al 2013, Jasińska et al
March
© 2019 Institute of Physics and Engineering in Medicine
2019

---
2017a, 2017b) that it is correlated with the stage of the development of metabolic disorders of the human tissues.
Therefore, an image of the average lifetime and production probability of the positronium formed inside the
human body during the routine PET imaging (Moskal et al 2013), as well as the fraction of its annihilations into
three-photons (Kacperski and Spyrou 2005, Jasińska and Moskal 2017) may deliver information complementary
to the standardized uptake value (SUV) and useful for the diagnosis. In this article we explore a possibility of
positronium average lifetime imaging with the J-PET tomograph constructed from plastic scintillators (Moskal
et al 2016, Niedźwiecki et al 2017). We concentrate on the decay of the ortho-positronium (o-Ps) into three pho-
tons and applications of radiopharmaceuticals emitting prompt gamma.
In more than 40% cases, the electron-positron annihilation proceeds in the tissue via creation of positro-
nium (Harpen 2004, Jasińska et al 2017b) and even in water the ortho-positronium is formed with the prob-
ability of about 25% (Stepanov et al 2010). Ortho-positronium decays in vacuum into three photons, however,
in the human body, due to the large probability of the pick-off processes (Garwin 1953), the three-photon events
(both from ortho-positronium decays and direct
 processes) constitute altogether only about 0.5%
→ 3γ
of electron-positron annihilations (Jasińska and Moskal 2017, Jasińska et al 2017b). However, such a relatively
low fraction of three-photon annihilations occurring in the human tissues may be compensated by increased
detection efficiency of the PET detector and by improved spatial resolution of the single-event annihilation point
reconstruction (Kacperski and Spyrou 2005, Gajos et al 2016). At present, PET tomographs record and use for
image reconstruction less than 1% of annihilation events occurring in the body (Cherry et al 2017a, 2017b), In
this context it is important to emphasize that there is an ongoing development of novel modalities for the whole
and total-body PET imaging which shall be characterized by about 40 times higher sensitivity (Cherry et al 2017a,
2017b, Viswanath et al 2017, Zhang et al 2017) with respect to the presently available PET systems (Slomka et al
2016, Vandenberghe et al 2016). Moreover, additional improvement in the time resolution to 100 ps, combined
with the total-body PET, may considerably increase the sensitivity even by a factor of 200 (Cherry et al 2017b).
Therefore, an efficient positronium imaging may become feasible in the nearest future.
The reconstruction of positronium lifetime requires determination of times of its creation and annihilation.
These can be achieved when applying radiopharmaceuticals labeled with isotopes as e.g. Scandium-44 which
after emission of the positron changes into a daughter nucleus in an excited state (Szkliniarz et al 2015, Walczak
et al 2015). The daughter nucleus subsequently de-excites through emission of one or several gamma quanta
(often referred to as prompt gamma). Thus, as a result of the decay both: e
 and
 are emitted e.g. via the follow-
γ
∗
44
ing process:
, where,
 denotes a neutrino whose interaction probability
Sc →
Ca
ν
→
Ca γ e
ν
ν
with the patient and detector is negligible.
β
Recently, applications of radiopharmaceuticals labeled with metallic
 emitters as radioisotopes of scan-
43
dium (
Sc,
Sc) have rapidly expanded, particularly for oncologic imaging purposes (Hofman et al 2012, Kra-
jewski et al 2013, Huclier-Markai et al 2014, Walczak et al 2015). For example the
Sc labeled DOTATATE due to
its high binding affinity for the somatostatin receptors can enhance significantly diagnostics quality of neuroen-
docrine cancers (Singh et al 2017).
In the current PET imaging the prompt gamma constitutes a source of background since when reacting via
Compton effect it may give a signal in the detector which can be misclassified as signal from a 511 keV photon
from the electron-positron annihilation. Analogously, in typical clinical PET, positron-electron annihilation to
three photons (
) constitutes a source of unwanted background which is discarded together with the
→ 3γ
multi-photon events due to the random coincidences and scatterings in the detectors.
However, the triple coincidences originating from
 processes or occurring when utilizing the
→ 3γ
 emitters may be useful for imaging. Since more than a decade, there is also an ongoing research aiming at
β
γ
improvement of spatial resolution of the annihilation point reconstruction when using
 emitters by com-
β
γ
bining standard PET tomography technique with the various types of Compton cameras for the registration of
the prompt gamma (Grignon et al 2007, Donnard et al 2012, Lang et al 2012, 2014, Oger et al 2012, Thirolf et al
2015). Also a fraction of detector scatter coincidences may be applied to improve the image by recovering true
Lines of Response (LORs) from triple coincidences (Cal-González et al 2015, Lin et al 2016).
Image reconstruction based on the three-photon annihilation was also proposed more than decade ago by
Kacperski et al (2004), Kacperski and Spyrou (2005) and Abuelhia et al (2007). The proposed method enables a
reconstruction a place of the
 annihilation based on the measurement of energies of three photons
→ 3γ
and positions of their interactions in the detector material. The application of energy and momentum conserva-
tion laws allows to reconstruct the annihilation point on an event-by-event basis. However, the precise measure-
ment of photon energies requires applications of high-resolution semiconductor detectors (such as e.g. CdZnTe
(Kacperski and Spyrou 2005, Abuelhia et al 2007)) rendering the implementation of the full scale PET scanner
relatively expensive.
In this article we present a cost-effective method enabling reconstruction of the density distribution of the
 annihilation points, as well as reconstruction of the image of ortho-positronium properties, such
→ 3γ
as its average lifetime and production probability. The proposed imaging of ortho-positronium properties pro-
2

---
Figure 1.
Photo of the prototype of the Jagiellonian positron emission tomograph (J-PET) scanner. The J-PET scanner is built from
strips of plastic scintillator with dimensions of
, forming cylindrical layers (Moskal et al 2014, 2015, Niedźwiecki
7 × 19 × 500
mm
et al 2017) with the inner diameter of 85 cm and the axial field-of-view (AFOV) of 50 cm. Light signals from each strip are converted
to electrical signals by Hamamatsu R9800 vacuum tube photomultipliers placed at opposite ends of each strip. Note that the present
J-PET prototype consists of 192 scintillator strips only, with empty spaces between the scintillators.
duced in the living organism is based on the time of signals registered with the plastic scintillators (Moskal et al
2013). Thus, the precision of the proposed method relies predominantly on the time resolution of the tomograph
since the measurement of energy losses of registered photons is used only to disentangle signals between prompt
gamma and annihilation photons. The measured times are used to reconstruct the hit-times and hit-positions of
registered photons which enables the reconstruction of the time and position of the annihilation point. A proto-
type of the tomograph enabling application of this method has been built (figure 1) (Niedźwiecki et al 2017,
Dulski et al 2018) and validated for classical PET imaging (Pawlik-Niedźwiecka et al 2017).
Here we explore the possibility of performing positronium imaging using dedicated Monte-Carlo simula-
tion procedures which are described in detail in Gajos et al (2016) and Kamińska et al (2016). The simulations
are carried out assuming an ideal case, that the tomograph is built from four concentric layers of axially arranged
plastics scintillator strips adjacent to each other as indicated in figure 2, left panel.
In order to reconstruct a mean lifetime of ortho-positronium atoms in each voxel of the patient it is necessary
to determine the times of the ortho-positronium creation and decay (see figure 3). The scheme of the detector
with recorded
 annihilation is shown in figure 4. The time of creation is determined using the prompt
o-Ps → 3γ
gamma e.g. from
Sc, since nuclear deexcitation is much faster (few picosends) than positronium lifetime.
The annihilation point and time of
 processes is reconstructed in an analytical way based on the
→ 3γ
measured times and positions of the three photons (X
,Y
,Z
,T
), i
  1,2,3 and knowing that (due to the momen-
=
i
i
i
i
x, y, z
tum conservation principle) all three photons momentum vectors and the annihilation point (
) are lying in
a single plane, referred to as the decay plane (Gajos et al 2016). Co-planarity of the
 reduces the reconstruction
3γ
2
2
2
2
to a two dimensional problem of three equations: (T
  t)
 (
)
  (
)
, where
,
, T
 (for
X
− x
Y
− y
X
Y
c
i
i
i
i
i
i
y
x
) denote hit-coordinates and hit-times in the decay plane reference system,
,
, t denote the annihila-
i = 1, 2, 3
tion point’s coordinates and time in this frame, and c stands for the speed of light. Solving this system of three
equations defined for
 yields the time and location point of the
 annihilation.
i = 1, 2, 3
→ 3γ
The registration of the prompt gamma, in addition to the annihilation photons, permits to measure the dif-
ferences between the time of formation and the time of annihilation of positronium on the event-by-event basis
(Moskal et al 2013, Gajos et al 2016, Kamińska et al 2016). Reconstructed time of the prompt gamma emission
is used as the formation time of the positronium. This approximation leads to negligible uncertainties, since the
time between the emission of the positron and the formation of the positronium, as well as the average lifetime
∗
Ca
of the excited nuclei relevant for PET examinations (e.g.
), are both in the order of a few picoseconds (see
table 1).
Finally, it is worth to note that if the time resolution would be improved in the future down to 10 ps, as
advocated by Lecoq (2017), then
 pick-off annihilations may be used to determine the annihilation time and
2γ
position for each event, which in coincidence with the registration of prompt gamma would allow to determine
lifetime spectra for each voxel and, hence, to reconstruct the ortho-positronium lifetime based on
o-Ps → 2γ
pick-off annihilations. Such ortho-positronium pick-off annihilations for the human body (with the o-Ps life-
τ
time of about
2 ns) are about 70 times (
− 1
) more frequent than
 events. Therefore,
τ
∼
o-Ps → 3γ
vacuum
tissue
τ
tissue

---
Figure 2.
Left: visualization of simulated cylindrical shape detector. Detector consists of four layers and each of them is filled with
scintillators. In this article the simulations were performed assuming scintillators dimensions of
 and the inner
7 × 19 × 500 mm
layer radius of 43 cm. However, for clarity the elements are shown not to scale. In each layer only a selected number of scintillators
with increased x-y dimensions is shown. Inside the detector a cylindrically shaped phantom is placed. Right: visualization of
simulated phantom (R
  10.0 cm, R
  10.1 cm, z
  10.0 cm) with coordinate system as described in the text. The phantom
inner
outer
is composed of various materials in which the lifetime of ortho-positronium atoms varies from 2200 ps to 2300 ps as expected for the
positronium in the human body. A comprehensive discussion of the known results on o-Ps lifetime can be found in section 2.
β
Figure 3.
Scheme of scandium
 decay that leads to formation of ortho-positronium state. The
Sc isotope emits a positron
which thermalizes in the body and forms an ortho-positronium bound state. Independently, the remaining
Ca
 nucleus deexcites
*
after a few picoseconds (NUCLEIDE-LARA—Library for Alpha, X and Gamma Emissions 2017) emitting a prompt gamma with
characteristic energy. Both the time scale of thermalization (Schrader and Jean 1988) and
Ca excited state lifetime are in the order
of picoseconds, while the lifetime of created ortho-positronium is in the order of nanoseconds. This allows to treat the de-excitation
photon emission time as the time when the ortho-positronium was formed. The o-Ps annihilation time is reconstructed based on
gamma photons hit-times and hit-positions in the detector.
Figure 4.
A scheme of a simulated detector and phantom located inside it with an exemplary
 annihilation event.
→ 3γ
For clarity only a single layer of detector is shown with decreased number of scintillators and increased x-y dimensions. Red lines
represent the photons originating from ortho-positronium annihilation, while green line shows prompt gamma originating from
∗
β
 emitter de-excitation, e.g. from
 decay chain. Photon recording strips are marked with
Sc →
Ca
ν →
Ca γ e
ν
corresponding colors.
the proposed method of positronium imaging shall become practical with the advent of total-body PET and
improved time resolutions.
The article is organized as follows: in section 2 the positronium properties in vacuum and in the human body
are briefly reviewed based on the performed experiments up to date. This section concludes with the estima-
4

---
Table 1.
Selected physical characteristics of the exemplary beta-plus isotopes use for both: PET imaging and positron annihilation lifetime
spectroscopy (PALS) investigations. Data were adapted from National Nuclear Data Center (2016).
β
E
Isotope
Half-life
 decay
 (MeV)
Excited nuclei lifetime
γ
Sc
4.0 h
1.16
2.61 ps
Sc →
Ca + e
+ ν
+ γ
14
O
70.6 s
2.31
67.8 fs
14
14
O →
N + e
+ ν
+ γ
tion of the range and values of lifetimes of ortho-positronia produced in the various tissues of the human body
(
1.8 ns –
  2.5 ns), as well as in the expected differences in the ortho-positronium lifetime between the healthy
∼
∼
and cancerous tissues (
50 ps –
  200 ps). The section 3.1 contains a general description of the J-PET detec-
∼
∼
tor in order to illustrate the possibility of applying the introduced method for the ortho-positronium lifetime
imaging simultaneously to the regular PET scan. This section defines also the properties of the detection system
for which the feasibility study of the positronium imaging will be performed in section 3.3. The description of
the Monte Carlo simulation given in section 3.3 comprises: (i) description of the four-layer detector built from
axially arranged plastic strips adjacent to each other, (ii) three-photon annihilation as predicted by the theory of
quantum electrodynamics, (iii) gamma photon interaction within detector build from plastic scintillators as a
function of incident gamma photon energy, (iv) resolution for reconstructing position, time and energy loss of
gamma photons in the detector, (v) reconstruction of annihilation position based on the reconstructed gamma
photon hit information. Next, section 4 includes details on positronium mean lifetime image reconstruction,
discussion about the achievable mean-lifetime resolution and positronium imaging sensitivity.
2.
Positronium in the human body
Similarly to the hydrogen atoms, positronium may be formed in a singlet or triplet spin states, referred to as para-
positronium and ortho-positronium, respectively. Due to the charge conjugation symmetry conservation, the
para-positronium annihilates with emission of an even number of photons and ortho-positronium annihilation
is subject to emission of an odd number of photons. Predominantly these are two and three photons for para-
and ortho-positronium atoms, respectively. The decay time of formed positronium depends on its quantum
state: the observed lifetime for para-positronium localized in vacuum amounts to
, while, due to
∼ 125 ps
the smallness of the fine structure constant and the phase-space volume available, in the decays of the ortho-
positronium its lifetime in vacuum is three orders of magnitude greater (
  142 ns) (Al-Ramadhan and
τ
vac.
Gidley 1994, Jinnouchi et al 2003, Vallery et al 2003). When formed inside matter, positronium may be trapped
inside volumes of lower electron density. However, the positronium trapped in a free volume is susceptible to
the pick-off processes (Garwin 1953), in which positron from the positronium may annihilate with one of the
surroundings electrons. The smaller are the free voids between the atoms the larger is the probability of the
pick-off process and, hence, the smaller is the mean lifetime of the trapped positronium. This effect is significant
in case of ortho-positronium for which the mean lifetime may change in the range from about 142 ns to even
below 1 ns. The correlation is described in the framework of well tested models which relate quantitatively the
mean ortho-positronium lifetime with the size of the free void in the range from 0.2 nm up to 100 nm (Tao
1972, Eldrup et al 1981, Goworek et al 1997, 1998). In addition it is well known from the positron annihilation
lifetime spectroscopy (PALS) studies that the intensity of the production of ortho-positronium atoms is strongly
correlated with the free voids size distributions and their concentration (Goworek et al 1982, Schrader and Jean
1988, Kobayashi et al 1989, Coleman 2000, Jasińska and Dawidowicz 2003, Jasinska et al 2003).
In the human body, positronium atoms can be created and trapped both, in the dense tissue and in the bio-
fluids—importantly, including water. In purified water the o-Ps lifetime amounts to 1.8 ns and its production
probability is equal to about 25
. There are few works demonstrating the application of PALS technique to
%
model and living biological systems reporting differences between normal and cancerous cells and changes of
the o-Ps lifetime during dynamical processes undergoing in cells (Jean and Ache 1977, Jean et al 2006, 2007, Liu
et al 2007, 2008, Yas et al 2012, Axpe et al 2014, Kubicz et al 2015). Recently, significant differences in PALS param-
eters between normal and diseased tissues were also observed in samples of uterine leiomyomatis and normal
muscles tissues taken from women-patients after surgery of uterus tumors (Jasińska et al 2017a, 2017b). The
results of above quoted research indicate that in all investigations performed to date there is a difference in PALS
parameters determined for healthy and cancerous tissues. Thus, as suggested in Moskal et al (2013) and Jasińska
and Moskal (2017), measurements of properties of ortho-positronium atoms (such as lifetime and production
probability, or
 rate ratio) which are formed inside the human body during a routine PET imaging may
3γ/2γ
deliver information useful for the diagnostics. In this article as a first step on a way towards realisation of such
measurements, we present a feasibility study of the reconstruction method of the image of ortho-positronium
lifetime proposed in Moskal et al (2013).
5

---
The mean ortho-positronium lifetimes in the tissues varies from about 1.8 ns (as in pure water) to about
2.5 ns (or even up to about 4 ns as measured for the human skin with low energy positron beam (Chen et al
2012)). Whereas, the mean ortho-positronium lifetime differences for healthy and cancerous tissues are in the
range of about 50 ps to about 200 ps (Liu et al 2007, Pietrzak et al 2013, Jasińska et al 2017b).
Taking into account the above information, the feasibility study of the positronium lifetime imaging will be
performed assuming the formations of ortho-positronium atoms in a set of point-like sources and in a virtual
phantom formed from tissues in which ortho-positronium lifetime varies from 2.0 ns to 3.0 ns. The phantom
will be inserted into an idealized version of the J-PET tomograph consisting of four cylindrical layers. In the
simulations the properties of the detection system, which are briefly described in the following section, will be
taken into account.
3.
Materials and methods
3.1.
Principles of operation of the J-PET detector made of plastic scintillators
The J-PET tomograph (shown in figure 1) is a cylindrically shaped PET detector made of plastic scintillators.
The light signals are converted to electric pulses by photomultipliers connected optically at the two ends of a
scintillator. In plastic scintillators, which are composed of elements of low atomic number, incident photon
interacts mainly through the Compton scattering leading to a continuous energy loss distribution. Figure 6
presents an example of the energy loss spectra for photons from
 annihilation and for prompt gamma
→ 2γ
14
from
Sc and
O.
In the J-PET tomograph the electric signals from photomultipliers are processed by the front-end electron-
ics based on field programmable gate arrays (FPGA) (Pałka et al 2017), collected by means of the triggerless
data acquisition system (Korcyl et al 2016, 2018), and analysed using a dedicated analysis software framework
(Krzemień et al 2015).
The uncertainty of the hit-time determination and hence the time-of-flight (TOF) resolution depends on the
length of the scintillator strip and the coincidence resolving time (CRT) varies from CRT
 0.230 ns to CRT
≈
≈
0.440 ns when the length changes from 15 cm to 100 cm (Moskal et al 2016), and it may be improved by using a
matrix of silicon photomultiplier (SiPM) readout. In Moskal et al (2016) it was shown that e.g. for the
 SiPM
2 × 5
matrix at two ends of the scintillator strip the coincidence resolving time changes from CRT
 0.170 ns to CRT
≈
≈
0.365 ns when extending an axial field-of-view (AFOV) from 15 cm to 100 cm (Moskal et al 2016).
It is worth emphasizing that even higher time resolution of CRT
  0.100 ns was already achieved with the
small
 LaBr3:Ce(5%) crystals read out with the SiPMs (Schaart et al 2010). More recently a detec-
3 × 3 × 5 mm
tor design with SiPM has been reported, with CRT
  0.085 ns for
 LSO:Ce codoped 0.4%Ca
2 × 2 × 3 mm
crystals and CRT of 0.140 ns for
 crystals with the length as used in the current PET tomographs
2 × 2 × 20 mm
(Nemallapudi et al 2015). Therefore, in section 3.3 we will perform feasibility studies simulating the possibility
of positronium imaging under conservative assumption that CRT is equal to 0.140 ns, as already achieved for the
solutions applicable in the current PET detectors.
3.2.
Event selection criteria and background reduction
An event useful for reconstruction of the o-Ps image is shown pictorially in figure 4. It consists of four registered
gamma interaction points in the scintillator strips (later on referred to as hits), three of them originating from
the
 annihilation and the fourth one from the prompt gamma. After identification of the prompt
o-Ps → 3γ
gamma (e.g. using energy loss criterion E
  370 keV), the three remaining hits with E
  370 keV become
>
<
dep
dep
then candidates of photons originating from the
 decays useful for the reconstruction of annihilation
o-Ps → 3γ
position and time as indicated in figure 5. However, due to the secondary Compton scattering in plastic scintillator
(discussed in more detail in Kowalski et al (2015, 2016), Kowalski et al (2018) and Kamińska et al (2016)), the
2γ
events may mimic registration of
 ones, as indicated pictorially in figure 7. Yet, we can disentangle true and
3γ
false
 events based on the (i) relation between hit positions and hit times and (ii) angular correlation of relative
3γ
angles between the direction of annihilation photons.
For example, events for which both photons from
 annihilation are registered and one of them is
→ 2γ
◦
scattered causing signals in two detectors (middle panel of figure 7) satisfy the condition
 and
θ
+ θ
= 180
12
23
are concentrated along the anti-diagonal of the
 versus
 plot. Detailed studies were done in Kamińska et al
θ
θ
12
23
(2016) and show that the application of the angular correlation criterion suppresses the background from
 by
2γ
4
the factor of about 10
 while rejecting only 3% of true
 signal events and that further requirements based on
3γ
the relation between hit position and hit time may reduce the background by another few orders of magnitude.
Therefore, even though the
 events are expected to constitute only about 0.5% of
 events, the back-
3γ
2γ
ground due to the
 annihilation associated with the secondary scattering in the detector can be reduced to a
2γ
negligible level. Hence, in the next section we present simulations assuming that the hits from the
o-Ps → 3γ
6

---
Figure 5.
A scheme of the detector showing
 annihilation. For clarity only a single layer with registered hits is shown.
o-Ps → 3γ
Red lines represent the gamma photons from ortho-positronium annihilation. Figure shows a single layer of the detector in 3D view
with cross-section of the scintillators shown in grey and their inner side in blue. The trilateration method is used to determine the
annihilation position and time
(x
, y
, t)
 along the annihilation plane. For each recorded photon a circle, which is a set of possible
photon origin points, centered in the hit-position and parameterised with the unknown o-Ps annihilation time is considered. The
intersection of the three circles corresponds to the
 annihilation point.
o-Ps → 3γ
×
10
-
E
 = 511 keV (e
)
140
γ
E
 = 1160 keV (
Sc)
γ
120
14
E
 = 2310 keV (
O)
γ
100
80
60
40
20
0
0
500
1000
1500
2000
2500
E
 [keV]
dep
Figure 6.
Simulated energy loss spectra for annihilation photons (511 keV) and prompt gamma from the decay of isotopes
indicated in the legend. The shown distributions include the energy resolution of the J-PET detector (Moskal et al 2014).
annihilation can be identified and we will focus on the possibilities of the positronium lifetime image reconstruc-
tion, assuming that the time resolution of the detector amounts to CRT
  0.140 ns.
3.3.
Monte Carlo simulations
The feasibility studies of positronium imaging are performed based on the virtual cylindrical phantom placed
inside the J-PET tomograph as shown in figure 2. A cylindrically shaped phantom with radii of R
  10.0 cm,
inner
R
  10.1 cm and z
  10.0 cm is placed along the z axis of the detector (see figure 2, right panel). The volume
outer
of the phantom was divided into 2 parts, each corresponding to the tissues characterized with different average
lifetimes of ortho-positronium atoms. The values of average lifetimes of 2200 ps–2300 ps were chosen. These
values are within the range of lifetimes expected for the ortho-positronium in the human body (as discussed in
sections 1 and 2).
7
Counts

---
Figure 7.
Typical topology of
 (left panel) and background events originating from
 annihilations (middle and right
2γ
o-Ps → 3γ
panels). Red solid lines denote actual photons, the red arrow indicates not registered photon while green dotted lines indicate
reconstructed photon candidates. The numbering of photons is such that
.
θ
< θ
< θ
12
23
31
9
For the studies presented in this article about
 events of o-Ps decays into three photons were generated
3 × 10
and, subsequently, the response of the J-PET tomograph was simulated and then the image of the mean lifetime
of positronium was reconstructed. The simulations were conducted in the following steps: for each event (i) the
position of the annihilation was generated, assuming a homogeneous density distribution within the cylindrical
phantom; (ii) the lifetime of a given o-Ps atom was generated with the exponential probability density distribu-
tion with the mean lifetime depending on the position in the phantom; (iii) the momentum vectors of pho-
8
tons from the
 decay were generated according to the predictions of the quantum electrodynamics
;
o-Ps → 3γ
(iv) hit-position and energy deposition of each photon in the plastic scintillators of the J-PET tomograph were
simulated taking into account the cross sections for the Compton interactions of gamma photons in the plastic
scintillators; (v) the experimental resolutions for the registration of time and position of gamma photons were
accounted for by smearing the values of generated hit-times and hit-positions with the resolution functions of
J-PET, assuming the CRT value of 0.140 ns; (vi) for each registered event, a time and position of o-Ps annihilation
was reconstructed using the trilateration method (Gajos et al 2016) and the image of mean o-Ps lifetime was cre-
ated. The simulation methods for the above listed steps from (i) to (v) were described in details in the previous
publications (Moskal et al 2014, 2015, Moskal et al 2016, Kamińska et al 2016).
At the step of generation of the position of annihilation and the decay time of o-Ps in the phantom, the area of
the phantom was divided into pixels with dimensions of
. The generated mean lifetime distribution
2 cm × 2 cm
z − ρ
as a function of the
 coordinates along the cylinder is shown in figure 8. A colour code indicates the mean
life-time of o-Ps calculated in each pixel, using lifetime values of ortho-positronium atoms generated in the pixel.
The differences in the mean lifetimes observed within each of two sections are due to the finite statistics of gener-
ated events. Thus, figure 8 presents an image of o-Ps mean lifetimes which could be reconstructed in the case of
an ideal detector, assuming that all of the annihilations events were registered with an ideal resolution.
In addition, we preformed a simulation positioning six point annihilation sources in the arrangement
required by the National Electrical Manufacturers Association (NEMA 2012) norm: a six sources (cylinders
r
  0.5 mm, z
  1 mm) were placed as presented in table 2. Simulations included response of the detector using
time and position resolutions as discussed in the text of the article.
4.
Results
As a result of simulated measurements and data selection described in the previous sections for each registered
 event a hit-time and hit-position of the annihilation and deexcitation photons are determined. For
o-Ps → 3γ
each event these enable a reconstruction of time and position of the annihilation point using the trilateration
method (Gajos et al 2016), as well as the lifetime of ortho-positronium estimated as a difference between the
emission time of annihilation photons and emission time of the prompt gamma. For the purpose of positronium
mean-lifetime image reconstruction an object is divided into voxels in which
  annihilations
o-Ps → 3γ
are recorded. In each voxel the mean lifetime is estimated by the arithmetic mean of individual lifetimes.
Reconstruction time-resolution of a single ortho-positronium lifetime strongly depends on detector timing
capabilities (see left panel of figure 9) while the resolution of mean lifetime determination depends additionally
8
The positronium has a small boost along the initial positron momentum. Due to this the annihilation photons’ momenta
slightly deviate from co-planarity in the detector frame of reference. This effect is included in the simulations and its
contribution is negligible with respect to the resolution achieved so far (Gajos et al 2016).
8

---
30
2.8
20
2.6
10
0
2.4
10
2.2
20
2
30
1.8
5
4
2
1
0
1
2
4
5
Z [cm]
Figure 8.
Generated distribution of mean ortho-positronium lifetime in the phantom as a function of z and
 coordinates. Pixel size
ρ
is 2 cm
 2 cm. Radiopharmaceutical is uniformly distributed, however the phantom is composed of two different materials in which
×
the lifetime of ortho-positronium atoms is equal to 2300 ps (green area) and 2200 ps (blue area).
Table 2.
Coordinates of simulated ‘point-like’ sources positioned according to the NEMA norm. Each source is characterized by a
different o-Ps lifetime.
Position
Coordinates (cm)
Simulated o-Ps lifetime (ns)
1
(1,0, 0 )
2.0
2
(10,0, 0)
2.4
(20,0, 0 )
2.8
4
(1,0, 18.75)
2.2
5
(10,0, 18.75)
2.6
6
(20,0, 18.75)
3.0
40
160
35
140
30
25
120
20
100
15
80
10
60
5
×
10
20 30 40 50 60 70 80 90 100
0
100
200
300
σ
(t
) [ps]
N
rec
hit
Figure 9.
Left: resolution of ortho-positronium reconstructed decay time obtained by the trilateration method as a function of
detector time resolution. Achievable resolution strongly depends on detector timing capabilities. Right: resolution of the mean
lifetime determination as a function of detected entries in a single voxel. An example for
σ(t
)
= 40
 ps (corresponding to
hit
τ
= 2010
CRT
 140 ps) and
 ps is presented.
≈
on the number of detected events (see right panel of figure 9). The mean-lifetime resolution of about 40 ps is
achievable for about 3000 events per voxel.
The left panel of figure 10 shows the image of positronium lifetime reconstructed, assuming the CRT of
140 ps (corresponding to
 of about 40 ps). The original (ideal) image is shown in figure 8. The recon-
σ(t
)
hit
structed mean lifetime image (figure 10) indicates, that for CRT of about 140 ps the method allows to distinguish
9
 [cm]
ρ
 [ns]
τ
) [ps]
rec
(t
σ
) [ps]
τ
(
σ

---
τ
30
Gen
τ
2.8
2.8
20
Rec
2.6
2.6
10
0
2.4
2.4
10
2.2
2.2
20
2
2
30
1.8
1.8
5
4
2
1
0
1
2
4
5
30
20
10
0
10
20
30
ρ
Z [cm]
 [cm]
Figure 10.
Left: reconstructed image of the phantom composed of two materials in which the lifetime of ortho-positronium
atoms is equal to 2200 ps (blue area) and 2300 ps (green area). Generated image is shown in figure 8. The o-Ps annihilation time is
reconstructed by using the trilateration method and assuming time resolution of
σ(t
)
= 40
 ps corresponding to CRT of about
hit
140 ps, as discussed in the introduction. Right: comparison between generated (black line) and reconstructed (red line) o-Ps lifetime
image shown as a function of transverse radius.
between o-Ps lifetime differing by 100 ps. Right panel of figure 10 compares projections of the generated and
simulated images. It indicates that the reconstructed positronium lifetime image is in a very good agreement with
the generated one. The achieved time resolution for positronium mean lifetime includes also uncertainties due to
the depth of interaction (which can be estimated to about 70 ps per event (Moskal et al 2016). There is still a room
for the future improvement of both the time resolution and accuracy of DOI determination (see e.g. van Dam
et al (2011), Pizzichemi et al (2016) and Lecoq (2017).
Results of the test of the discussed image reconstruction method on the ‘point-like’ sources arranged accord-
ing to the NEMA norm is presented in figure 11. In figure 11 the coronal (XZ) cross-section along y
  0 cm is
shown, for all six positions of the source. In each source a different o-Ps lifetime were assumed (see table 2).
Reconstructed and generated ortho-positronium lifetime are in agreement (compare figures 11(a) and (b)).
Spread originates from uncertainty of the reconstructed spatial coordinates of the annihilation point.
The conventional PET imaging on simulated
 events (see figure 11(d)) and trilateration-based reconstruc-
2γ
tion of ortho-positronium decays into
 photons (see figure 11(c)) were preformed. The filtered back projec-
3γ
tion (FBP) is a standard approved by NEMA for the estimation of spatial resolution. Therefore, we employed 3D
Reprojection (3DRP) FBP algorithm for image reconstruction, implemented in STIR framework (Thielemans
et al 2012). STIR does not support multilayer geometry, hence all events were remapped onto a single layer of
radius R
  43.73 mm, comprising the same 7-mm wide strips, 384 in total (an ideal scanner, defined in previous
works (Shopa et al 2017). After the smearing of times and Z-positions of hits were applied, a SAFIR package was
utilised (Becker et al 2017) for data conversion from list-mode into a STIR-compatible format. In figure 11(d),
the coronal (XZ) cross-section along y
  0 cm is shown, for all six positions of the source, aggregated into one
image. It is important to note that transverse spatial resolution is expected to be higher for the initial 4-layer scan-
ner, since it provides better sinogram sampling in projection space.
The efficiencies for the registration of three photons from the
 annihilation with the four-layer
o-Ps → 3γ
J-PET with AFOV of 50 cm and 200 cm are shown in the figure 12. The efficiencies are plotted as a function of the
energy loss threshold. For the threshold of 25 keV the determined efficiency for
 amounts to 2% and
o-Ps → 3γ
8
9% for AFOV of 50 cm and 200 cm, respectively. Thus, about
 events can be collected assuming that: (i)
3γ
2 × 10
activity of 370 MBq (10 mCi) is administered to the patient (Cherry et al 2017b), (ii) the data is acquired during
20 min according to the standard total body protocol, (iii) the sensitivity for
 detection with total-body J-PET
3γ
amounts to 9%, (iv) the fraction of 3
 annihilations is equal to 0.5%.
γ
In addition, taking into account the sensitivity for the detection and selection (energy loss larger than
370 keV) of the prompt gamma equal to about 27% (blue solid line in the right panel of figure 12), one obtains
7
 registered
 events in coincidence with the prompt gamma. This gives a promising perspective, espe-
3γ
5 × 10
cially in view of the recent works showing that for the current PET with AFOV
  25.5 cm (Siemens mMR PET/
6
7
MR scanner) even 10
 counts in the scan could be sufficient to provide good images (Yan et al 2016).
 reg-
5 × 10
istered events result in, on average, about 700 events per cubic centimeter of the examined patient. However, the
effective imaging sensitivity may be significantly enhanced by improving spatial resolution of the determination
10
 [cm]
ρ
 [ns]
τ
 [ns]
τ

---
30
30
25
25
2.5
2.5
20
20
2
2
15
15
1.5
1.5
10
10
1
1
5
5
0.5
0.5
0
0
5
5
0
0
5
0
5
10
15
20
25
30
5
0
5
10
15
20
25
30
X [cm]
X [cm]
6
×
10
30
0.4
0.35
25
0.3
20
0.25
15
0.2
10
0.15
5
0.1
0
0.05
5
5
0
5
10
15
20
25
30
X [cm]
Figure 11.
A six point sources were placed according to the positions denoted in table 2. Each source has a different ortho-
positronium lifetime (a). Application of
 trilatelation-based algorithm (Gajos et al 2016) allows to reconstruct distribution of
3γ
spatial coordinates of annihilation points (c) as well as the mean lifetime distribution of the ortho-positronium (b). The voxel size is
equal to
. For comparison result of standard PET imaging for
 events is shown (d).
2γ
2 cm × 2 cm × 2 cm
1
1
10
γ
prompt
; 200 cm strip
2
γ
10
prompt
; 50 cm strip
γ
; 200 cm strip
γ
; 50 cm strip
10
4
10
5
10
6
10
7
10
0
50
100 150 200 250 300 350
E
 [keV]
th
Figure 12.
Registration efficiency for prompt gamma (upper lines) and
 as a function of the energy-loss threshold
o-Ps → 3γ
(determined at the center of the tomograph, taking into account geometrical acceptance, probability of gamma photons registration
in the plastic scintillator and J-PET detector resolution). The dashed and solid lines indicate efficiencies for 50 cm and 200 cm strip
length, respectively.
11
Z [cm]
 [ns]
τ
Z [cm]
 [ns]
τ
Z [cm]
Counts
det
ε

---
of the annihilation position for a single event. The signal-to-noise ratio (SNR) in the conventional PET image is
√
proportional to the square root of sensitivity S,
, and it may be increased effectively by improving the
SNR ∼
S
uncertainty in the localization of the annihilation point
 (SNR
) (Conti 2009, Eriksson and Conti
∆X
∼
1/∆X
∆X = 7.5
2015), where
 is proportional to the time resolution (e.g. CRT
  0.5 ns gives
 cm). For the activ-
∆X
∆X = 7.5
ity distributed homogeneously in 40 cm diameter phantom the SNR increases with
 cm by a factor
of about 2.3 (Townsend 2008), corresponding to the effective sensitivity growth by a factor of about 5.3. Thus,
improving
 to 1 cm would lead to a gain in sensitivity by a factor of about 40 with respect to non-TOF PET
∆X
systems. The three photon iterative image reconstruction methods remains to be elaborated, nevertheless, this
challenge is far beyond the scope of this article. Thus, presently, in the case of the three photon reconstruction dis-
cussed in this article, it is impossible to precisely evaluate the gain in effective sensitivity due to the improvement
of the spatial resolution of annihilation point reconstruction. However, the estimations of the sensitivity of about
9% for three photon registration with the 200 cm long J-PET, and the estimated spatial resolution for a single
event of about 1 cm (Gajos et al 2016) give promising perspectives for the future application of the positronium
imaging even though the fraction of three photons annihilations in human body is expected to be at the level of
about 0.5% only.
5.
Summary and perspectives
In this article we presented a cost-effective method enabling reconstruction of the density distribution of the
 annihilation points, as well as reconstruction of the image of ortho-positronium average lifetime.
→ 3γ
The proposed method is based on the time of signals registered with plastic scintillators and its precision relies
predominantly on the time resolution of the tomograph. The arrival times of signals measured at the ends of
scintillator strips are used to reconstruct the hit-times and hit-positions of registered gamma photons. Next, the
annihilation point is reconstructed analytically, on an event by event basis, using trilateration method and taking
advantage of the fact that all three photons from the decay of ortho-positronium are lying in single plane.
We explored the possibility to apply the introduced method for the reconstruction of images of mean lifetime
of ortho-positronium which are created in patients body during the routing PET investigations. In the patient
about 40% of
 annihilations proceed via formation of positronium atoms with about 0.5% decaying into
 (Harpen 2004, Jasińska et al 2017b). Though the relative rate of
 annihilations is low, we have argued that
3γ
3γ
the positronium imaging may become feasible with the advent of the total-body PET scanners. The average life-
time of positronium, as well as the fraction of its annihilations into three-photons, depend on the size of the free
volumes between atoms and there are experimental indications (Liu et al 2007, Pietrzak et al 2013, Jasińska et al
2017a, 2017b) that it is correlated with the metabolic disorders of the human tissues. Therefore, imaging of the
properties of positronium inside the body may deliver new in vivo diagnostic information complementary to the
standardized uptake value (SUV index).
In this article we focused our attention on the feasibility studies of imaging of the ortho-positronium mean
lifetime which requires application of radiopharmaceuticals labeled with isotopes emitting a prompt gamma
(e.g.
Sc). The prompt gamma is used to determine the creation time of positronium and thus the lifetime may
be determined event-by-event based on the reconstructed time difference between annihilation and creation.
However, the introduced method may be also used to determine images of the ratio of the rate of
 to
 anni-
3γ
2γ
hilations which is also sensitive to the size and concentration of free voids between atoms (Jasińska and Moskal
2017). In this case the prompt gamma is not required and such an image may be created for all radiopharmaceu-
β
ticals, independently of the kind of the used
 tracer.
Acknowledgments
The authors acknowledge technical and administrative support of A Heczko, M Kajetanowicz and W
Migdał. This work was supported by The Polish National Center for Research and Development through
grant
INNOTECH-K1/IN1/64/159174/NCBR/12,
the
Foundation
for
Polish
Science
through
the
MPD
and TEAM/2017-4/39 programmes, the National Science Centre of Poland through grants no. 2016/21/B/
ST2/01222, 2017/25/N/NZ1/00861, the Ministry for Science and Higher Education through grants no. 6673/
IA/SP/2016, 7150/E-338/SPUB/2017/1 and 7150/E-338/M/2017, 7150/E-338/M/2018 and the Austrian Science
Fund FWF-P26783.
ORCID iDs
P Moskal
 https://orcid.org/0000-0002-4229-3548
D Kisielewska
 https://orcid.org/0000-0002-1429-3209
12

---
References
Abuelhia E, Kacperski K and Spyrou N 2007 Three-photon annihilation in PET: 2D imaging experiments J. Radioanal. Nucl. Chem.
271 489–95
Al-Ramadhan A H and Gidley D W 1994 New precision measurement of the decay rate of singlet positronium Phys. Rev. Lett. 72 1632–5
Axpe E, Lopez-Euba T, Castellanos-Rubio A, Merida D, Garcia J A, Plaza-Izurieta L, Fernandez-Jimenez N, Plazaola F and Bilbao J R 2014
Detection of atomic scale changes in the free volume void size of three-dimensional colorectal cancer cell culture using positron
annihilation lifetime spectroscopy PLoS One 9 1–5
Becker R et al 2017 The SAFIR experiment: concept, status and perspectives Nucl. Instrum. Methods Phys. Res. A 845 648–51
Cal-González J, Lage E, Herranz E, Vicente E, Udias J M, Moore S C, Park M A, Dave S R, Parot V and Herraiz J L 2015 Simulation of triple
coincidences in pet Phys. Med. Biol. 60 117
Chen H M et al 2012 Applications of positron annihilation spectroscopy to life science Defect and Diffusion Forum 331 275–93
Cherry S R, Badawi R D, Karp J S, Moses W W, Price P and Jones T 2017a Total-body imaging: transforming the role of positron emission
tomography Sci. Trans. Med. 9 eaaf6169
Cherry S R, Jones T, Karp J S, Qi J, Moses W and Badawi R 2017b Total-body pet: maximizing sensitivity to create new opportunities for
clinical research and patient care J. Nucl. Med. 59 3–12
Coleman P 2000 Positron Beams and their Applications (Singapore: World Scientific)
Conti M 2009 State of the art and challenges of time-of-flight pet Phys. Medica 25 1–11
Donnard J et al 2012 Compton imaging with liquid Xenon and 44Sc: recent progress toward 3 gamma imaging Nucl. Med. Rev. 15 64–7
Dulski K et al 2018 Commissioning of the J-PET detector in view of the positron annihilation lifetime spectroscopy Hyperfine Interact.
239 40
Eldrup M, Lightbody D and Sherwood J 1981 The temperature dependence of positron lifetimes in solid pivalic acid Chem. Phys. 63 51–8
Eriksson L and Conti M 2015 Randoms and tof gain revisited Phys. Med. Biol. 60 1613–23
Gajos A et al 2016 Trilateration-based reconstruction of ortho-positronium decays into three photons with the J-PET detector Nucl.
Instrum. Meth. A 819 54–9
Garwin R L 1953 Thermalization of positrons in metals Phys. Rev. 91 1571–2
Goworek T, Ciesielski K, Jasińska B and Wawryszczuk J 1997 Positronium in large voids. Silicagel Chem. Phys. Lett. 272 91–5
Goworek T, Ciesielski K, Jasińska B and Wawryszczuk J 1998 Positronium states in the pores of silica gel Chem. Phys. 230 305–15
Goworek T, Rybka C, Wasiewicz R and Wawryszczuk J 1982 Positronium trapping at thermally produced defects in naphthalene Phys. Status
Solidi b 113 K9–13
Grignon C et al 2007 Nuclear medical imaging using
 coincidences from 44Sc radio-nuclide with liquid Xenon as detection medium
β + γ
Nucl. Instrum. Methods Phys. Res. A 571 142–5 (Proc. of the 1st Int. Conf. on Molecular Imaging Technology)
Harpen M D 2004 Positronium: review of symmetry, conserved quantities and decay for the radiological physicist Med. Phys. 31 57–61
Hofman M S, Kong G, Neels O C, Eu P, Hong E and Hicks R J 2012 High management impact of Ga-68 dotatate (gatate) PET/CT for imaging
neuroendocrine and other somatostatin expressing tumours J. Med. Imaging Radiat. Oncol. 56 40–7
Huclier-Markai S, Kerdjoudj R, Alliot C, Bonraisin A, Michel N, Haddad F and Barbet J 2014 Optimization of reaction conditions for the
radiolabeling of dota and dota-peptide with 44m/44Sc and experimental evidence of the feasibility of an in vivo PET generator Nucl.
Med. Biol. 41 e36–43
Jasińska B and Dawidowicz A 2003 Pore size determination in vycor glass Radiat. Phys. Chem. 68 531–4 (Proc. of the 7th Int. Conf. on Positron
and Positronium Chemistry)
Jasinska B, Dawidowicz A L and Pikus S 2003 Application of positron annihilation lifetime spectroscopy in studies of crystallization
processes Phys. Chem. Chem. Phys. 5 3289–93
Jasińska B and Moskal P 2017 A new PET diagnostic indicator based on the ratio of 3gamma/2gamma positron annihilation Acta Phys. Pol.
B 48 1577
Jasińska B et al 2017a Human tissue investigations using PALS technique—free radicals influence Acta Phys. Pol. A 132 1556
Jasińska B et al 2017b Human tissues investigation using PALS technique Acta Phys. Pol. B 48 1737
Jean Y C and Ache H J 1977 Studies of molecular association in biological systems by positron annihilation techniques J. Am. Chem. Soc.
99 1623–5
Jean Y, Chen H, Liu G and Gadzia J E 2007 Life science research using positron annihilation spectroscopy: UV-irradiated mouse skin Radiat.
Phys. Chem. 76 70–5 (Proc. of the 8th Int. Workshop on Positron and Positronium Chemistry)
Jean Y, Li Y, Liu G, Chen H, Zhang J and Gadzia J E 2006 Applications of slow positrons to cancer research: search for selectivity of positron
annihilation to skin cancer Appl. Surf. Sci. 252 3166–71 (Proc. of the 10th Int. Workshop on Slow Positron Beam Techniques for Solids and
Surfaces)
Jinnouchi O, Asai S and Kobayashi T 2003 Precision measurement of orthopositronium decay rate using SiO(2) powder Phys. Lett. B
572 117–26
Kacperski K and Spyrou N M 2005 Performance of three-photon pet imaging: Monte Carlo simulations Phys. Med. Biol. 50 5679
Kacperski K, Spyrou N M and Smith F A 2004 Three-gamma annihilation imaging in positron emission tomography IEEE Trans. Med.
Imaging 23 525–9
Kamińska D et al 2016 A feasibility study of ortho-positronium decays measurement with the J-PET scanner based on plastic scintillators
Eur. Phys. J. C 76 445
Kobayashi Y, Zheng W, Meyer E F, McGervey J D, Jamieson A M and Simha R 1989 Free volume and physical aging of poly(vinyl acetate)
studied by positron annihilation Macromolecules 22 2302–6
Korcyl G et al 2016 Sampling FEE and Trigger-less DAQ for the J-PET Scanner Acta Phys. Pol. B 47 491
Korcyl G et al 2018 Evaluation of single chip real time tomographic data processing on FPGA SoC devices IEEE Trans. Med. Imaging
37 2526–35
Kowalski P et al 2015 Multiple scattering and accidental coincidences in the J-PET detector simulated using GATE package Acta Phys. Pol. A
127 1505–12
Kowalski P et al 2016 Scatter fraction of the J-PET tomography scanner Acta Phys. Pol. B 47 549
Kowalski P et al 2018 Estimating the nema characteristics of the J-PET tomograph using the GATE package Phys. Med. Biol. 63 165008
Krajewski S et al 2013 Cyclotron production of 44Sc for clinical application Radiochim. Acta Int. J. Chem. Asp. Nucl. Sci. Technol. 101 333–8
Krzemień W et al 2015 Analysis framework for the J-PET scanner Acta Phys. Pol. A 127 1491–4
Kubicz E et al 2015 Studies of unicellular micro-organisms Saccharomyces cerevisiae by means of positron annihilation lifetime
spectroscopy Nukleonika 60 749–53
13

---
Lang C, Habs D, Parodi K and Thirolf P G 2014 Sub-millimeter nuclear medical imaging with high sensitivity in positron emission
tomography using
 coincidences J. Instrum. 9 P01008
β
γ
β+
Lang C et al 2012 Submillimeter nuclear medical imaging with a Compton Camera using triple coincidences of collinear
 annihilation
photons and
-rays (arXiv:1202.0397)
γ
Lecoq P 2017 Pushing the limits in time-of-flight pet imaging IEEE Trans. Radiat. Plasma Med. Sci. 1 473–85
Lin H H, Chuang K S, Chen S Y and Jan M L 2016 Recovering the triple coincidence of non-pure positron emitters in preclinical pet Phys.
Med. Biol. 61 1904
Liu G, Chen H, Chakka L, Cheng M L, Gadzia J E, Suzuki R, Ohdaira T, Oshima N and Jean Y 2008 Further search for selectivity of positron
annihilation in the skin and cancerous systems Appl. Surf. Sci. 255 115–8 (Proc. of the 11th Int. Workshop on Slow Positron Beam
Techniques for Solids and Surfaces)
Liu G, Chen H, Chakka L, Gadzia J E and Jean Y C 2007 Applications of positron annihilation to dermatology and skin cancer Phys. Status
Solidi c 4 3912–5
Moskal P et al 2013 TOF-PET tomograph and a method of imaging using a TOF-PET tomograph, based on a probability of production and
lifetime of a positronium US Patent 9851456 (2017), PL 227658 (2013)
Moskal P et al 2014 Test of a single module of the J-PET scanner based on plastic scintillators Nucl. Instrum. Methods A 764 317–21
Moskal P et al 2015 A novel method for the line-of-response and time-of-flight reconstruction in TOF-PET detectors based on a library of
synchronized model signals Nucl. Instrum. Methods A 775 54–62
Moskal P et al 2016 Time resolution of the plastic scintillator strips with matrix photomultiplier readout for J-PET tomograph Phys. Med.
Biol. 61 2025
National Nuclear Data Center 2016 www.nndc.bnl.gov/ (Accessed: 20 January 2016)
NEMA 2012 ‘Performance measurements of positron emission tomographs’ National Electrical Manufacturers Association (NEMA NU
2-2012)
Nemallapudi M V, Gundacker S, Lecoq P, Auffray E, Ferri A, Gola A and Piemonte C 2015 Sub-100 ps coincidence time resolution for
positron emission tomography with lso:ce codoped with ca Phys. Med. Biol. 60 4635
Niedźwiecki S et al 2017 J-PET: a new technology for the whole-body PET imaging Acta Phys. Pol. B 48 1567
NUCLEIDE-LARA—Library for Alpha, X and Gamma Emissions 2017 www.nucleide.org/NucData.htm (Accessed: 10 February 2018)
Oger T et al 2012 A liquid xenon TPC for a medical imaging Compton telescope Nucl. Instrum. Methods A 695 125–8
Pałka M et al 2017 Multichannel FPGA based MVT system for high precision time (20 ps RMS) and charge measurement J. Instrum.
12 P08001
Pawlik-Niedźwiecka M et al 2017 Preliminary studies of J-PET detector spatial resolution Acta Phys. Pol. A 132 1645
Pietrzak R, Borbulak S and Szatanik R 2013 Influence of neoplastic therapy on the investigated blood using positron annihilation lifetime
spectroscopy Nukleonika 58 199–202
Pizzichemi M et al 2016 A new method for depth of interaction determination in PET detectors Phys. Med. Biol. 61 4679.
Schaart D R et al 2010 LaBr3:Ce and SiPMs for time-of-flight PET: achieving 100 ps coincidence resolving time Phys. Med. Biol. 55 N179
Schrader D M and Jean Y C 1988 Positron and Positronium Chemistry (Amsterdam: Elsevier)
Shopa R Y et al 2017 Three-dimensional image reconstruction in J-PET using filtered back-projection method Acta Phys. Pol. B 48 1757
Singh A et al 2017 First-in-human PET/CT imaging of metastatic neuroendocrine neoplasms with cyclotron-produced 44sc-dotatoc: a
proof-of-concept study Cancer Biotherapy Radiopharmaceuticals 32 124–32
Slomka P J, Pan T and Germano G 2016 Recent advances and future progress in PET instrumentation Semin. Nucl. Med. 46 5–19 (recent
advances and future perspectives in nuclear medicine)
Stepanov S, Zvezhinskiy D, Duplâtre G, Byakov V, Batskikh Y Yu and Stepanov P 2010 Incorporation of the magnetic quenching effect into
the blob model of ps formation. Finite sized ps in a potential well Materials Science Forum 666 109–14
Szkliniarz K et al 2015 Medical radioisotopes produced using the alpha particle beam from the warsaw heavy ion cyclotron Acta Phys. Pol. A
127 1471
Tao S J 1972 Positronium annihilation in molecular substances J. Chem. Phys. 56 5499–510
Thielemans K, Tsoumpas C, Mustafovic S, Beisel T, Aguiar P, Dikaios N and Jacobson M W 2012 STIR: software for tomographic image
reconstruction release 2 Phys. Med. Biol. 57 867
Thirolf P G, Lang C and Parodi K 2015 Perspectives for highly-sensitive PET-based medical imaging using
 coincidences Acta Phys. Pol.
β
γ
A 127 1441–4
Townsend D W 2008 Multimodality imaging of structure and function Phys. Med. Biol. 53 R1–39
Vallery R S, Zitzewitz P W and Gidley D W 2003 Resolution of the orthopositronium-lifetime puzzle Phys. Rev. Lett. 90 203402
van Dam H T et al 2011 A practical method for depth of interaction determination in monolithic scintillator PET detectors Phys. Med. Biol.
56 4135
Vandenberghe S, Mikhaylova E, D’Hoe E, Mollet P and Karp J S 2016 Recent developments in time-of-flight PET EJNMMI Phys. 3 3
Viswanath V et al 2017 Development of PET for total-body imaging Acta Phys. Pol. B 48 1555
Walczak R et al 2015 Cyclotron production of 43Sc for pet imaging EJNMMI Phys. 2 33
Yan J, Schaefferkoetter J, Conti M and Townsend D 2016 A method to assess image quality for low-dose PET: analysis of SNR, CNR, bias and
image noise Cancer Imaging 16 26
Yas R M, Al-Mshhdani A H, Elias M M and Al-Sheibani Z T 2012 Detection of line shape parameters in normal and abnormal biological
tissues Iraqi J. Phys. 10 77–82
Zhang X, Zhou J, Cherry S R, Badawi R D and Qi J 2017 Quantitative image reconstruction for total-body pet imaging using the 2-meter
long explorer scanner Phys. Med. Biol. 62 2465
14