Signature:©PolJRadiol,2009;74(4):69-76
REVIEWARTICLE
OptimisationoftheCTparameterswithevaluationof
Received:2009.07.17
Accepted:2009.10.20
MDCTdouble-scanimagesintheplanningofthedentalimplanttreatment
WadimWojciechowski1,AndrzejUrbanik1,PatrykKownacki2,StanisławKownacki2,AgataSowa3
1DepartmentofRadiology,CollegiumMedicum,JagiellonianUniversity,Cracow,Poland2DentalClinic"Kowdent",Cracow,Poland
3AGHUniversityofScienceandTechnology,FacultyofPhysicsandAppliedScience,DepartmentofMedicalPhisicsand
Biophysics,Cracow,Poland
Author’saddress:WadimWojciechowski,RadiologyDepartment,CollegiumMedicum,JagiellonianUniversity,Kopernika19Str.,31-501Cracow,Poland,e-mail:wwadik@poczta.onet.pl
Summary
TheaimofthepresentstudywasoptimisationoftheexaminationparametersandevaluationofreliabilityoftheMDCTdouble-scanimagesobtainedwithcomputernavigationfordentalimplanttreatment.
WiththeuseofaMDCTscannerSOMATOMSensation(Siemens),CT-imagesofaphantomwereperformed:slice-collimation(10×0.75mm,10×1.5mm),slice-thickness(0.75,1,2,3,5mm),pitch(0.5,1,1.5).Additionally,theanalysisonvariousfiltersfromH20ftoH60fwasperformed.Forstudyusedaphantomofthehumancadaverhead
QualitativeanalysiswasdoneusingNobelGuide(NobelBiocare,Sweden),assessingpossibleartefactsontheimages,andmeasurementsofthebonestructureonallfiltersincomparisonwiththerealimage.
Thequalityofthephantomimageswasassessedasoptimalfortheslicesthickness0.75and1mm.Theuseofvariousvaluesofthepitchdidnothavestatisticallysignificantdifferenceontheimagequality.Applicationofvariousfiltersdidnotaltertheparametersofthebonestructure,howevertheuseoflowerfilters(H30fandH40f)hadabeneficialeffectonthequalityof3Dreconstruction.Thearrangementofthe“window”parametersinCTseemedtohaveagreaterinfluenceonthemeasurementandevaluationofthebonestructure.
Slice-collimationandslice-thicknessarethemostimportantparametersinselectionoftheoptimalscan-protocol.Itisrecommendedtouseinthepostprocessing,thementionedaboveparametersuccessionwiththeapplicationofvariousfilters(H30fandH60f)atastablearrangementofthe“window”intheCTexamination.
Background:
Material/Methods:
Results:
Conclusions:
Keywords:
computersimulation•dentalimplants•humancadaver•three-dimensional•computer-assisted•computedtomography
PDFfile:
http://www.polradiol.com/fulltxt.php?ICID=900385
Background
thecreationofsystemsofcomputer-aideddetectionthatmaketheevaluationofimagescarriedoutbyradiologistseasierandmoreprecise,byrevealingpathologicallesions.ThewidestapplicationofCADcanbefoundinmam-mography.However,thesesystemsarealsobecoming
ThedevelopmentofthesoftwareusedforCTdatapro-cessingallowedforintroductionofnewsolutionsinthediagnosticevaluationofimages.Animportantstepwas
69

ReviewArticle
©PolJRadiol,2009;74(4):69-76
increasinglypopularincomputedtomography,magneticresonanceimaginganddentalexaminagions[1,2].Theymakethediagnosismoreobjective.
ModerntechniquesofCTimagingwieldenormousinflu-enceonthedevelopmentofimplantology.Atpresent,everyprocedureofdentalimplantationisfullybasedondataobtainedinthecourseofCT.Thisallowsforapreciseeval-uationoftheosseousstructureinlocationoftheplannedprocedure.CTprovidesuswithallpossiblemeasurementsofthebone,sothatitispossibletousethelongestimplants,guaranteeingthebeststabilisation[3].Thesafetyoftheproceduregetssignificantlyincreasedwhenwelocalisesuchstructuresasthesinuses,nervesandvessels(theircourse)andwhenwedetectpossiblepathologies.Forthelastfewyears,thesystemsofcomputer-plannedimplanta-tionhavebeenthebasisofimplantology.Theychangedthewayofprocedureperformanceandintroducedmanynewsolutions.ThefirstsystemofthatkindwasNobelGuide™,presentedin2003byaSwedishcompany,NobelBiocare.Simulationoftheimplantprocedureintheprogrambasesona3Dreconstructionofthemaxillaorthemandible[4].Afteranexactanalysisanddeterminationofanoptimallocationforanimplantwithinthebone,specialsurgicaltemplatesareprepared,ensuringtheprecisionofthepro-cedureandminimisingpotentialsurgicalmistakes.Systemsofvirtualimplantationareconnectedwithanintroductionoftechniquesofimmediateloading.Theclassicprocedureinvolvestwotreatmentstages:first,thereistheintroduc-tionoftheintraosseouselement,andthen,aftersometimerequiredforaproperosteointegration,proceedsattachmentoftheexternalpartandfixationoftheprostheticimplant.Owingtothesystemofvirtualplanning,theprostheticimplantisdesignedbeforetheprocedure.Withnocontrain-dications,theimplantisfixedrightafteritsplacement.Thistechniqueisabreakthroughinimplantology,asitshortensthetreatmenttime,reducesthedegreeofinvasivenessandisaestheticallyadvantageousforthepatients[5].
Figure1.Double-scanCT.(A)phantomwitharadiologicaltemplateinthemouth,(B)radiologicaltemplate.
obtainedintheMDCTdouble-scantechnique,intheplan-ningofdentalimplantationprocedures.
MaterialandMethods
Studymaterialandequipment
CTexaminationswereconductedwiththeuseofahumancadaverhead,inthedouble-scantechnique,witha10-rowSomatomSensation10scanner(Siemens).
Thefollowingconstantparameterswereused:
- Lampvoltage–120kV,
- Lampcurrentintensity–80mAs,
- Lamprotationtime–0.5sec.
Double-scantechnique
Thistechniqueinvolvestwoacquisitionsofthemaxillaandthemandibleinaspecificway(Figure1).
Incaseofpatients,thefirststageinvolvesCTvisualisationofthephantomwithradiographictemplateinthemouth.Inthesecondstage,theradiographictemplateisscannedalone.Thepreconditionofacorrectperformanceofthesecond-phasescanningisaproperplacementoftheradio-graphictemplateinthegantries,exactlyinthesameposi-tionasitwasseeninthemouthoftheexaminedperson.Thisrequires application of a pad made of material absorb-ingx-rays,e.g.polystyrenefoam.Scanningoftheradio-graphictemplateusesthesameparametersasinthefirststageofexamination.
Moderntechniquesofimmediateloading,basedonthree-dimensionalimagesofthemaxillaorthemandible,enabletheperformanceofavirtualprocedurewiththeuseofspecialisedprograms.Computerplanningoftheprocedureinvolves:quantitativeandqualitativeevaluationoftheosseousstructures,determinationofthecourseofnervesandvesselsandlocationofthesinuses.Thisanalysisallowsfortheestablishmentofanoptimalintraosseousimplantlocation,whichincreasestheoddsofacorrectosteointegra-tionandenhancesthesafetyandprecisionoftheprocedure.
Inthisstudy,thetemplatewasplacedinthephantom’smouth.
Conductedexaminationsandanalysismethods
Inordertoexaminetheinfluenceofsomeselectedacquisi-tionparametersonthequalityofobtainedCTimages,weanalysed:slicethickness,pitch,reconstructioninterval,and collimation.Different filters,used inthe processof sec-ondaryreconstruction,andtheirinfluenceonimagequal-ity,wasalsoevaluated.Thisrequiredsocalleddigitalpro-cessingofthedata,performedwithadiagnosticconsoleofCT,afterthescanningstage.Digitalprocessingletusobtainsecondaryreconstructionsoftheimages,forthefollowingvaluesoftheexaminedparameters:
•Slicethickness:0.75,1,1.5,2,3,4i5mm,
•Pitch:0.5,1and1.5,
- Reconstructioninterval:50%and100%,
- Collimation:10x0.75mmand10x1.5mm,
- Filters:H20f,H30f,H40f,H50fandH60f.
Virtualplanningbasesona3DprocessingofthedataobtainedduringCT.Properperformanceof theCTexamina-tionisessentialforareliableanalysisandafullapplicationofthepossibilitiesoftheplanningsystem.Thefulfilmentoftheaforementionedconditionsdependsontherightchoiceoftheexaminationparameters,whichwillresultintheproductionofimagesofthehighestquality,withthepossi-blylowestrateofradiationdoseatthesametime[6].
Aimofthework
Theaimoftheworkwastooptimisetheexaminationparametersandtoevaluatethereliabilityoftheimages
70
A	B

©PolJRadiol,2009;74(4):69-76
WojciechowskiWetal–OptimisationoftheCTparameterswithevaluation…
Figure2.Examplesoftheimagesratedas(fromtheleft):optimal,good,poor;atpitch
=0.5and(A)slicethickness=0.75mm,overlap50%,(B)slicethickness2mm,overlap50%,(C)slicethickness=2mm,overlap=100%.
Table1.Qualitativeanalysis.Number ofoptimal,good,and poorimages,depending onthepresenceofgeometrical distortions andartefacts.Theresultswereaddedupforallpitchvalues,reconstructionintervalsof50%and100%.Theappliedcollimationamountedto10×0.75mm.
Slicethickness[mm]
Geometricaldistortion
Artefacts
Finalevaluationresult

Optimal
Good
Poor
Optimal
Good
Poor
Optimal
Good
Poor
0.75
6
2
4
8
4
1
3
3
3
3
6
6
1.5
1
5
2
2
2
3
7
2
3
3
2
1
3
2
4
6
3
1
5
6
1
11
4
6
6
12
5
6
6
12
Weobtainedimagesatallpossiblecombinationsof:slicethickness,pitch,reconstructioninterval,andcollimation.AllthesereconstructionsappliedbonefilterH60f.Intwoselectedimageswiththebestparameters,allpossiblefil-ters wereused: fromH20ftoH60f.As aresult,we obtained76protocols.Next,a3Dreconstructionofthedatawascarriedoutwiththeuseoftwoalgorithms:SSD-surfaceshadedisplay,andVR–volumerendering.
Poorimage–imageoflowquality;visibledistortionsandartefactsofsubstantialdegree.
Theevaluationwasperformedbytwoindependentobserv-ers.Next,theexaminationresultsweremergedandobjec-tifiedona jointsession.
ThestatisticalanalysiswascarriedoutwiththeuseoftheFishertest.
Datapresentedinthisway,intheDICOMformat,wererecordedonCD-ROMsandtransferredonPCs.
Results
Influenceof theslice thickness onthe qualityof CTimages
Thequalitativeanalysisapplied3DreconstructionsofthedataandwascarriedoutintheeFilmprogram,Workstation1.5.3,MERGEHealthcarecompany.Thispro-gramallowsforimagereview,comparison,enlargementofsomechosenstructures,andperformanceofmeasure-ments,inordertoestablishpossibleimagedistortions.
JuxtapositionoftheresultsofphantomimagingwiththenumberofgeometricaldistortionsandartefactswaspresentedinTable1.Scansofthehighestqualitywereobtainedwhentheslicesmeasured0.75and1mm.Theslicethicknessof1.5mmwasalsoevaluatedasgood.Whencomparedtotheslicethicknessof1mm,someminorartefactsonthesurfaceoftheevaluatedstructuresinthe1.5-milimitersliceswererevealed.However,theywerenotsignificant.Weconcludedthat“slicethickness’parameterhasthehighestinfluenceonthequalityoftheobtainedimages.
Weanalysedpotentialgeometricaldistortionsandarte-facts.Geometricaldistortionsweredefinedaschangesofimagedimensions(elongationorshorteningoftheimage),whiletheartefacts–asdistortionsoftheimagesurface(creasesandfolds).
Theevaluationwasbasedonathree-stagescale(Figure2),definedintheworkbyLaghietal.[7]:
- Optimalimage–imageofthehighestquality,sharpout-lines,withoutgeometricaldistortionsorartefacts,
- Goodimage–imageworsethantheoptimalone,butwithoutsignificantdeviations.Minordistortionsorarte-factsmaybepresent,
Slicethicknessof2mmisthehighestonethatcanbeusedforthepurposesof3Dreconstructions.Thecompari-sonoftheresultsobtainedwiththebestslices(0.75mm,1mm)andwiththe2-milimeterslices,confirmedbyanexactFishertest(P=0.007),pointstoastrongcorrelationbetweenthequalityoftheimagesandslicethickness.This
71
A	B	C

ReviewArticle
©PolJRadiol,2009;74(4):69-76
Table2.Comparisonofthequalityofimagesperformedwithreconstructionintervalequalto50%and100%,withslicethicknessof2mm.
Geometricaldistortion
Artefacts
Reconstructioninterval
Optimal
Good
Poor
Optimal
Good
Poor
50%
3
2
1
100%
3
3
Table3.Qualitativeanalysis.Numberofoptimal,good,andpoorresults,withcollimationof10×0.75andreconstructionintervalof50%and100%.
Reconstructioninterval50%and100%
Geometricaldistortion
Artefacts
Slicethickness[mm]
Optimal
Good
Poor
Optimal
Good
Poor
50%
100%
50%
100%
50%
100%
50%
100%
50%
100%
50%
100%
0.75
3
3
1
1
2
2
1
3
3
3
3
1.5
1
2
3
2
1
1
2
2
3
3
2
1
3
3
1
2
3
3
3
4
3
3
3
3
5
3
3
3
3
Figure3.Imagesatslicethicknessamountingto2mm,pitch0.5,reconstructionintervalof(A)50%,(B)100%.
Influenceofthereconstructionintervalonimagequality
meansthatthe2-milimitersliceshouldnotbeusedforthecreationof3Dimagesandforthepurposesofthetreatmentprocedure.Theanalysisofallimagesobtainedatthisslicethicknessshowedthattheirqualityishighlydependentonotherparameters,especiallythereconstructioninterval.
Table3presentsthejuxtapositionofthephantomimagingresultswithrespecttothenumberofgeometricaldistortionsandartefacts,atcollimation10×0.75mm,atthreevaluesofpitchandatreconstructionintervalof50%and100%.
Thesine qua noncondition for theuse of 2-milimeterslicesistheapplicationofthereconstructionintervalof50%,atwhichsuchsliceswereevaluatedasgood.However,averyspecificcomparisonofthestructuresinthe2-milimeterslicesandinthinneronesshowedsomeappearingartefactswithintheobservedstructures,aswellasminorgeometri-caldistortions.Althoughtheimagewasgood,theobservedartefactscouldresultinmisrepresentationsandinaccura-ciesinthevirtualpreoperativeplanning.Thatiswhy,theslicesmeasuring0.75mm,1mm,and1.5mmaremuchmoreadvantageous(Table2).
Inthescansperformedatreconstructionintervalof100%,thereweremuchlargerdistortions,foldsanddefectsinbonestructure–eveninthethinnestslices.Whentheslic-eswere2mmthick,theimageperformedwitha100-per-centintervalbecamecompletelyunclear.
Ourmeasurementsconfirmedtherequiredvalueofthereconstructionintervals:50%(Figure3).
Influenceofthepitchvalueonimagequality
Irrespectiveofotherexaminationparameters,slicethick-ness equalto 3,4,and 5mmresults inimagesof averylowquality.Suchimagescannotbeusedintheplanningofden-talimplanttreatment.
Theanalysisoftheinfluenceofthe‘pitch’parameteronimagequalitydidnotconfirmourtheoreticalpredictions.Itwasassumedthatwithanincreasingpitch,i.e.withtableincrementduringlamprotation,andwithdecreasing
72
A	B

©PolJRadiol,2009;74(4):69-76
WojciechowskiWetal–OptimisationoftheCTparameterswithevaluation…
Figure4.Slicethicknessof1mm,reconstructionintervalof50%and(A)pitchof0.5,(B)pitchof1.5.
Table4.Qualitativeanalysis.Thenumber ofoptimal,good,and poorresults,depending on thepresenceof geometrical distortions and artefacts.
Reconstructionintervalof50%,collimationof10×0.75mm.
Pitch0.5
Pitch1
Pitch1.5
Slicethickness[mm]
Optimal
Good
Poor
Optimal
Good
Poor
Optimal
Good
Poor
0.75
2
1
1
1
1
1
2
2
2
1.5
2
2
1
1
2
1
1
1
1
2
3
1
1
2
2
4
2
2
2
5
2
2
2
Table5.Qualitativeanalysis.Thenumberofoptimal,good,and poorimages,depending onthepresenceofgeometricaldistortions andartefacts.
Reconstructionintervalof50%.
Collimation10×0.75mm
Collimation10×1.5mm
Slicethickness[mm]
Optimal
Good
Poor
Optimal
Good
Poor
2
2
4
1
4
1
3
6
6
4
6
6
5
6
6
Influenceofcollimationonimagequality
timeoftheexamination,thequalityoftheimagewoulddecrease.
Thenextstageoftheexaminationswastoanalysetheinfluenceofincreasingcollimationonimagequality.WecomparedtwooptionsoftheCTscanner:collimation10×0.75mmand10×1.5mm.Itwasexpectedthattheuseofthe2-milimiterslicewithcollimation10×1.5mmwouldpro-ducequalitativelysatisfactoryimages(sliceof2mmisthethinnestoneatcollimationequalto10×1.5mm).Whencom-paringtheimagesobtainedwiththesetwocollimations,atslicethicknessof2mm,reconstructionintervalof50%,andatallpitchvalues,itturnedoutthatthechangeincollima-tionvaluedoesnotimprovethequalityofthescans(p=0.5).
Scanningtimeatdifferentpitchvalueswasasfollows:
•Pitch0.5–19.29 sec,
Pitch1–15.35sec,
•Pitch1.5–7.07sec.
However,onthebasisoftheperformedexaminations,itwasconcludedthattheincreasingpitch,withotherparam-etersremainingstable,hadaninsignificantinfluenceonthequalityofthevisualisedstructures.Figure4presentsan imageofa bonestructureat differentpitchvalues.
Imagesatpitch0.5weretheonesthatobtainedthebestscores.Imagesatpitch1and1.5wereevaluatedasworsebutthedecreaseofqualityisinsignificant.Andthus,high-erpitchvaluesmaybeusedwithoutanyrisksofqualityreduction.Thecomparisonofthequalityofimagesatdif-ferentpitchvaluesandslicethicknesseswaspresentedinTable4.
Table5presentsthecomparisonofimagequalityatcol-limation10×0.75mmand10×1.5mm,addedupforallthreepitchvalues.
Thecomparedimagesdidnotdiffermuch,andinsomecases,atcollimationequalto10×1.5mm,theywereevenslightlyworse.
73
A	B

ReviewArticle
©PolJRadiol,2009;74(4):69-76
analysisoftheobtainedresultsshowedthatthekeyparam-eterfortheCTimagequalityistheslicethickness.Yasuhikoetal.[8]analysedtheimpactoftheCTparametersonthequalityofimagesonthebasisofwhichtheimplantationproceduresareplanned.Theexaminationswereperformedwiththeuseofthephantomofthemandible.Influenceofthefollowingparameterswasanalysed:slicethickness,tablespeedandlampcurrentintensity.Itwasshownthatimagequalitydecreaseswithincreasingthicknessoftheslice.Withthesametablespeed,amuchbetterimagequal-itywasobtainedincaseofthinnerslices,1mmthick,thanincaseofthethickerones,2mmthick.Slicethicknessequalto2mmwasrejectedbytheauthorsasinsufficientandunsatisfactoryforthepreoperativeanalysis.
Figure5.Influenceoffiltersonthequalityofvisualisedslice:(A)H30f,(B)H60f.
Theinfluenceofthefiltersonimagequality
Thelaststageoftheanalysiswasthecomparisonoftheimpactofdifferentfiltersonthequalityofobtainedimag-es.Wechose twoscans, with the followingparameters:
1. Slicethickness:0.75mm,pitch:1.5,reconstructioninterval: 50%, collimation:10×0.75 mm,
2. Slicethickness:2mm,pitch:1.5,reconstructioninter-val:50%,collimation:10×1.5mm.
Laricietal.[9]cametosimilarconclusionswhenevalu-atingtheimpactof,i.a.,slicethicknessinthedetermina-tionofthesizeofnoduleswithinthelungs,onthebasisofthree-dimensionalreconstructionsofimagesofaspe-ciallymadechestphantom.Twovaluesofslicethicknesswerecompared:1.25mmand2.5mm.Irrespectiveofotherparameters,onlyatslicethicknessof1.25mmalllesionswereproperlydistinguishedbytheprogramanalysingtheobtaineddata.
Reconstructionswereperformedwiththeuseoffilters,i.e.H20f,H30f,H40f,H50f,andH60f.FilterH60fwastreatedbythesoftwareproducerasareferentialparameter,andthusotherfilterswerecomparedtoit.
Loubeleetal.[10],whendevelopingtheprotocolthatwoulddecreasepatient’sradiationdose,appliedtheslicethicknessof0.75mminalltrials.Althoughthickersliceswouldallowforthereductionofthedose,theauthorsdidnotdecidetoincreasethisparameter,duetotheneedofaveryhighscanningprecisionintheplanningofthedentalimplantprocedures.
Atahighfilter,H60f,theobtainedtransverseimageswereverysharp,withmanydistinctlydelineateddetails(Figure5).
Whenperformingthe3Dreconstructions,weconcludedthatfilterH60fproducesanartificialsharpeningofana-tomicaldetailswithinbones.
Gahleitneretal.[11],intheirreviewpresenting(amongothers)thetechnique of CTexaminationsindentistry,whentryingtodescribethecriteriaof choiceofaproperprotocol,pointed out that the slicethicknesscannotexceed1.5mm.
Therefore, the authors decidedthatinordertocreatethree-dimensionalreconstructionsnecessaryforvirtualplanning,itismuchmoreadvantageoustouselowerfil-ters,suchasH30forH40f,whichsmoothenthestructures,andremoveinsignificantelementsthatmaketheimagemoreunclear(Figure6).
Apartfromtheabovementionedstudies,therewerealsomanyarticlesonvirtualendoscopythatunderscoredthesignificanceofacorrectlyselectedthicknessofthesliceandtheevaluationof3Dreconstructiveoptions.
Discussion
Introductionofthemethodsofcomputerimageanalysisandgraphicalprogramsenabling3DreconstructionsofdifferentstructuresexaminedwithCT,creatednewoptionsofimageprocessingandinterpretation.Appropriatesystemsenablethecreationof(interalia)detailedplansofimplanttreat-ment,socalledvirtualimplantation,andvirtualendoscopyofmanydifferentorgans:virtualbronchoscopy,gastrosco-py,colonoscopy,angioscopy,cystoscopy,ureteroscopy,andendoscopyofparanasalsinusesandvertebralcanal.Withtheintroductionoftechniquesofdiagnosticinformationacquisitiononthebasisof3DCTreconstructions,therehasappearedaneedfordeterminationofimagereliability.The
Lingetal.[12],intheirstudyonvirtualcolonography,evaluatedtheinfluenceof:thesizeofsomeprotrudingelementswithinthephantomofthelargeintestine,posi-tioningofthephantom,thicknessoftheappliedslices,andscannertype(single-ormulti-detectorrow).Itwasfoundthatthelargestelementsofthephantom(10mm),clinicallymostsignificant,werealwaysvisible,inallpositions,andatallpossiblevaluesofslicethickness.Thethickestslicesinthat studyamounted to5mm.
Intheirwork,Laghietal.[6]presentednotonlytheinden-tationswithinthelumenofthephantombutalsothe
Figure6.Three-dimensionalreconstructionswiththeuseoffilters:(A)H30f,(B)H60f.
74
A	B

A	B

©PolJRadiol,2009;74(4):69-76
WojciechowskiWetal–OptimisationoftheCTparameterswithevaluation…
elementssurpassingitsoutlines.Thatishowtheauthorsweretryingtoshowtheimagesofartificialpolypsanddiverticulaofdifferentsizes,presentinthelargeintestine.Bydefiningoptimalparametersofthemulti-detectorrowCT,authorsrevealedagraduallydecreasingscannersensi-tivitywithanincreasingthicknessoftheslices.Theworstresultsinthequotedworkwereobtainedatslicethicknessequalto 5 mm, and thebest ones at 1 mm.
Inclinicalcases,whenweneedtoevaluatepreciselythestructureofthebones,itisrecommended(accordingtoourresults)tousethecollimationof10×0.75mmandthelow-estvaluesofslicethickness(0.75and1mm),withthepitchvalueequalto1.5,althoughtheparameterof1mmwouldbesufficientnottolosethequalityoftheimage.
SimilarconclusionsweremadebyLackneretal.[13],intheirworkevaluatingtheimpactoftheacquisitionparam-etersonthedistinguishabilityofstructuresofdifferentdensities,duringthedetectionofintraocularforeignbod-ies.Theauthorsfoundoutthatthedecreaseinthecollima-tionvalue,from16×1.5mmto16×0.75mm,resultedinatwo-foldimprovementintheimage–noiseratio.
Similarresultswereobtainedinourwork,wherethehigh-estnumberofoptimalobservationswasfoundwithslices
0.75mmand1mmthick.However,afteradetailedanaly-sisoftheseparameters,nodifferenceinimagequalitywasrevealedwiththeslices0.75mmand1mmthick.Duetothatfact,thisslightlyhighervalueofslicethicknessmaybeusedwithoutanyqualityreduction.Thatiswhy,therecommendedandsufficientslicethicknessinimplantationplanningisthatof1mm.
AsimilarpositionwaspresentedbyWesslingetal.[14],aftertheirstudyaimingattheselectionofthebestparam-etersofCTcolonography.Theyusedthephantomofthelargeintestinewithstructurescorrespondingtothepolyps,2–12mminsize.Theyanalysedtheimpactofathinner(4×1mm)andathicker(4×2.5mm)collimation.Thestud-iesrevealedahighernumberofdistortionsandartefactsinthethickercollimation.Polypssmaller than6mm werevis-ibleonlyatcollimationequalto4×1mm.
ResultsofourownexperimentindicatethattheseriesofimagesobtainedinCTexaminationswithstandardproto-cols,settingthethicknessoftheslicesfor3,4,and5mm,areunquestionablyuselessfor3Dreconstructionsandvir-tualplanningofdentalimplantprocedures.Theseprotocolsleadtovisibleartefactsandimagedistortions.Andthus,inmulti-detectorrowtomography,settingtherightthicknessofthesliceisofthehighestimportanceforthechoiceofoptimalexaminationparameters.
Inmulti-detectorrowtomography,thechoiceofsuchvaluesasthereconstructionintervalandpitchisoflowersignifi-cance.Haraetal.[15]definedthereconstructionintervalof50%asoptimal,butpointedoutthatthisvaluewouldbringonahighnumberoftransverseimages,whichextendstheprocessofreconstruction,thetransfer,andthetimeofasec-ondary3Dreconstructiononaworkstation.Thisstudywasperformedinmid-90softhe19thcentury.However,today,inthetimesofaconstantaugmentationofthecomputing powerofthecontemporarycomputersandmodernisationoftheavailablecomputerprogramsforsecondarydataprocessing,ahighernumberofobtainedimagesisnotsuchabigproblem.
Atthispoint,weshouldunderscoreonesignificantdifferencebetweenthesingle-andthemulti-detectorrowCT.In thesin-gle-detectorrowCT,thetermcollimationthicknessandslicethicknessareequivalent–assumingthesamevalues.Inthemulti-detectorrowCT,ontheotherhand,thereiscollima-tionthickness,whichistheproductofrowsandwidthofthedetectors,andthereisslicethickness,whichistheparameterofthereconstructionprocessinfurtherdataprocessing.
Owingtothat,themulti-rowCTprovidesuswithmanymorepossibilities,ascomparedtothesingle-rowCT–thereisapossibilityofamoreelasticprocessingofthe‘raw’dataobtainedintheacquisitionprocess.Moreover,anincreasednumberofrowsandashortenedrotationtime(ofthelamp)reducedthetimeofacquisitionandimprovedthequalityoftheimages.
TheconclusionsofHaraetal.[15]wereconfirmedbyotherauthors,who,intheirpapers,definitelyrecommendedusingthereconstructionintervalequalto50%[10,16].Considerabledifferencesinthequalityof3Dreconstruction,performedwithintervalsequalto50%,75%and100%,wereshownbyShinetal.[17].Theyconfirmedthatthebestresultsareobtainedwiththereconstructionintervalsamountingto50%.
Themostimportantfactorinthemulti-detectorrowCTistherightchoiceofapropercollimationthickness.Inthiswork,themulti-rowCTscannerhadtwocollimationstochose:10×0.75mmand10×15mm.
Hopperetal.[18]pointedtoahigheffectivenessoftheintervalofreconstructionrangingfrom25%to75%.
Theresultsofourstudyoverlapwiththeabovequotedconclusions.Thatiswhy,therecommendedvalueofthereconstructionintervalinthisworkamountsto50%.
Withthefirstcollimationused,weobtainedslices0.75–5mmthick.Withthecollimationof10×1.5mm,thethinnestpossibleslicehad2mm,andthethickest–5mm.Inthissecondoption(10×1.5mm),thequalityoftheobtainedimageswasunsatisfactory,evenifthesliceswere2mmthickandthereconstructionintervalwas50%.Whenweincreasethethicknessoftheslice,thenumberofthepoor-qualityimagesroseaswell.Owingtothat,theseriesofimagesobtainedinCTexaminationsapplyingprotocolswithslicethicknessequalto2–5mmandcollimationof10×1.5areuselessintheplanningofvirtualimplantpro-cedures,inthemulti-detectorrowtechniques.
Whenevaluatingtheimpactofthepitchvalueonimagequality,Neumannetal.[19]suggestedtouse(intheproto-col)thethinnestsliceoptionofascanner,withamaximalpitch.Theyclaimedthatthequalityofthereconstructionwouldimprovethen.
AsimilarconclusionwasmadebyLaghietal.[6]whorevealedstep-likeartefacts,withthepitchvalueincreas-ingto1.5.Atthispitchvalue,theartefactsweremorepronounced,buttheydecreasedwithadecreasingslice
75

ReviewArticle
©PolJRadiol,2009;74(4):69-76
thickness.Wealsofoundthattheslicethicknessof1mmandthereconstructionintervalof50%resultedinforma-tionofimagesthatwerenearlyfreefromdistortionsandartefacts,despitethepitchvaluesof1.5.
ofdentalimplanttreatmentcanbecarriedout,should,duetotheresultsofownexperiment,involvethefollowingparameters:
- Slicethickness:0.75or1mm,
- Reconstructioninterval:50%,
- Pitch:1.5,
- Collimation:10×0.75mm,
- Filter:H30f–for3Dreconstructions,andH60f–fortheevaluationofbonestructureintransversesections.
ObservationspresentedintheworkbyHopperetal.[18]confirmthehighqualityofthescans,irrespectiveoftheappliedpitchvalues(1or1.5).Atthesametime,Schornetal. [20],whenexaminingthe valuesof pitch,ranging from1to3,foundoutthatvalueshigherthan2leadtoanevidentdeteriorationoftheimagequality.
Conclusions
Inourwork,changesinpitchvaluedidnotinfluenceimagequalitysignificantly.Withpitchincrease,onlythethickerslices(3,4,5mm)producedstep-likeartefacts.Withthin-nerslices,thepitchdidnothaveanyimpactonthequalityofthereconstructionbutrevealedasignificantinfluenceontheacquisitiontime.Forexample,examinationtimeofthephantomatpitchequalto0.5amountedto19.29seconds,atpitch=1–15.35seconds,andatpitch=1.5,only7.07seconds.NobelBiocarecompanyrecommendsusingonlythesharpeningbonefilterH60f,whencreatingthethree-dimensionalreconstructionsforvirtualplanningoftheimplanttreatment[6].However,ourstudyresultsrevealedthattheuseoflowerfilters,andespeciallyH30fandH40f,ismuchmoreadvantageousinsubsequent3Dprocessing[5].Smootheningofthestructuresmakestheimagesmoreclearandundisturbedwithsmall,insignificantdetails.
Correct performance oftheexaminationinthe double-scantechnique,withtheuseofproperparameters,resultsinacomprehensiveandapreciseimageofthestudiedstruc-tures.Inordertouseallthepossibleoptionsoftheplan-ningsystemandtoensurereliabilityoftheinformationobtainedwithit,itisnecessarytoperformtheCTscanningprecisely,inthedouble-scantechnique,complyingwiththerecommendationsandchoosinganexaminationprotocolappropriateforagiventypeofscanner.
Themethodofmandibleandmaxillaexamination,withtheuseofthetechniqueofCTvirtualpreoperativeplanning,shouldapplythefollowing,optimalparameters:themostsignificantparameterforimagequalityistheslicethick-ness.Inthisstudy,theoptimalvalueofthisparameterwas
0.75mmand1mm.Thisthicknessmaybeobtainedonlyatcollimationamountingto10×0.75mm.Optimalpitchandreconstructionintervalwereasfollow:1.5and50%,respectively.
Summingup,anoptimalprotocolforCTexaminations,onthebasisofwhich3Dreconstructionsandvirtualplanning
References:
1. SteenbergheD,MalevezC,CleynenbreugelJetal:Accuracyofdrillingguidesfortransferfromthree-dimensionalCT-basedplanningtoplacementofzygomaimplantsinhumancadavers.ClinOralImplantsRes,2003;14(1):131–36
2. VanAN,vanSD,GuerreroMEetal:Accuracyofimplantplacementbasedonpre-surgicalplanningofthree-dimensionalcone-beamimages:apilotstudy.JClinPeriodontol,2007;34(9):816–21
3. CavalcantiMG,RuprechtA,VannierMW:3DvolumerenderingusingmultisliceCTfordentalimplants.DentomaxillofacRadiol,2002;31(4):218–23
4. ŁomżyńskiŁ,Mierzwińska-NastalskaE:NatychmiastoweobciążenieimplantówzzastosowaniemtechnologiiNobelGuide.ProtetStomatol,2006;56(2):136–39
5. WojciechowskiW,KownackiP,KownackiSetal:VirtualplanningofdentalimplantplacementusingCTdouble-scantechnique–ownexperience.PolJRadiol,2007;72(4):44–49
6. NobelBiocare:ConceptRadiologist,Clinicans,LaboratiriesandotherworkingwithNobelGuide.ProceduresandProducts–poweredbyProceraNobelGuide,2005;(1):39–42
7. LaghiA,IannacconeR,MangiapaneFetal:ExperimentalcolonicphantomfortheevaluationoftheoptimalscanningtechniqueforCTcolonographyusingamultidetectorspiralCTequipment.EurRadiol,2003;13(3):459–66
8. YasuhikoM, Eiichi H,HirokoI: Effect of ImagingParametersofSpiralCTScanningonImageQualityfortheDentalImplants.Visualevaluationusingasemi-anthropomorphicmandiblephantom.DentalRadiol,2002;42(4):259–73
9. LariciAR,StortoML,TorgeMetal:AutomatedvolumetryofpulmonarynodulesonmultidetectorCT:influenceofslicethickness,reconstructionalgorithmandtubecurrent.Preliminaryresults.RadiolMed.,2008;113(1):29–42
10. LoubeleM,JacobsR,MaesFetal:Radiationdosevs.imagequalityforlow-doseCTprotocolsoftheheadformaxillofacialsurgeryandoral implantplanning. RadiatProtDosimetry,2005;117(1–3): 211–16
1. GahleitnerA,WatzekG,ImhofH.DentalCT:imagingtechnique,anatomy,andpathologicconditionsofthejaws.EurRadiol,2003;13(2):366–76
2. LingSH,SummersRM,LoewMHetal:Computer-aideddetectionofpolypsinacolonphantom:effectofscanorientation,polypsize,
collimation,anddose.JComputAssistTomogr,2002;26(6):1013–18
1. LacknerB,SchmidingerG,FunovicsM:Diagnosisofintraocularforeignbodiesinmultidetector-CT:influenceofacquisitionparameterson radiation dose and detectability].Rofo,2006; 178(1):90–95
2. WesslingJ,FischbachR,MeierNetal:CTcolonography:
Protocoloptimizationwithmulti-detectorrowCT-studyinananthropomorphiccolonphantom.Radiology,2003;228(3):753–59
1. HaraAK, Johnson CD,ReedJE et al:Reducingdata size andradiationdoseforCTcolonography.AJRAmJRoentgenol,1997;168(5):1181–84
2. KasalesCJ,HopperKD,AriolaDNetal:ReconstructedhelicalCTscans:improvementinz-axisresolutioncomparedwithoverlappedandnonoverlappedconventionalCTscans.AJRAmJRoentgenol,1995;164(5):1281–84
3. ShinJH,LeeHK,ChoiCGetal:Thequalityofreconstructed3Dimagesinmultidetector-rowhelicalCT:experimentalstudyinvolvingscanparameters.KoreanJRadiol,2002;3(1):49–56
4. HopperKD,PierantozziD,PotokPSetal:Thequalityof3Dreconstructionsfrom1.0and1.5pitchhelicalandconventionalCT.JComputAssistTomogr,1996;20(5):841–47
5. NeumannK,WintererJ,KimmigMetal:Real-timeinteractivevirtualendoscopyofthetracheo-bronchialsystem:influenceofCTimagingprotocolsandobserverability.EurJRadiol,2000;33(1):50–54
6. SchornC,VisserH,HermannKPetal:DentalCT:imagequalityandradiationexposureinrelationtoscanparameters].Rofo,1999;170(2):137–44
76