# DETERMINATION OF SELENIUM IN HUMAN BLOOD USING ATOMIC FLUORESCENCE SPECTROMETRY
Renata WIETECHA1, Paweł KOŚCIELNIAK1-2, Teresa LECH2,
Maciej RYMANOWSKI1
1. Faculty of Chemistry, Jagiellonian University, Cracow
2. Institute of Forensic Research, Cracow
ABSTRACT: In the present publication, a method developed for the determination
of selenium in human blood is described. A blood sample was completely digested by
nitric acid with use of a microwave oven. The analytical signal for selenium was mea-
sured by means of atomic fluorescence spectroscopy with the hydride generation sys-
tem (AFS-HG). The method appeared to be typified by a detection limit of ca. 0.1 pgl-1
and precision of ca. 1.0% (RSD). The accuracy of the analytical results was evaluated
at a level of ca. 2% (RE). It was verified with the use of reference materials of whole
human blood and human serum. The procedure is very simple, fast and reliable,
hence it is suggested to be applied to routine clinical and toxicological analyses.
KEYWORDS: Selenium; Analysis of blood; Atomic fluorescence spectrometry; Hy-
dride generation; Microwave digestion.
Problems of Forensic Sciences, vol. LII, 2002, 21-36
Received 25 October 2002; accepted 30 December 2002
INTRODUCTION
Selenium has aroused increasing interest in recent years — it has become
the subject of special attention for persons connected with various branches
of industry and environmental protection, and also persons connected with
medicine, biochemistry and toxicology.
For many years selenium was considered solely as an element that was
toxic and carcinogenic for people and animals. Today, it is known to be neces-
sary for the normal functioning of various living organisms: bacteria, plants,
animals and people [2], in spite of its toxicity. In the literature one can find
many examples of therapeutic applications of selenium, e.g. as a treatment
for atherosclerosis, in cases ofcardiac anoxia and diseases which are charac-
terised by development of inflammation and depression. Moreover, sele-
nium is an inactivator of toxic heavy metals, such us mercury, cadmium or
silver [15] and organic compounds liberated during infections, injuries and
stresses. Selenium and its compounds reveal antibacterial and fungicidal
activity, which is made use of in the production of cosmetics for skin and

22
R. Wietecha, P. Koscielniak, T. Lech, M. Rymanowski
hair. Furthermore, it neutralises the activity of aflatoxins, which are known
for their carcinogenic influence on the organism [10,14,19].
In contrast to other bio-elements, selenium is characterised by a narrow
range between its therapeutic and toxic doses. For this reason it is named
“the most toxic element necessary for life” [1]. Because selenium is used in
industry, its compounds can pass through to the environment and constitute
a danger to health and life. The most frequent cases of poisonings occur
among people working in copper smelting [9]. Cases of suicide by the con-
sumption of technical preparations containing large amounts of selenium
compounds are also known [11]. Moreover, incorrect supplementation by se-
lenium preparations may be the cause of chronic poisonings. The toxic dose
for man is 700 pg per day [1].
Therefore, a reliable determination of selenium in various tissues is very
important from a clinical and toxicological point of view.
Analysis of selenium in biological materials is a relatively difficult prob-
lem, which has not to this day been entirely satisfactorily resolved. The com-
plete analytical procedure of selenium determination in this type of material
consists in two principle steps: mineralisation of the sample and measure-
ment of analytical signals using a suitable instrument. The analytical chem-
ist encounters significant problems at both stages.
The main problems of the mineralisation stage are the low speed of this
process and the risk of occurrence of selenium losses from a sample because
of its volatility. Both these effects are especially important in the case of
“wet” mineralisation in the open system — a universally used method. Vari-
ous mixtures of acids (HNO3, H2SO4, HClO4, HF etc.) and peroxide (H2O2) are
usually used for this aim, and the process itself is performed under normal
air pressure and temperature. The above difficulties can to a large extent be
avoided by applying closed-systems that use microwave energy, which en-
able total decomposition of the sample matrix [12].
Various instrumental methods are applied to the measuring ofanalytical
signals of selenium. The oldest one, commonly used to this day, is spectro-
photometry. In this case, quantitative determination of selenium is usually
realised by measurement of absorbance of coloured complex combinations of
selenium (IV) with aromatic o-diamines, i.e. piazoselenoles. The detection
limit (0.003-0.3 mg l-1) of the spectrophotometric method is often insuffi-
cient for determination of trace amounts of selenium in biological mate-
rial [7]. A more sensitive method is spectrofluorometry, where measurement
of molecular fluorescence at a suitable wavelength is performed after a pre-
liminary reduction of selenium and a complexing reaction. However, the
photo-chemical reaction used in this method requires special conditions
(a dark room or yellow light). Moreover, both methods are quite complicated
from the chemical point of view, which is a cause of the relatively large ran-

Determination of selenium in human blood
23
dom analytical errors [7,17]. These drawbacks are also characteristic for the
voltamperometric method of selenium determination [4].
Other analytical methods, such as activation analysis and X-ray fluores-
cence are only used sporadically for determination of selenium, because,
above all, of the large costs of equipment and of the analysis performed [8].
The most popular method of selenium determination - beside spectro-
photometry — is atomic absorption spectrometry coupled with hydride genera-
tion (AAS-HG). It is characterised by high speed and simplicity of analysis, as
well as a low detection limit and good precision of analytical results. A draw-
back of this method, which is especially significant when samples with a com-
plex matrix are analysed (such as biological materials), is the problem of back-
ground, which cannot always be corrected to a sufficient degree.
There has been a noticeable increase in interest in atomic fluorescence
spectrometry (AFS) in recent years. This method coupled with the technique
of hydride generation (HG) is used for determination of elements that are
important from the toxicological point of view, such as arsenic, mercury,
lead, and selenium as well. The AFS-HG method is characterised by advan-
tages such as good selectivity, a wide linearity range, a very high sensitivity,
a low noise level which enables achievement of very low detection limits, and
also relatively simple and cheap equipment requirements [6, 16].
In the presented paper, an analytical procedure which consists of a com-
bination of mineralisation in a micro-wave system with a process of sele-
nium hydride generation and measurement of the analytical signal by the
AFS method was applied to selenium determination. This method was
tested on samples of human blood. Research focused mainly on optimising
conditions of mineralisation of blood samples and on defining analytical pa-
rameters of the elaborated method. The accuracy of the method was verified
by analysis of reference human blood samples.
MATERIALS AND METHODS
## Reagents
The following reagents were used during analysis: a standard solution of
selenium of 1 mg ml-1 Se concentration (analytically pure, Merck), conc.
HNO3 (analytically pure, Merck), conc. HCl (analytically pure, POCh), conc.
H3PO4 (analytically pure, POCh), conc. HCOOH (analytically pure, POCh),
30% H2O2 (analytically pure, POCh), 98% NaBH4 (analytically pure, Sigma-
Aldrich Chemie), 0.5% NaOH (analytically pure, Zakłady Chemiczne Oświę-
cim). Deionised water (conductance approx. 1 pS cm-1) was used for prepara-
tion of solutions.

24
R. Wietecha, P. Koscielniak, T. Lech, M. Rymanowski
Samples
Samples of human blood (two series, 50 ml each) taken from volunteer
(unpaid) blood donors from the Regional Centre of Blood Donation and Blood
Health Care, Cracow, were analysed. The blood samples were stored in spe-
cial vials sputtered with anticoagulant (K2EDTA) at —20oC before analysis.
Two reference materials were also analysed: a sample of whole human blood
“Seronom Trace Elements” (LOT 404109, SERO AS, Billingstad) as
a lyophilic solution and a sample of blood serum obtained from the Depart-
ment of Bromathology, Collegium Medicum, Jagiellonian University. The
content of the second sample was determined as part of interlaboratory re-
search entitled “Interlaboratory comparision for copper-zinc-selenium in
blood serum” co-ordinated by the French Society for Clinical Biology,
Vendeuvre Les Nancy.
Analytical equipment
A Mars 5X (CEM, USA) 12-position microwave mineraliser was used for
sample preparations. The process of mineralisation was performed in a high
pressure teflon vessel (type XP-1500) with simultaneous temperature and
pressure control. A double-channel atomic fluorescence spectrometer cou-
pled with a system of hydride generation (AFS-230 Haiguang Co., China)
was applied to the measurement of the analytical signal. A selenium lamp
with a hollow cathode (current — 100 mA) as a source of radiation and
a photo-multiplier (working at 300 V) as a detector were installed in the
spectrometer.
RESULTS AND DISCUSION
Preliminary research consisted in determination of optimal conditions
for the preparation of blood samples for analysis. To this end, the concentra-
tion of selenium was determined in blood samples taken from the same set,
which were subjected to various mineralisation conditions. Establishment
of these conditions was partially based on data gathered from scientific pa-
pers [8, 9]. The composition of the mineralised mixture and the method of
sample treatment after mineralisation are presented in Table I.

Determination of selenium in human blood
25
TABLE I. RESULTS OF SELENIUM DETERMINATION IN BLOOD SAMPLES OBTAINED
UNDER VARIOUS CONDITIONS OF SAMPLE MINERALISATION AND PREPA-
RATION FOR ANALYSIS
Analytical / series	Reagents used for mineralisation	Reagent used after mineralisation	Concentration of selenium / [^g l 1]
I	3.0 ml conc. HNO3 + 0.5 ml 30% H2O2	12.5 ml 6 M HCl	27.3 ±1.2
II	6.0 ml conc. HNO3 + 1.5 ml 30% H2O2	12.5 ml 6 M HCl	29.0 ±2.6
III	7.0 ml conc. HNO3	12.5 ml 6 M HCl	56.6 ±1.2
IV	7.0 ml conc. HNO3	12.5 ml conc. HCl*	40.7 ±0.8
V	1.5 ml conc. HNO3 + 1.0 ml conc. H3PO4	1.5 ml 30% H2O2 + 2 ml conc. HCOOH + 2 ml conc. HCl**	41.0 ±2.3

* The mineralised sample was heated for 10 min at 220oC.
Peroxide was added to the mineralised sample and heated for 25 min at 220oC, formic acid
was added after cooling and the sample was heated for 2 min at 120oC, at the end conc. hydro-
chloric acid was added and the sample was heated for 20 min at 120oC.
Three blood samples and a blind sample (containing reagents) of volume
about 0.5 ml were used in each series for mineralisation. Each sample was
transferred into a teflon vessel and an appropriate volume of mineralisation
mixture was added. After that the process of mineralisation was carried out
in three steps, differing in terms of conditions of maximum temperature and
pressure: a) T = 160oC, P = 130 psi; b) T = 180oC, P = 190 psi; c) T = 200oC,
P = 290 psi. Each step lasted 8 min. After cooling, the mineralised sample
was exposed to hydrochloric acid (without or with addition of formic acid)
with the aim of reduction of Se(VI) to Se(IV), and then it was transferred to
a volumetric flask and filled up to 25 ml. The prepared solution was ready for
introduction into a fluorescence spectrometer.
The generation of selenium hydride took place in a flowing system, after
the sample had been introduced into a solution containing 2% sodium
tetraborane and 0.5% sodium hydroxide. This solution was added at a rate of
6 ml min-1 toa3M solution of hydrochloric acid serving as a carrier, and was
transported to a separator at a rate of 11 ml min-1. After that it was trans-
ported to the argon-hydrogenate flame (temperature: 200oC) of the fluores-
cence spectrometer in a stream of argon at 800 ml min-1.
A series of standard solutions containing an analyte at concentrations
from 0 to 3.48 |Tg l-1 was prepared for calibration purposes. These solutions
were prepared by dilution of a basic standard solution up to a uniform vol-
ume by 6 M hydrochloric acid. Results of selenium determination in blood
samples are presented in Table I.
In further research, analysis of blood samples was performed in condi-
tions which were established in series III of determinations (see Table I).

26
R. Wietecha, P. Koscielniak, T. Lech, M. Rymanowski
This choice was mainly governed by the value of selenium determined in
samples, which was significantly higher in this series of determinations. It
was therefore supposed that the possible loss of selenium caused by the sam-
ple mineralisation process was the smallest under these conditions. More-
over, it is also ofgreat significance that the selected conditions ofmineralisa-
tion were very simple: the use of only one reagent in this case, and prepara-
tion of the mineralised sample for measurement did not require heating, but
only the addition of hydrochloric acid at a suitable concentration. Therefore,
the mineralisation procedure is not only relatively cheap, but also carries
a minimum risk of contamination of sample by reagents added.
The determined limit of selenium detection was about 0.1 p.g l-1 under the
selected analytical conditions and the precision (encompassing the
reproducibility of the whole analytical procedure) was at the 1% level (RSD).
Two calibration methods were applied to selenium determination in
a blood sample with the aim of studying the interference effect: the stan-
dard-series method and the standard addition method. In the case of the
standard addition method, solutions were prepared in the following way:
several portions of a sample were mineralised under the same conditions
and then an analyte, in the form of aqueous standard solution, was added to
the samples. Selenium concentrations in the added standard solutions were
00; 0.63; 1.25 and 1.88 ugl-1. The mutual location of the calibration curves
created by both methods (Figure 1A) indicated the occurrence of the interfer-
ence effect.
Chemical interference in the case of selenium determination in biological
materials by the AAS-HG method after mineralisation by nitric (V) acid has
been reported several times [3, 5, 18].
A hypothesis was put forward [5] that this effect is caused by nitrogen ox-
ides (NO and NO2) remaining in samples after mineralisation, which pre-
vents reduction of Se(VI) to Se(IV). As a consequence, the concentration of
the analyte in the sample introduced into the spectrometer in the form of
Se(IV) was lowered, and because this form has the highest ability to gener-
ate hydrides, the analytical signal did not correspond to the true value of se-
lenium in the sample. One may suppose that, even though another analyti-
cal method (AFS-HG) was applied, the interference effects observed in this
research can be explained in the same way.
In the scientific literature, application of chemical compounds is recom-
mended to eliminate the influence of nitrogen oxides on selenium determi-
nation by mean of the AAS-HG [5, 13, 18]. A new approach, hitherto not ap-
plied by other authors, was tried, which would not involve complicating the
analytical procedure by adding additional reagents to a sample. It was de-
cided to decrease the concentration of nitrogen oxides in the water phase by
removing them from above the samples by a strong stream of argon or nitro-

Determination of selenium in human blood
27
A
Concentration of Se [ g l-1]
B
Concentration of Se [ g l-1]
Fig. 1. Calibration curves calculated for selenium by the standard-series method (a)
and by the method of standard addition (b) without the expelling of nitric oxides from
headspace phase of mineralised sample (A) and after the expelling of nitric oxides in
a stream of argon for five minutes (B).
gen. This procedure was performed directly after mineralisation and the
cooling of samples to room temperature. Samples prepared in this additional
way were analysed by application of two calibration methods. The location of
calculated calibration curves (Figure 1B) indicated that the interference ef-
fect was eliminated by a suitably long action of gases (approx. 10 min).
The effectiveness of reduction of interference effects is illustrated by re-
sults of determination of selenium in a sample of whole human blood, pre-
sented in Table II. Results obtained for reference samples showed that it
was possible to totally eliminate these effects. As can be seen, the applied an-
alytical method allows determination of selenium in blood with accuracy de-
fined by a relative error not greater than 2%.

28
R. Wietecha, P. Koscielniak, T. Lech, M. Rymanowski
TABLE II. RESULTS OF SELENIUM DETERMINATION IN BLOOD SAMPLES WITH UN-
KNOWN SELENIUM CONCENTRATION AND IN REFERENCE SAMPLES
Type of sample	Gas used for the expelling of nitric oxides	Time of expelling [min]	Concentration of Se [ p.l-1]
the expected value	the obtained value
Whole blood - a sample	-	0	?	53.5
Whole blood - a sample	Ar	5	?	56.0
Whole blood - a sample	Ar	10	?	62.7
Whole blood - a sample	N2	10	?	64.4
Whole blood - reference material	Ar	10	79.0	80.5
Whole blood - reference material	n2	10	79.0	78.0
Serum - reference material	Ar	5	35.0	35.5

SUMMARY AND CONCLUSIONS
The analytical method developed for determination of selenium in hu-
man blood involves several experimental steps:
- sample mineralisation in a microwave oven in nitric acid;
- expelling nitric oxides from the headspace of the mineralised sample
  and the transferring of the sample into hydrochloric acid;
- calibration of the determination by the standard series method using
  standard aqueous solutions;
- the generation of hydrides of selenium in the flow system;
- the measurement of analytical signals by means of the atomic fluores-
  cence spectrometer.
The procedure is characterised by the following advantages:
- it is relatively fast and simple;
- it only needs several reagents, therefore it is relatively cheap and safe;
- it allows determination of selenium at concentrations occurring in hu-
  man blood;
- it ensures the gaining of analytical results with very good precision
  and accuracy.
One limitation of this procedure is the necessity of using relatively expen-
sive specialist equipment (a microwave oven) and apparatus that is not eas-
ily available (the AFS spectrometer). The obtained results, however, show
that the proposed method of selenium determination in blood samples is
very competitive in relation to other methods used for the same purpose, and
it can be whole-heartedly recommended for use in toxicological and clinical
laboratories.

Determination of selenium in human blood
29
Acknowledgements:
The authors are sincerely grateful to Dr. Stanisław Walas for his valuable discus-
sions and helpful comments concerning atomic fluorescence spectrometry.
References:
1. Andrzejak R., Groch J., Jurga M., Rola selenu w patofizjologii
   człowieka, Postępy Higieny i Medycyny Doświadczalnej 1996, t. 50, s. 293-307.
2. Barceloux G., Selenium, Clinical Toxicology 1999, vol. 37, pp. 145-172.
3. Brown R., Fry R., Moyers J. [et al.], Interference by volatile nitrogen ox-
   ides and transition-metal catalysis in the preconcentration of arsenic and sele-
   nium as hydrides, Analytical Chemistry 1981, vol. 53, pp. 1560-1566.
4. Bryce D., Izquierdo A., Luque de Castro M., Flow-injection anodic
   stripping voltammetry at a gold electrode for selenium(IV) determination,
   Analytica Chimica Acta 1995, vol. 308, pp. 96-101.
5. Cutter G., Elimination of nitrite interference in the determination of sele-
   nium by hydride generation, Analytica Chimica Acta 1983, vol. 149,
   pp. 391-394.
6. d’Ulivo A., Bramanti E., Lampugnani L. [et al.], Improving the ana-
   lytical performance of hydride generation non-dispersive atomic fluorescence
   spectrometry. Combined effect of additives and optical filters, Spectrochimica
   Acta. Part B 2001, vol. 56, pp. 1893-1907.
7. Gawłowska A., Masłowska J., Spektralne metody oznaczania śladowych
   ilości selenu w wodach naturalnych oraz sokach owocowych, Bromatologia
   i Chemia Toksykologiczna 2000, t. 1, s. 91-97.
8. Jędrzejczak R., Metody oznaczania selenu i arsenu w materiale biolo-
   gicznym, [w:] Arsen i selen w środowisku - problemy ekologiczne i metodyczne,
   Kabata-Pendias A., Szteke B. [red.], Zeszyty Naukowe nr 8, Polska Akademia
   Nauk, Warszawa 1994, s. 194-201.
9. Klewska A., Poziomy fizjologiczne selenu w materiale biologicznym, Archi-
   wum medycyny sądowej i kryminologii 1978, t. IV, s. 307-312.
10. Maehira G., Luyo G. A, Oshiro M., Alterations of serum selenium con-
    centrations in the acute phase of pathological conditions, Clinica Chimica Acta
    2002, vol. 316, pp. 137-146.
11. Matoba R., Kimura H., Uchima E. [etal.], An autopsy case of acute sele-
    nium (selenious acid) poisoning and selenium levels in human tissues, Forensic
    Science International 1986, vol. 31, pp. 87-92.
12. Matusiewicz H., Metody przygotowania próbek w analizie śladowej, [w:]
    Problemy jakości analizy śladowej w badaniach środowiska przyrodniczego,
    Kabata-Pendias A., Szteke B. [red.], Wydawnictwo Edukacyjne Zofii Dob-
    kowskiej, Warszawa 1998.
13. Morales Flores E., Longhi Cirne da Silva L., Barin J. [et al.],
    Minimization of volatile nitrogen oxides interference in the determination of
    arsenic by hydride generation atomic absorption spectrometry, Spectroche-
    mica Acta. Part B 2001, vol. 56, pp. 1883-1891.

30
R. Wietecha, P. Kościelniak, T. Lech, M. Rymanowski
1. Navarro-Alarcon L., Lopez de la Serrana H., Perez-Valero V.
   [et al.], Serum selenium levels as indicators of body status in cancer patients
   and their relationship with other nutritional and biochemical markers,
   The Science of Total Environmental 1998, vol. 212, pp. 195-202.
2. RaymanM., The importance of selenium to human health, The Lancet 2000,
   vol. 356, pp. 233-241.
3. Segura M., Madrid Y., Camara C., Evaluation of atomic fluorescence
   and atomic absorption spectrometric techniques for the determination of ar-
   senic in wine and beer by direct hydride generation sample introduction, Jour-
   nal of Analytical Atomic Spectrometry 1998, vol. 14, pp. 131-135.
4. Watkinson J., Fluorometric determination of selenium in biological mate-
   rial with 2,3-diaminonaphthalene, Analytical Chemistry 1966, vol. 38, pp. 92-97.
5. Wu C., ChenP., YagM., Simultaneous determination of arsenic, mercury,
   antimony and selenium in biological materials with prior collection of gaseous
   products followed by neutron activation analysis, Journal of Radioanalytical
   Nuclear Chemistry 1987, vol. 112, pp. 140.
6. Zachara A., Marchaluk-Wiśniewska E., Maciąg A. [et al.], De-
   creased selenium concentration and glutathione peroxidase activity in blood
   and increase of these parameters in malignant tissue of lung cancer patients,
   Lung 1997, vol. 175, pp. 321-332.

# METODA OZNACZANIA SELENU WE KRWI LUDZKIEJ Z ZASTOSOWANIEM ATOMOWEJ SPEKTROMETRII FLUORESCENCYJNEJ
Renata WIETECHA, Paweł KOŚCIELNIAK, Teresa LECH,
Maciej RYMANOWSKI
WPROWADZENIE
Selen budzi w ostatnich latach coraz większe zainteresowanie — staje się przed-
miotem szczególnej uwagi nie tylko osób związanych z różnymi gałęziami przemysłu
i ochroną środowiska, ale także z medycyną, biochemią i toksykologią.
Bardzo długo selen był uważany za pierwiastek wyłącznie toksyczny i rakotwór-
czy dla ludzi i zwierząt. Dziś wiadomo, że pomimo swej toksyczności jest on bardzo
istotny dla prawidłowego funkcjonowania różnych organizmów żywych: bakterii,
roślin, zwierząt i ludzi [2]. W literaturze przedmiotu znaleźć można wiele przykła-
dów terapeutycznego zastosowania selenu, m.in. do leczenia miażdżycy czy też
w przypadkach niedotlenienia mięśnia sercowego oraz chorób charakteryzujących
się rozwojem stanów zapalnych i depresyjnych. Ponadto selen jest inaktywatorem
toksycznych metali ciężkich, takich jak np. rtęć, kadm czy srebro [15] oraz związków
organicznych uwalnianych w czasie zakażeń, urazów i stresów. Selen i jego związki
wykazują działanie przeciwbakteryjne i grzybobójcze, co jest wykorzystywane w pro-
dukcji preparatów pielęgnujących skórę i włosy. Ponadto neutralizuje działanie
aflatoksyn, które są znane z rakotwórczego wpływu na organizm [10, 14, 19].
W odróżnieniu od innych mikroelementów selen charakteryzuje się tym, że
rozpiętość między jego dawką terapeutyczną a toksyczną jest niewielka. Z tego
powodu nazywany jest „najbardziej toksycznym niezbędnym do życia pierwiast-
kiem” [1]. Ze względu na zastosowanie selenu w przemyśle, jego związki mogą
przedostawać się do środowiska i stanowić zagrożenie dla zdrowia i życia. Naj-
częstsze przypadki zatrucia są spotykane wśród osób pracujących przy wytopie
miedzi [9]. Znane są również przypadki samobójstw w wyniku spożycia preparatów
technicznych zawierających duże ilości związków selenu [11]. Ponadto często nie-
właściwa suplementacja preparatami selenowymi może być przyczynąprzewlekłych
zatruć. Toksyczna dawka selenu dla człowieka wynosi 700 ąg na dobę [1].
Kwestia wiarygodnego oznaczania selenu w różnych tkankach jest zatem bardzo
ważnym zagadnieniem z klinicznego i toksykologicznego punktu widzenia.
Badania analityczne selenu w materiałach biologicznych stanowią stosunkowo
trudny i do tej pory nie do końca rozwiązany problem. Na pełną procedurę analitycz-
ną oznaczania selenu w tego typu materiałach składają się dwa zasadnicze etapy:
mineralizacja próbki i pomiar sygnałów analitycznych przy użyciu odpowiedniego
instrumentu. Na obu tych etapach analityk napotyka spore trudności.
Głównymi problemami etapu mineralizacji jest zarówno mała szybkość tego pro-
cesu, jak i ryzyko występowania strat selenu w próbce na skutek jego dużej lotności.
Z oboma tymi efektami należy się liczyć przede wszystkim w przypadku powszechnie

32
R. Wietecha, P. Kościelniak, T. Lech, M. Rymanowski
stosowanej mineralizacji „na mokro” w układzie otwartym. Zwykle używa się do tego
celu różnych mieszanin kwasów (HNO3, H2SO4, HClO4, HF i inne) oraz nadtlenku
wodoru (H2O2), a sam proces prowadzi się w normalnych warunkach ciśnienia i tem-
peratury. Powyższych trudności można w dużej mierze uniknąć, używając systemów
zamkniętych z wykorzystaniem energii mikrofalowej, które umożliwiają całkowity
rozkład matrycy próbki [12].
Do pomiaru sygnałów analitycznych selenu stosuje się różne metody instrumen-
talne. Najstarszą z nich, powszechnie do dnia dzisiejszego wykorzystywaną, jest
metoda spektrofotometryczna. W tym przypadku ilościowe oznaczanie selenu opiera
się najczęściej na pomiarze absorbancji barwnych kompleksowych połączeń sele-
nu (IV) z o-diaminami aromatycznymi, tzw. piazoselenolami. Granica wykrywalnoś-
ci metody spektrofotometrycznej wynosząca 0,003—0,3 mg l—1 jest często niewystar-
czająca do oznaczania śladowych ilości selenu w materiale biologicznym [7]. Bardziej
czułąmetodąjest metoda spektrofluorymetryczna, gdzie po wstępnej redukcji selenu
i reakcji kompleksowania wykonuje się pomiar fluorescencji molekularnej przy
odpowiedniej długości fali. Jednak wykorzystywana w tej metodzie reakcja foto-
chemiczna wymaga zastosowania specjalnych warunków (ciemne pomieszczenie lub
żółte światło). Obie metody sąponadto dość skomplikowane od strony chemicznej, co
jest przyczyną stosunkowo dużych przypadkowych błędów analitycznych [7, 17].
Podobnymi ujemnymi cechami charakteryzuje się woltamperometryczna metoda
oznaczania selenu [4].
Inne metody analityczne, takie jak np. analiza aktywacyjna i fluorescencja
rentgenowska, stosowane są do oznaczania selenu sporadycznie ze względu — przede
wszystkim — na duży koszt aparatury i samych analiz [8].
Najpowszechniej stosowaną metodą oznaczania selenu — obok metody spektrofo-
tometrycznej — jest metoda atomowej spektrometrii absorpcyjnej w połączeniu z ge-
neracją wodorków (AAS-HG). Charakteryzuje się ona duża szybkością i prostotą
wykonywania analiz a także stosunkowo niską granicą oznaczalności i dobrą precy-
zją wyników analitycznych. Ujemną stroną tej metody ujawniającą się szczególnie
mocno w przypadku analiz materiałów o złożonej matrycy (a więc także materiałów
biologicznych) jest problem tła, nie zawsze możliwego do skorygowania w dostatecz-
nym stopniu.
W ostatnim czasie daje się zauważyć coraz większe zainteresowanie metodą
atomowej spektrometrii fluorescencyjnej (AFS). W połączeniu z techniką generacji
wodorków (HG) stosuje się jądo oznaczenia takich pierwiastków ważnych z toksyko-
logicznego punktu widzenia, jak arsen, rtęć, ołów, a także selen. Metoda AFS-HG
charakteryzuje się takimi zaletami, jak duża selektywność, szeroki zakres liniowości
wskazań, bardzo wysoka czułość, niewielki poziom szumu umożliwiający osiąganie
bardzo niskich granic wykrywalności, a także relatywnie proste i tanie wymagania
aparaturowe [6, 16].
W obecnej pracy do oznaczania selenu w materiale biologicznym zastosowano
procedurę analityczną stanowiącą połączenie mineralizacji w systemie mikrofalo-
wym z procesem generacji wodorków selenu i pomiarem sygnału analitycznego
metodą AFS. Metodę testowano na przykładzie analizy próbek krwi ludzkiej. Bada-
nia skupiały się głównie na optymalizacji warunków mineralizacji próbek krwi oraz
na określeniu parametrów analitycznych opracowanej metody. Dokładność metody
zweryfikowano na podstawie analizy referencyjnych próbek krwi ludzkiej.

Metoda oznaczania selenu we krwi ludzkiej
33
MATERIAŁY I METODY
### Odczynniki
W badaniach wykorzystano następujące odczynniki: roztwór wzorcowy selenu
stężeniu 1 mg ml-1 Se (cz.d.a., Merck), stężony HNO3 (cz.d.a., Merck), stężony HCl
(cz.d.a., POCh), stężony H3PO4 (cz.d.a., POCh), stężona HCOOH (cz.d.a. POCh),
30% H2O2 (cz.d.a. POCh), 98% NaBH4 (cz.d.a., Sigma-Aldrich Chemie), 0,5% NaOH
(cz.d.a., Zakłady Chemiczne Oświęcim). Do przygotowania roztworów stosowano
wodę dejonizowaną o przewodnictwie właściwym ok. 1 ąS cm-1.
### Materiał do badań
Materiał do badań stanowiły próbki krwi ludzkiej (dwie partie po 50 ml) po-
chodzące od honorowych krwiodawców z Regionalnego Centrum Krwiodawstwa
Krwiolecznictwa w Krakowie. Próbki krwi przechowywano do czasu analizy w spe-
cjalnych fiolkach napylanych środkiem antykrzepliwym (K2EDTA) w temperaturze
—20oC. Analizom poddano również dwa materiały referencyjne krwi: próbkę pełnej
krwi ludzkiej „Seronom Trace Elements” (LOT 404109, SERO AS, Billingstad) w po-
staci liofilizowanej oraz próbkę surowicy krwi otrzymanej z Zakładu Bromatologii Col-
legium Medicum UJ. Skład drugiej próbki wyznaczono w ramach badań międzylabo-
ratoryjnych pn. „Interlaboratory comparisons for copper-zinc-selenium in blood se-
rum” nadzorowanych przez French Society for Clinical Biology, Vendeuvre Les Nancy.
### Aparatura
W procesie przygotowania próbek wykorzystano 12-pozycyjny mineralizator
mikrofalowy Mars 5X (CEM, Stany Zjednoczone). Proces mineralizacji prowadzono
w wysokociśnieniowych teflonowych naczyniach typu XP-1500 z równoczesną kon-
trolą temperatury i ciśnienia.
Do pomiaru sygnałów analitycznych zastosowano dwukanałowy spektrometr
fluorescencji atomowej z systemem do generacji wodorków AFS-230 (Haiguang Co.,
Chiny). W spektrometrze zainstalowano selenową lampę z katodą wnękową (o natę-
żeniu 100 mA) jako źródło promieniowania i fotopowielacz (pracujący pod napięciem
300 V) jako element detekcyjny.
WYNIKI I DYSKUSJA
Badania wstępne polegały na ustaleniu optymalnych warunków przygotowania
próbek krwi do pomiaru. W tym celu przeprowadzono oznaczanie selenu w próbkach
tej samej partii krwi w kilku seriach różniących się od siebie warunkami mineraliza-
cji. Przy ustalaniu tych warunków oparto się częściowo na danych zaczerpniętych
z literatury przedmiotu [8, 9]. Szczegółowy skład mieszaniny mineralizacyjnej oraz
sposób traktowania próbki po mineralizacji przedstawiono w tabeli I.
W każdej serii do mineralizacji pobierano trzy próbki krwi oraz próbkę ślepą (od-
czynnikową) o objętości ok. 0,5 ml. Każdąpróbkę przenoszono do naczynia teflonowe-
go, dodawano odpowiednią ilość mieszaniny mineralizującej, a następnie prowadzo-
no proces mineralizacji w trzech etapach różniących się warunkami maksymalnej
temperatury i ciśnienia: a) T = 160oC, P = 130 psi; b) T =180oC, P = 190 psi;

34
R. Wietecha, P. Kościelniak, T. Lech, M. Rymanowski
c) T = 200oC, P = 290 psi. Czas trwania kazdego etapu wynosił 8 min. Po ochłodzeniu
mineralizat poddawano działaniu kwasu solnego (bez dodatku lub z dodatkiem kwa-
su mrówkowego) w celu zredukowania Se(VI) do Se(IV), a następnie przenoszono do
kolby miarowej i uzupełniano do objętości 25 ml. Tak przygotowany roztwór gotowy
był do wprowadzenia do spektrometru fluorescencyjnego.
Generacja wodorków selenu następowała w układzie przepływowym po wprowa-
dzeniu próbki do roztworu zawierającego 2% tetraboran sodowy i 0,5% wodorotlenek
sodu. Roztwór ten łączył się z prędkością 6 ml min-1 z roztworem 3 M kwasu solnego
pełniącego rolę nośnika i był doprowadzany do separatora z prędkością 11 ml min-1,
a następnie do płomienia argonowo-wodorowego (o temperaturze 200oC) spektro-
metru fluorescencyjnego w strumieniu argonu z szybkością 800 ml min-1.
Do celów kalibracyjnych sporządzono serię roztworów wzorcowych zawierających
analit w stężeniach od 0 do 3,48 pg l-1. Roztwory te sporządzano przez rozcieńczenie
podstawowego roztworu wzorcowego 6 M roztworem kwasu solnego do jednakowej
objętości. Wyniki oznaczenia selenu w badanych próbkach krwi zebrano w tabeli I.
W dalszych badaniach analizę próbek krwi wykonywano w warunkach, jakie
ustalono w III serii oznaczeń (patrz tabela I). Przy wyborze kierowano się głównie
wartością wyznaczonego stężenia selenu w próbkach, która w przypadku tej serii
oznaczeń była zdecydowanie największa. Można było zatem sądzić, że ewentualny
ubytek selenu spowodowany procesem mineralizacji próbki był w tych warunkach
najmniejszy. Ponadto, istotne znaczenie ma również to, że wybrane warunki minera-
lizacyjne są niezwykle proste: stosuje się w tym przypadku pojedynczy odczynnik,
a przygotowanie mineralizatu do pomiarów nie wymaga ogrzewania, lecz jedynie
dodania kwasu solnego w odpowiednim stężeniu. Procedura mineralizacyjna jest
więc nie tylko mało kosztowna, ale związana z minimalnym ryzykiem zanieczyszcze-
nia próbki wprowadzanymi do niej odczynnikami.
W wybranych warunkach doświadczalnych wyznaczona granica oznaczalności
selenu wynosiła ok. 0,1 pg l-1, a precyzja oznaczeń (obejmująca powtarzalność pełnej
procedury analitycznej) kształtowała się na poziomie 1% (RSD).
W celu zbadania efektu interferencyjnego do oznaczenia selenu w danej próbce
krwi zastosowano dwie metody kalibracyjne: metodę serii wzorców i metodę dodat-
ków wzorca. W przypadku metody dodatków wzorca roztwory sporządzano w ten
sposób, że mineralizacji poddawano w jednakowych warunkach kilka porcji próbki,
a analit dodawano do próbek po mineralizacji w postaci wodnego roztworu wzorco-
wego. Stężenia selenu w dodanych roztworach wzorcowych wynosiły 0; 0,63; 1,25
i 1,88 pg l-1. Wzajemne położenie linii kalibracyjnych wyznaczonych obiema metoda-
mi (rycina 1A) świadczyło o występowaniu efektu interferencyjnego.
O istnieniu interferencji chemicznych w przypadku oznaczania selenu w materia-
łach biologicznych metodą AAS-HG po mineralizacji próbek w kwasie azotowym(V)
donoszono już kilkakrotnie [3, 5, 18]. Wysunięto hipotezę [5], że za efekt ten odpowie-
dzialne są tlenki azotu (NO i NO2), które, pozostając w próbce po procesie mineraliza-
cji, zapobiegają redukcji Se(VI) do Se(IV). W konsekwencji w próbce wprowadzanej do
spektrometru zmniejsza się stężenie analitu w formie Se(IV), a ponieważ ta forma
zdolna jest do tworzenia wodorków z największą wydajnością, sygnał analityczny nie
odpowiada faktycznej ilości selenu w próbce. Należy przypuszczać, że pomimo zastoso-
wania innej metody pomiarowej (AFS-HG), efekty interferencyjne zauważone w obec-
nych badaniach można tłumaczyć w ten sam sposób.

Metoda oznaczania selenu we krwi ludzkiej
35
W literaturze przedmiotu zaleca się stosowanie środków chemicznych do elimi-
nacji wpływów tlenków azotu na oznaczanie selenu metodą AAS-HG [5, 13, 18].
Chcąc jednak uniknąć komplikowania opracowywanej procedury analitycznej przez
wprowadzanie dodatkowych odczynników do próbki, podjęto próbę zabiegu nie stoso-
wanego dotąd przez innych autorów. Postanowiono mianowicie zredukować stężenie
tlenków azotu w fazie wodnej poprzez ich usuwanie znad próbek silnym strumieniem
argonu lub azotu. Zabieg ten wykonywano bezpośrednio po procesie mineralizacji
i ochłodzeniu próbek do temperatury otoczenia. Tak dodatkowo przygotowane próbki
analizowano z zastosowaniem dwóch metod kalibracyjnych. Położenie wyznaczonych
linii kalibracyjnych (rycina 1B) wskazywało, że w wyniku odpowiednio długiego
działania gazów (przez ok. 10 min) efekt interferencyjny ulegał eliminacji.
Efektywność redukcji efektów interferencyjnych ilustrują wyniki oznaczenia sele-
nu w próbce pełnej krwi ludzkiej przedstawione w tabeli II. O możliwości całkowitego
usunięcia tych efektów świadczą wyniki uzyskane dla próbek referencyjnych. Jak wi-
dać, zastosowana metoda analityczna pozwala na wyznaczenie stężenia selenu we
krwi z dokładnością określoną błędem względem nie większym niż 2%.
PODSUMOWANIE I WNIOSKI
Opracowana metoda analityczna dotycząca oznaczania selenu w krwi ludzkiej
obejmuje kilka etapów doświadczalnych:
- mineralizację próbki w piecu mikrofalowym w kwasie azotowym;
- odpędzenie tlenków azotu znad mineralizatu i przeprowadzenie próbki do śro-
  dowiska kwasu solnego;
- kalibrację oznaczenia metodą serii wzorców za pomocą wodnych roztworów
  wzorcowych;
- generację wodorków selenu w systemie przepływowym;
- pomiar sygnałów analitycznych za pomocą atomowego spektrometru fluore-
  scencyjnego.
Procedura charakteryzuje się następującymi zaletami:
- jest stosunkowo szybka i prosta;
- wymaga użycia zaledwie kilku odczynników, dzięki czemu jest stosunkowo ta-
  nia i bezpieczna;
- zapewnia możliwość oznaczania selenu na poziomach występujących w krwi
  ludzkiej;
- zapewnia uzyskanie wyników analitycznych o bardzo dobrej precyzji i dokład-
  ności.
Pewnym ograniczeniem w stosowaniu tej metody jest konieczność użycia stosun-
kowo drogiej (piec mikrofalowy) i trudnodostępnej (spektrometr AFS) aparatury spe-
cjalistycznej. Uzyskane wyniki dowodzą jednak, że proponowany sposób analizy
krwi na zawartość selenu jest w pełni konkurencyjny w stosunku do innych metod
stosowanych w tym samym celu i może być z przekonaniem zalecany do realizacji
w laboratoriach toksykologicznych i klinicznych.

36
R. Wietecha, P. Kościelniak, T. Lech, M. Rymanowski
Podziękowania:
Autorzy serdecznie dziękująPanu Drowi Stanisławowi Walasowi za cenne dyskusje
i pomocne uwagi dotyczące atomowej spektrometrii fluorescencyjnej.