Anal Bioanal Chem (2017) 409:4383–4393 DOI 10.1007/s00216-017-0382-y
Improving repeatability of capillary electrophoresis—a critical comparison of ten different capillary inner surfaces and three criteria of peak identification
Paweł Mateusz Nowak1 & Michał Woźniakiewicz1 & Marta Gładysz 1 & Magdalena Janus1 &
Paweł Kościelniak1
Received: 20 March 2017 / Revised: 18 April 2017 / Accepted: 25 April 2017 / Published online: 8 May 2017
# The Author(s) 2017. This article is an open access publication
Keywords Capillary electrophoresis . Capillary coating . Electrophoretic mobility . Micellar electrokinetic chromatography . Relative migration times . Repeatability
Abstract A poor repeatability of migration times caused by the fluctuations of electroosmotic flow (EOF) is an inherent weakness of capillary electrophoresis. Most researchers en- deavor to prevent this problem using relative migration times or various capillary coatings which are expensive and not easy in comparison. Herein, we present an original approach to this problem: we apply a model sample designed to induce signif- icant EOF instability, in order to critically compare ten capil- lary types with different physicochemical characteristics. Moreover, we accompany capillary modification with the evaluation of various criteria of peak identification: migration time, migration times ratio, and electrophoretic mobility. Our results show a great effectiveness of a dynamic coating in the stabilization of migration times, with the average RSD(%) value reduced from 3.5% (bare silica capillary) down to 0.5%. The good outcomes were also obtained for the surfactant-modified silica and amine capillaries. For the cap- illaries exhibiting significant instability of EOF, electrophoret- ic mobility turned out to be a more universal and reliable criterion of peak identification than relative migration time. It can be explained by an intrinsic dependency of migration times ratio on EOF change, which should always be consid- ered during the selection of an internal standard.
Abbreviations
AMI BGE CE DC EOF HSA MEKC SDS WAR
Amitriptyline Background electrolyte Capillary electrophoresis Dynamic coating Electroosmotic flow Human serum albumin
Micellar electrokinetic chromatography Sodium dodecyl sulfate
Warfarin
Introduction
Capillary electrophoresis (CE) is one of the most powerful analytical techniques of a wide and still not fully uncovered analytical potential. It enables separation of the structurally similar molecules, including enantiomers, and provides low consumption of sample material and buffers, short analysis time, a variety of factors that can be easily used to improve resolution, and high automation degree. Nevertheless, CE has its own Achilles’ heels which often overshadow its huge an- alytical power—relatively poor repeatability/reproducibility and sensitivity [1–4]. The former becomes a significant limi- tation when the RSD values obtained for migration times sur- pass the acceptable limits, often over 3%. The alterations of migration velocity may be caused by many factors: fluctua- tions of electroosmotic flow (EOF), unstable temperature, cur- rent, pH, ionic strength, presence of air bubbles, and siphoning effect [1–4]. EOF fluctuations often pose the main source of
Electronic supplementary material The online version of this article (doi:10.1007/s00216-017-0382-y) contains supplementary material, which is available to authorized users.
* Paweł Mateusz Nowak nowakp@chemia.uj.edu.pl
1
Faculty of Chemistry, Department of Analytical Chemistry, Jagiellonian University in Kraków, Ingardena St. 3,
30-060 Kraków, Poland
RESEARCH PAPER

4384
Nowak P.M. et al.
this instability. They result from the physical modification of capillary inner surface, caused most commonly by the analyte- wall interactions, formation of insoluble aggregates, and in- sufficient capillary rinsing. Insofar as capillary regeneration may be enhanced by using the longer rinsing times and the additional solvents, the dynamical analyte-wall interactions occurring during the electromigration constitute a bigger prob- lem and often hamper analysis. It is particularly frequently observed for the real samples with complex matrix and the protein-containing samples [5]. The large repeatedly ionized biomolecules possess many potential sites of such interac- tions. In such conditions, an average capillary lifetime is also appreciably reduced.
The most popular method for improving repeatability and for enabling transfer of methods in the qualitative analysis is the use of relative migration times, i.e., a ratio of migration times obtained for the analyte and the internal standard. Assuming an efficient temperature control, relative migration times do not depend on capillary dimensions, separation po- tential, and viscosity change. However, what can be overlooked by the less experienced users of the CE technique, effectiveness of relative migration times in reducing the EOF fluctuation-related error is by definition limited and depends strictly on a difference in mobility between analyte and inter- nal standard (see the next part of manuscript for more details). An alternative way is to determine electrophoretic mobility of the analytes instead of migration times ratio or to perform the transformation of the whole electropherogram from a time scale into a mobility scale. Effectiveness of both methods has been examined and confirmed in many studies [6–19]; nevertheless, electrophoretic mobility is still quite rarely used as a direct criterion of peak identification.
Another approach to this problem is the application of physicochemically modified capillaries, coated permanently or dynamically with chemically versatile agents. The literature concerning this subject is broad, and among numerous arti- cles, one may encounter the applications of the commercially offered ready-to-use capillaries, kits for dynamic coating, and other coating materials synthetized ab initio [5, 20–37]. The major advantage is that a modified capillary surface exhibits weaker affinity to analytes, especially macromolecules, stabi- lizes EOF, and thus allows one to enhance repeatability. Moreover, the use of polyamine positively ionized coating enables the EOF reversal, from cathodic (toward cathode) to anodic (toward anode), while the neutral capillaries provide its total elimination. On the other hand, the use of any type of coating significantly elevates the cost of analysis. The direct comparisons of various types of capillary modifications are rarely performed [27, 29, 33]. On that account, there is still little known about advantages and predispositions of the given capillary types in regard to the particular types of methods and experimental conditions, and this hinders their critical evalu- ation and selection.
This work offers a fresh look at the repeatability of the CE- based analyses affected by a variable EOF magnitude. We present an original approach, the use of a specially designed model sample inducing the significant flow instability and containing chemically varied analytes, combined with the concurrent examination of ten various capillary types of the different inner surfaces, and the comparison of three alterna- tive criteria of peak identification: migration times, relative migration times, and electrophoretic mobilities. Our results may be of importance for all CE users and for researchers without experience in CE who look for an alternative tech- nique for liquid and gas chromatography. The presented data help one to choose the optimal capillary type, and show that a careful consideration of many factors is crucial to select the optimal criterion of peak identification.
Materials and methods
Instrumentation
The experiments were performed with a P/ACE MDQ Capillary Electrophoresis (CE) System (Beckman-Coulter, Brea, CA, USA) equipped with a diode array detector. The following commercially available capillaries were applied: an uncoated bare fused-silica capillary (silica); amine eCAP™ polyamine-coated capillary (amine) providing the reversal of EOF; neutral polyacrylamide-coated capillary (neutral PAA): and neutral polyvinyl alcohol-coated capil- lary (neutral PVA) providing neutralization of EOF, all sup- plied by Beckman-Coulter, and a Celerity™ diol phase- coated capillary column (diol) supplied by MicroSolve Technology, Eatontown, NJ, USA. All capillaries were of 60-cm total length, 50-cm effective length, and 50-μm inter- nal diameter. In order to obtain other physicochemical mod- ifications, the silica and amine capillaries were subjected to a dynamic coating performed with the commercially available kit CEofix™ pH 6, containing the initiator™ and accelera- tor™ solutions (Beckman-Coulter): obtaining the CEofix™ dynamically coated bare silica capillary (DC-silica) and CEofix™ dynamically coated amine eCAP™ capillary (DC-amine). Coating by a successive multiple ionic layer (SMIL) was performed with the silica capillary (cationic layer was formed by initiator™ and anionic layer was formed by accelerator™), and by a single anionic- polymer layer with the amine capillary, pre-coated perma- nently by a polyamine layer. In addition, in the case of bare silica, amine, and neutral PAA capillaries, the background electrolyte (BGE) was modified by addition of 30 mM so- dium dodecyl sulfate (SDS). Since the presence of micelles may influence on all dynamic analyte-wall interactions, these variants were included in the comparison as the in- dependent physicochemical modifications. In the case of

Improving repeatability of capillary electrophoresis
4385
amine capillary, SDS molecules coat the capillary’s
wall
EOF is virtually entirely eliminated. Sample injection was performed using the pressure of 2.8 kPa (0.4 psi) for 4.0 s. The UV-vis absorption spectra were collected between 200 and 600 nm; 200 nm was the analytical wavelength for plotting electropherograms. The analysis of electropherograms was done using the Origin 9.1. software (OriginLab Corporation, Northampton, MA, USA).
forming a characteristic admicelle layer, reversing its charge [38]. Schematic illustration of all ten capillary types is shown in Fig. 1.
The sample trays and capillaries were conditioned at 25 °C. The rinsing of capillaries between the runs was done by applying a pressure of 137.9 kPa (20 psi). The procedures were either developed in our laboratory or followed the instructions provided by the suppliers. They are presented in detail in Table 1. BGE was com- posed of the phosphate buffer (Na2HPO4/NaH2PO4) of the 50 mM ionic strength and pH 6.0, except the dynam- ic coating where BGE (accelerator) was supplied as a kit component and was not prepared ab initio in the labora- tory. It was also the phosphate buffer; its pH was equil- ibrated before use to 6.0. Its ionic strength was un- known, but probably similar, as it follows from the com- parison of current measured during separations. BGE based on the phosphate buffer was selected to provide the same conditions for all capillaries, including the com- mercial kit (CEofix™). All aqueous solutions were pre- pared using a deionized water (Milli-Q system, Merck- Millipore, Darmstadt, Germany), filtered through a 0.45-μm regenerated cellulose membrane and then degassed by sonication and centrifugation. The separa- tion voltage of 20 kV was always applied. The reverse polarity (cathode at inlet) was used in the case of amine, neutral, and SDS-modified neutral capillaries, while the normal polarity (anode at inlet) in all other cases. The
Sample composition
The sample composition was selected taking several issues into account: (i) the presence of macromolecule intrinsically prone to interactions with capillary inner surface; (ii) the pres- ence of small molecules of different charge, to include analytes migrating both before and after EOF marker; (iii) the presence of neutral EOF marker; and (iv) simplicity, to avoid potential problems with peaks overlapping, especially in the surfactant-modified capillaries. The final composition
was as follows: human serum albumin (HSA), molecular weight (MW)≈66,500 Da—macromolecule known for its nonspecific adsorption/adhesion to silica surface via various
types of interactions, positively charged amitriptyline (AMI), MW = 277 Da, pKa≈9.4, negatively charged warfarin (WAR), MW = 308 Da, pKa≈5.0 [39], and EOF marker—dimethyl
sulfoxide (DMSO), MW = 78 Da, all dissolved in BGE. In
the case of DC-silica, DC-amine, and SDS-modified capil- laries, the standard BGE without kit components and buffer additives was used for sample preparation. All analytes were supplied by Sigma-Aldrich (St. Louis, MO, USA). Three con-
centration levels of each analyte were used: 500, 250, and 125 μg mL−1. To keep the sample composition as simple as possible, only one macromolecule was used, but instead,
current values were between 45 and 60 μA
for all
cap-
illary types. To avoid significant increase of separation time, the positive pressure of 2.8 kPa (0.4 psi) was used in case of the neutral capillaries (without SDS) where
Fig. 1 Schematic illustration of ten physicochemically different capillaries including their inner surface and the direction of EOF. Symbols: - negative charge; + positive charge; ( ) dynamic polyionic layer
(renewable); N1 polyacrylamide (PAA) layer; N2 polyvinyl alcohol (PVA) layer; II hydrophobic tail-tail interactions. MEKC-SDS method is presented as a distinct capillary modification (see the text)

4386
Nowak P.M. et al.
Table 1	The procedures used for capillary rinsing, before the first use on a given working day and between the following runs
Before the first usea
Full capillary name/(abbreviation)
Between runs
Uncoated bare fused-silica capillary/(silica)
Methanol: 20 min
0.1 M HCl: 5 min Deionized H2O: 5 min
0.1 M NaOH: 20 min BGE: 20 min Deionized H2O: 20 min
Amine regenerator solution (supplied): 20 min BGE: 20 min
Deionized H2O: 20 min BGE: 20 min
Deionized H2O: 20 min BGE: 20 min
Deionized H2O: 20 min
0.1 M HCl: 5 min Deionized H2O: 5 min
0.1 M NaOH: 20 min BGE: 20 min Methanol: 20 min
0.1 M HCl: 5 min Deionized H2O: 5 min
0.1 M NaOH: 20 min BGE (normal): 20 min Deionized H2O: 20 min
Amine regenerator solution (supplied): 20 min BGE (normal): 20 min
Methanol: 20 min
0.1 M HCl: 5 min Deionized H2O: 5 min
0.1 M NaOH: 20 min BGE—with SDS: 20 min Deionized H2O: 20 min
Amine regenerator solution (supplied): 20 min BGE—with SDS: 20 min
Deionized H2O: 20 min BGE—with SDS: 20 min
0.1 M NaOH: 3 min BGE: 3 min
Amine eCAP™ polyamine-coated capillary/(amine)
Amine regenerator solution (supplied): 3 min BGE: 3 min
Neutral polyacrylamide-coated capillary/(neutral PAA)
Deionized H2O: 3 min BGE: 3 min
Deionized H2O: 3 min BGE: 3 min
0.1 M NaOH: 3 min BGE: 3 min
Neutral polyvinyl alcohol-coated capillary/(neutral PVA)
Celerity™ diol phase-coated capillary/(diol)
CEofix™ dynamically coated bare silica capillary/(DC-silica)
0.1 M NaOH: 3 min
Initiator—poly-cation (supplied): 3 min BGE—with accelerator (supplied):
3 min
CEofix™ dynamically coated amine eCAP™ capillary/(DC-amine)
Amine regenerator solution (supplied): 3 min BGE—with accelerator (supplied):
3 min
0.1 M NaOH: 3 min BGE—with SDS: 3 min
SDS-modified bare silica capillary/(SDS-silica)
SDS-modified amine eCAP™ capillary/(SDS-amine)
Amine regenerator solution (supplied): 3 min BGE—with SDS: 3 min
SDS-modified neutral polyacrylamide-coated capillary/(SDS-neutral (PAA))
Deionized H2O: 3 min BGE—with SDS: 3 min
a For the fresh capillary conditioning, the duration of each individual step was doubled
Results and discussion
relatively high concentration range was applied. The repeti- tions number amounted to 6 for each concentration level. The new vials with fresh BGE solutions were used for each con- centration level, to minimize buffer depletion.
Theoretical background
In CE, migration times are inversely proportional to the ap- parent mobility, which derives directly from the sum of two vectors: electrophoretic and electroosmotic ones. It is to be noted that the change of EOF affects only the electroosmotic vector and entails disproportionate shifts of migration times for different peaks. It is shown in Fig. 2A, where we consider the separation of four distinct model analytes with the conven- tional bare silica capillary. The same change in the EOF vector induces different relative changes in the total vector for the given analytes. Therefore, there is no proportionality in the shift of migration times; the smallest change is observed for the fastest compound (cationic), and the largest one for the slowest compound (second anionic). In this respect, the analytes of a large negative electrophoretic mobility are
Calculations
The values of electrophoretic mobility (μ ) were obtained
ep
from Eq.1:
(
\
 L ∙L
 1
  1
tot eff
V
−
μ ¼
ð1Þ
ep
ttot
teof
where Ltot and Leff are the total and effective capillary lengths (cm), V is the nominal separation voltage set up in the software (kV); ttot is the total (absolute) migration time of analyte (min); while teof is the time measured for the neutral marker of EOF—DMSO (min).

Improving repeatability of capillary electrophoresis
4387
representative electropherograms are shown in Fig. 3, while the values of migration times are shown in Table S1 in the Electronic Supplementary Material (ESM). The first issue is the effect of capillary clogging, which was noted for the sev- eral independent bare silica capillaries. Every time the capil- lary got clogged after around 8–12 separations, making further analysis is impossible. It was observed as the appearance of artifacts on the electropherogram, significant drift of the base- line and the lack of the signal from analytes. This effect was not observed with the other capillaries, suggesting that the interaction between HSA and Bunprotected^ silanol groups
may be the main obstacle hampering analysis. A similar effect
was observed by us in our recent work, when HSA was used as an pKa-shift inducer [33].
The bare silica capillary (before clogging), two dynamical- ly coated capillaries, and silica and amine SDS-modified cap- illaries enabled the detection of all sample ingredients (Fig. 3). The HSA peak was not observed with the amine capillary, most likely because of the strong protein-wall interactions caused by the positive charge of the wall and the negative net charge of the macromolecule. On the other hand, AMI was not detected using the neutral capillaries since the me- chanically induced flow of buffer was too weak to register a cationic compound upon application of the reverse polarity. DMSO peak was not observed with the SDS-modified neutral capillary owing to the lack of EOF. The diol-modified capil- lary, in turn, exhibited most probably a strong retention of HSA and AMI, yielding only the peaks of DMSO and WAR. In this context, regarding the clogging of the bare silica capillary, the dynamically coated and SDS-modified silica and amine capillaries seem the best option for performing analysis of the chemically varied samples containing macromolecules as well as cationic and anionic compounds.
As regards the RSD values obtained for migration times, an interesting relation is observed for the bare silica and amine capillaries. In the former case, the RSD values rise in the order AMI > DMSO > HSA > WAR, the same as the migration order observed with this capillary. In the case of the amine capillary, WAR > DMSO > AMI order is noted, also consis- tent with the increase of migration times. This observation fully agrees with the expectation that due to the increasing EOF fluctuation-related error, discussed previously, repeat- ability of migration times should drop in accordance with migration order.
The dynamically coated silica capillary is characterized by the very low RSD values, averagely 0.78% for HSA and 0.20% for the rest of compounds, respectively. This excellent outcome is not preserved by the dynamically coated amine capillary, for which the RSD values start to increase from
the second concentration level (250 μg × mL−1), with the
increasing injections number. In this case, the positively
charged polyionic layer is permanently present on the wall, unlike the dynamically coated silica capillary; hence, the HSA
Fig. 2 Theoretical simulation of the potential shifts of parameters in the qualitative analysis caused by the EOF change, (A) Using migration times obtained for four different compounds (1–4) exhibiting different migration velocity. (B) Using migration time and relative migration times obtained for one compound (4) and considering three different internal standards (IS, 1–3). The shifts were calculated as a relative change of the parameter upon the given EOF alteration. The inset graphics present schematic electropherogram and adding of vectors for the particular analytes (the initial electroosmotic mobility equals 100 and
it varies +/− 10; electrophoretic mobilities are constant and they equal
+20, 0, −20, −40 for the given analytes; the values were chosen arbitrarily
to visualize the discussed phenomenon)
inherently much more susceptible to EOF variation than the cationic compounds.
It is also seen that due to the non-proportional shift of migration times, the relative migration times are also chang- ing, and their utilization only partially eliminates the EOF change-related error (Fig. 2B). The more similar is mobility of analyte and standard; the lower error can be expected. It would be entirely eliminated then, and only then, if internal standard had the same migration time as analyte.
Repeatability of migration times
After performing separation of the model sample using ten different capillary types, the repeatability of migration times was expressed by the RSD values, depicted in Table 2. The

4388
Nowak P.M. et al.
Table 2	The RSD (%) values (n = 6) obtained for various analytes on three concentration (conc.) levels (500, 250, 125 μg× mL−1) using ten different capillaries, and calculated for three alternative parameters: t—migration times, t/tIS—relative migration times (calculated in respect to DMSO as internal
standard), μep—electrophoretic mobilities
Parameter
Analyte
Conc.
Silica
Amine
DC-silica
DC-amine
Neutral PAA
Neutral PVA
SDS-silica
SDS-amine
SDS-neutral (PAA)
Diol
t
AMI
500
250
125
500
250
125
500
250
125
500
250
125
500
250
125
500
250
125
500
250
125
500
250
125
500
250
125
500
250
125
1.2
C C 5.9 C C 2.2 C C 4.9 C C 1.1 C C 3.8 C C 1.6 C C 1.3 C C 0.8 C C 2.3 C
C
6.2
5.9
6.3
1.2
1.3
1.2
2.5
2.5
2.3
X X X 4.0
3.4
4.0
1.3
1.2
1.2
X X X 2.0
0.2
0.1
1.7
1.3
1.2
X X X
0.2
0.0
0.2
0.4
0.1
0.2
0.3
0.2
0.2
1.2
0.3
0.8
0.2
0.2
0.1
0.2
0.2
0.2
1.0
0.3
0.7
0.3
0.2
0.3
1.0
1.0
0.8
2.7
0.9
1.8
0.0
1.0
4.1
0.5
1.6
5.2
0.3
1.4
4.8
0.6
2.7
3.3
0.3
0.6
0.9
0.2
0.3
0.4
0.4
1.3
1.8
0.4
0.7
2.9
1.0
0.9
2.2
0.9
2.8
1.2
X X X 1.1
1.5
1.1
0.2
3.2
0.8
0.8
1.0
1.9
X X X 1.1
1.8
0.4
0.7
0.6
0.4
X X X 2.2
0.2
1.4
1.8
0.8
1.3
X X X 0.8
1.5
0.9
0.1
2.9
1.3
4.2
1.9
1.7
X X X 0.8
1.6
0.6
4.3
1.1
0.6
X X X 1.6
0.9
0.2
4.9
1.5
1.4
11.4
0.3
0.2
10.3
0.3
0.4
3.4
0.4
0.0
8.6
0.8
0.1
7.7
0.2
0.2
6.6
0.2
0.4
5.0
0.6
0.1
0.4
0.6
0.1
0.7
0.6
0.3
0.5
0.7
0.1
8.4
3.7
0.9
5.5
3.8
0.6
2.5
0.9
0.4
5.6
3.2
0.7
6.0
2.9
0.7
3.1
3.0
0.6
3.2
2.4
0.7
0.3
0.6
0.4
1.0
1.1
0.5
0.6
1.1
0.7
2.3
1.2
3.7
2.9
2.8
4.4
X X X 1.6
0.7
2.7
X X X X X X X X X X X X X X X X X
X
X X X 7.3
8.4
8.2
3.1
2.6
3.6
X X X X X X 4.2
6.0
4.7
X X X X X X 1.3
1.9
0.8
X X X
WAR
DMSO
HSA
t/tIS
AMI
WAR
HSA
μep
AMI
WAR
HSA
The capillary name abbreviations are consistent with Table 1; C—capillary clogging effect, X—lack of the corresponding peak due to adsorption of analyte on capillary inner surface or very low apparent mobility
molecules are probably being retained on this layer from run to run and a specific Bsurface passivation effect^ takes place. As it comes to the neutral capillaries, the RSD values are
moderate and no clear trend is visible. Nevertheless, the neu- tral PAA capillary yields better results than the neutral PVA capillary. The SDS-modified silica and amine capillaries are characterized by the high RSD values on the first concentra-
tion level (500 μg× mL−1), but on the second and third levels (250 and 125 μg × mL−1) applied chronologically later, the
RSD values become significantly reduced. This is an opposite effect to the dynamically coated amine capillary where the higher RSD values were observed on the second and third concentration levels; here, a specific surface stabilization and
Bmaturation^ may occur as a result of the increasing injection
number. Importantly, the high RSD values noted on the first concentration level also agree with the effect discussed in the previous section. In this separation medium, electrophoretic mobility of all the analytes, except DMSO, is negative and significant due to the interactions with the negatively ionized SDS molecules. For this reason, the RSD values obtained for AMI, WAR, and HSA are much higher than those obtained for DMSO. The SDS-modified neutral capillary displays a differ- ent characteristic. The RSD values are not changing along the sequence of runs, and they are rather moderate. Finally, the diol-modified capillary yields a weak repeat- ability, comparable for all three concentration levels and, again, significantly worse for the anionic WAR detected long after DMSO.

Improving repeatability of capillary electrophoresis
4389
Fig. 3 Representative electropherograms obtained for all ten capillaries on the highest concentration level
(500 μg × mL−1). The negative
peaks observed for the DC-silica,
DC-amine and SDS-amine capil- laries may stem from the lack of kit components and SDS mole- cules in the sample solution. The EOF strength was measured using always the positive DMSO peak

4390
Nowak P.M. et al.
In conclusion, for this type of sample, the dynamic coating performed with the silica capillary is appreciably better than any other capillary type. In reference to the other published works [33, 36, 37], the effectiveness of the amine- and diol- modified capillaries in comparison to the bare silica capillary occurred to be noticeably worse. The dynamically coated sil- ica capillary, contrary to the previous observations [37], turned out to be more effective in the stabilization of migration times than the dynamically coated amine capillary. This effect seems to be specific for the sample containing HSA molecules possessing the negative net charge.
reaches even 8.8%. It can be explained by the nonspecific protein-wall interactions, leading to the adsorption of macro- molecule and the significant peak broadening. Regarding AMI, WAR, and DMSO as possible internal standards, the differences are noticeable for the bare silica-, amine-, and both SDS-modified capillaries. There is a tendency that the more similar the mobilities of the analyte and the internal standard are, the lower RSD is noted. For example, DMSO is a better standard than WAR for the identification of the AMI peak in the bare silica and amine capillaries, whereas much worse in the SDS-modified capillaries where WAR and AMI migrate with a similar velocity (see electropherograms in Fig. 3). This confirms the theoretical divagations presented previously.
Another issue is the comparison between relative migration times and electrophoretic mobilities (see Table 2 for the RSD values and Table S1 in ESM for the calculated mobilities). It is interesting than when we consider the highest RSD values obtained for the absolute migration times in this experiment, those over 5%, the use of electrophoretic mobility allows to improve repeatability much more effectively than the use of relative migration time calculated in respect to DMSO (i.e., using the same input value as in the case of electrophoretic mobility where DMSO plays a role of the EOF marker). This effect is visible for WAR identified with the bare silica- and diol-modified capillaries, AMI in case of the amine capillary, and for all compounds in case of the SDS-modified capillaries. Interestingly, the consistent observation was also reported in our recent work, when we used the uncoated capillary and the sample devoid of protein [40]. However, as it was mentioned, in the SDS-modified capillaries, both AMI and WAR (detect- ed within a comparable time) are better standards than DMSO. In consequence, some RSD values obtained for relative mi- gration times with different standards are lower, similar to those obtained for electrophoretic mobilities (compare Tables 2 and 3). This is, however, not observed for the bare
Other criteria of peak identification
As it is seen in Table 2, the repeatability of relative migration times is in general better than the absolute times, irrespective of the analyte and capillary type. This confirms the well- documented observations that a ratio of migration times is less sensitive to EOF change than absolute migration time, inde- pendent of the analyte charge and size [2, 3, 6, 7]. However, the differences observed in case of the dynamically coated silica and both neutral capillaries are small, and they should be regarded as statistically insignificant. The comparison of the RSD values obtained for relative migration times using the various internal standards is shown in Table 3. This compar- ison includes all possible combinations, to present how repeat- ability of data is correlated with the difference in electropho- retic mobility and the nature of internal standard.
Taking into account four possible internal standards, the differences observed for the dynamically coated and neutral capillaries are low, within the margin of the random error. In the other capillaries, almost in each case the worst repeatabil- ity and the highest RSD values are observed for HSA used as an internal standard. The RSD value obtained for WAR as an analyte and HSA as a standard in the bare silica capillary
Table 3 The RSD(%) values averaged from three concentration levels (500, 250,
125 μg × mL−1) obtained for the
relative migration times using
different capillaries and calculated for various internal standards (IS)
Analyte
IS
silica
Amine
DC-
silica
DC-
amine
Neutral PAA
Neutral PVA
SDS-
silica
SDS-
amine
Diol
AMI
WAR DMSO HSA AMI DMSO HSA AMI WAR HSA AMI WAR
DMSO
4.7
1.1
5.2
4.7
3.8
8.8
1.1
3.8
5.8
5.4
9.5
1.6
5.0
3.8
–
5.0
1.3
–
3.8
1.3
–
–
–
–
0.2
0.2
0.7
0.2
0.2
0.6
0.2
0.2
0.7
0.7
0.6
0.7
0.8
0.6
1.2
0.8
0.3
1.2
0.6
0.3
1.1
1.2
1.2
1.1
0.6
2.7
1.2
0.6
2.4
0.8
2.5
2.2
1.8
1.1
0.8
1.9
1.5
3.2
1.4
1.5
2.2
0.7
3.2
2.2
2.1
1.4
0.7
2.1
–
–
–
–
1.1
0.8
–
1.1
1.2
–
0.8
1.2
–
–
–
–
1.0
1.9
–
1.0
2.1
–
1.9
2.0
–
–
–
–
5.0
–
–
5.1
–
–
–
–
WAR
DMSO
HSA
The capillary name abbreviations are consistent with Table 1

Improving repeatability of capillary electrophoresis
4391
silica and amine capillaries; there, the application of other standards than DMSO deteriorates the repeatability. These effects stem from the intrinsic dependency of relative migra- tion times on EOF magnitude, which can be minimized by selecting the internal standard of a similar mobility as the analyte (see section BTheoretical background^).
It should also be pointed out that the repeatability of abso-
lute migration times, relative migration times, and electropho- retic mobilities may be very similar, as in the case of AMI detected with the bare silica capillary, WAR detected with the amine capillary, or all molecules detected with the dynamical- ly coated silica capillary. In this situation, the influence of EOF instability is low, and the other random errors are deci- sive. Most importantly, the repeatability of relative migration times and electrophoretic mobilities may be decreased by the fact that two migration times are required to calculate these parameters, obtained for the analyte and the standard/marker, respectively. This introduces the increased random uncertainty as compared to the absolute migration times, stemming, e.g., from peak broadening and its asymmetry, instability of current and temperature during separation, and analyte-wall interac- tions [2, 6]. Therefore, a thorough consideration of all these factors is crucial for the selection of an appropriate criterion of peak identification.
amine capillary exhibits the similar electroosmotic mobility as the uncoated capillary, unlike the separations performed in the different buffer systems [35]. This outcome confirms that when the amine capillary is applied with the phosphate buffer, EOF is reduced due to the specific interaction between the positively charged wall and the phosphate buffer compo- nents [37, 41]. Therefore, aside from the other benefits, all modified capillaries tested in this experiment are rather useless in decreasing a total separation time in this experimental setup.
Conclusions
A detrimental impact of the analyte-wall interactions on the capillary lifetime and the repeatability of migration times may be successfully prevented by the capillary coating or addition of surfactant to BGE. The dynamic coating (CEofix™) allows one to appreciably reduce the EOF fluctuations for the chem- ically varied and protein-rich samples. Noticeably, in this case, the use of electrophoretic mobilities and relative migration times instead of absolute times does not bring an improvement because of the prevalent impact of other random errors, in- creased by the use of standard/marker. The application of the dynamic poly-cationic layer (initiator solution™) is crucial for preserving functionality of coating; its replacement by the permanently present amine layer (eCAP amine capillary™) may have an unfavorable impact on the repeatability. After stabilization/pre-treatment of the inner surface, the SDS- modified silica and amine capillaries are also effective in the stabilization of migration times. The remaining capillaries are evidently more prone to the analyte-wall interactions; howev- er, any modified capillary allows one to avoid capillary clog- ging, which in this study, was inevitable with the bare silica capillary. When the variation of absolute migration times is significant, over 2–3%, the use of relative migration times or electrophoretic mobilities is advised to ensure the reliable peak identification. In the former case, it is crucial to choose the internal standard of a similar mobility to the analyte. It is caused by the fact that relative migration times are intrinsically dependent on EOF change, although they are less sensitive than absolute times. Therefore, electrophoretic mobilities seem to be a more universal and reliable criterion of peak identification when a large EOF variation is expectable, or when the mobility of the analytes differs considerably. Nevertheless, a suitable correction of the inherent systematic errors attributed to the mobility determination is required to provide the high reliability and broad applicability, in partic- ular the Joule heating-related effects [42]. The reliable elec- trophoretic mobility values may be useful in performing meth- od transfer between different experimental setups, e.g., capil- lary types, dimensions, or separation voltages. The presented data may be helpful in development of versatile CE-based
Electroosmotic mobility
The magnitude of EOF was calculated for each capillary and presented in Fig. 4. The significant difference in comparison to the bare silica capillary is observed only for the SDS- modified capillaries. Nevertheless, due to the interactions with the negatively ionized SDS micelles, the overall separation time is actually similar as in the other capillaries. It is also interesting that in the present experimental conditions, the
Fig. 4 The average values of electroosmotic mobility obtained for various capillaries, using the sample containing all analytes on the
highest analyte concentration level (500 μg × mL−1)

4392
Nowak P.M. et al.
using capillary electrophoresis. Macromolecules. 2009;42:3152– 61.
Hoagland DA, Arvanitidou E, Welch C. Capillary electrophoresis measurements of the free solution mobility for several model poly- electrolyte systems. Macromolecules. 1999;32:6180–90.
Cottet H, Biron JP. Charge- and size-based separations of polyelec- trolytes by heart-cutting two-dimensional capillary electrophoresis. Macromol Chem Phys. 2005;206:628–34.
Oudhoff K, Schoenmakers P, Kok WJ. Characterization of polyeth- ylene glycols and polypropylene glycols by capillary zone electro- phoresis and micellar electrokinetic chromatography. J Chromatogr A. 2003;985:479–91.
Tetsuo I, Jun K, Yasuyuki KJ. Electropherograms in capillary zone electrophoresis plotted as a function of the quantity of electric charge. J Chromatogr A. 1998;810:183–91.
Mammen M, Colton IJ, Carbeck JD, Bradley R, Whitesides GM. Representing primary electrophoretic data in the 1/time domain: comparison to representations in the time domain. Anal Chem. 1997;69:2165–70.
Chamieh J, Martin M, Cottet H. Quantitative analysis in capillary electrophoresis: transformation of raw electropherograms into con- tinuous distributions. Anal Chem. 2015;87:1050–7.
Schmitt-Kopplin P, Garmash AV, Kudryavtsev AV, Menzinger F, Perminova IV, Hertkorn N, Freitag D, Petrosyan VS, Kettrup A. Quantitative and qualitative precision improvements by effective mobility-scale data transformation in capillary electrophoresis anal- ysis. Electrophoresis. 2001;22:77–87.
Hajba L, Guttman A. Recent advances in column coatings for cap- illary electrophoresis of proteins. TrAC—trends. Anal. Chem. 2017;90:38–44.
Melanson JE, Baryla NE, Lucy CA. Dynamic capillary coatings for electroosmotic flow control in capillary electrophoresis. TrAC— trends. Anal. Chem. 2001;20:365–74.
Horvath J, Dolník V. Polymer wall coatings for capillary electro- phoresis. Electrophoresis. 2001;22:644–55.
Dolník V. Wall coating for capillary electrophoresis on microchips. Electrophoresis. 2004;25:3589–601.
Znaleziona J, Petr J, Knob R, Maier V, Ševčík J. Dynamic coating
agents in CE. Chromatographia Supplement. 2008;67:S5–S12. Hardenborg E, Zuberovic A, Ullsten S, Söderberg L, Heldin E, Markides KE. Novel polyamine coating providing non-covalent deactivation and reversed electroosmotic flow of fused-silica capil- laries for capillary electrophoresis. J Chromatogr A. 2003;1003: 217–21.
Danel C, Melnyk P, Azaroual N, Larchanché P-E, Goossens J-F, Vaccher C. Evaluation of three neutral capillary coatings for the determination of analyte-cyclodextrin binding constants by affinity capillary electrophoresis. Application to N,N′-disubstituted pipera- zine derivatives. J Chromatogr A. 2016;1455:163–71.
Pobozy E, Sentkowska A, Piskor A. Comparison of three modifi- cations of fused-silica capillaries and untreated capillaries for pro- tein profiling of maize extracts by capillary electrophoresis. J Sep Sci. 2014;37:2388–94.
Lucy CA, MacDonald AM, Gulcev MD. Non-covalent capillary coatings for protein separations in capillary electrophoresis. J Chromatogr A. 2008;1184:81–105.
Lanz C, Kuhn M, Bortolotti F, Tagliaro F, Thormann W. Evaluation and optimization of capillary zone electrophoresis with different dynamic capillary coatings for the determination of carbohydrate- deficient transferrin in human serum. J Chromatogr A. 2002;979: 43–57.
Bodnar J, Hajba L, Guttman A. A fully automated linear polyacryl- amide coating and regeneration method for capillary electrophore- sis of proteins. Electrophoresis. 2016;37:3154–9.
Iwata YT, Mikuma T, Kuwayama K, Tsujikawa K, Miyaguchi H, Kanamon T, Inoue H. Applicability of chemically modified
analytical and bioanalytical methods and in the critical com- parison of CE with other separation techniques. They show some effective strategies how to the overcome the inherent weakness of CE.
13.
14.
Acknowledgments P.M. Nowak gratefully acknowledges financial support received from the Ministry of Science and Higher Education of Poland (Iuventus Plus 2015-2017, grant no. IP2014 033273). The study was carried out using equipment purchased thanks to the European Regional Development Fund within framework of the Polish Innovation Economy Operational Program (contract no. POIG.0 2.01.00-12-0 23/08).
15.
16.
Compliance with ethical standards
17.
Conflict of interest The authors declare that they have no competing interests.
18.
Open Access This article is distributed under the terms of the Creative Commons Attribution 4 .0 International License (http:// creativecommons.org/licenses/by/4.0/), which permits unrestricted use, distribution, and reproduction in any medium, provided you give appropriate credit to the original author(s) and the source, provide a link to the Creative Commons license, and indicate if changes were made.
19.
20.
References
21.
1.
Faller T, Engelhardt H. How to achieve higher repeatability and reproducibility in capillary electrophoresis. J Chromatogr A. 1999;853:83–94.
Mayer BX. How to increase precision in capillary electrophoresis. J Chromatogr A. 2001;907:21–37.
Ali I, Aboul-Enein HY, Gupta VK. Precision in capillary electro- phoresis. Anal Lett. 2006;39:2345–57.
Schaeper JP, Sepaniak MJ. Parameters affecting reproducibility in capillary electrophoresis. Electrophoresis. 2000;21:1421–9.
Štěpánová S, Kašička V. Recent applications of capillary electromigration methods to separation and analysis of proteins. Anal Chim Acta. 2016;933:23–42.
Welch CF, Hoagland DA. Molecular weight analysis of polycations by capillary electrophoresis in a solution of neutral polymers. Polymer. 2001;42:5915–20.
Petersen N, Hansen S. Improving the reproducibility in capillary electrophoresis by incorporating current drift in mobility and peak area calculations. Electrophoresis. 2012;33:1021–31.
Miramon H, Cavelier F, Martinez J, Cottet H. Highly resolutive separations of hardly soluble synthetic polypeptides by capillary electrophoresis. Anal Chem. 2010;82:394–9.
Galbusera C, Thachuk M, Lorenzi ED, Chen DDY. Affinity capil- lary electrophoresis using a low-concentration additive with the consideration of relative mobilities. Anal Chem. 2002;74:1903–14. Fang N, Zhang H, Li J, Li HW, Yeung ES. Mobility-based wall adsorption isotherms for comparing capillary electrophoresis with single-molecule observations. Anal Chem. 2007;79:6047–54.
Souaïd E, Cottet H. Separation of living and dead polymers in synthetic polypeptide mixtures by nonaqueous capillary electro- phoresis using differences in ionization states. Electrophoresis. 2005;26:3300–6.
Lalwani S, Venditto V, Chouai A, Rivera G, Shaunak S, Simanek E. Electrophoretic behavior of anionic triazine and PAMAM dendrimers: methods for improving resolution and assessing purity
22.
2.
23.
3.
24.
4.
25.
5.
6.
26.
7.
27.
8.
9.
28.
10.
29.
11.
30.
12.
31.

Improving repeatability of capillary electrophoresis
4393
Nowak PM, Woźniakiewicz M, Michalik M, Fiedor L, Kościelniak
P. Capillary coating as an important factor in optimization of the off- line and on-line CE-MEKC assays of a highly hydrophobic enzyme chlorophyllase. Anal Bioanal Chem. 2017;409:1493–501.
Erny GL, Gonçalves BM, Esteves VI. Dynamically formed admicelle layer to control the amplitude of cathodic electroosmotic flow. J Chromatogr A. 2012;1256:271–5.
Nowak P, Olechowska P, Mitoraj M, Woźniakiewicz M, Kościelniak P. Determination of acid dissociation constants of war- farin and hydroxywarfarins by capillary electrophoresis. J Pharm Biomed Anal. 2015;112:89–97.
Woźniakiewicz M, Gładysz M, Nowak PM, Kędzior J, Kościelniak
P. Separation of 20 coumarin derivatives using the capillary elec- trophoresis method optimized by a series of Doehlert experimental designs. Talanta. 2017;167:714–24.
Coulter B. Instructions for eCAP™ amine capillary. Beckman Coulter: Brea; 2011.
Nowak PM, Woźniakiewicz M, Kościelniak P. Seven approaches to elimination of the inherent systematic errors in determination of electrophoretic mobility by capillary electrophoresis. Anal Chem. 2017;89:3630–8.
capillaries in chiral capillary electrophoresis for methamphetamine profiling. Forensic Sci Int. 2013;226:235–9.
Pattky M, Huhn C. Advantages and limitations of a new cationic coating inducing a slow electroosmotic flow for CE-MS peptide analysis: a comparative study with commercial coatings. Anal Bioanal Chem. 2013;405:225–37.
Nowak PM, Woźniakiewicz W, Garnysz M, Kościelniak P. A com- parative study of various physicochemically modified capillaries used in CE technique for the three distinct analytical purposes. J
Chromatogr B. 2016;1020:134–41.
Zeisbergerová M, Řemínek R, Mádr A, Glatz Z, Hoogmartens J, Van Schepdael A. On-line drug metabolites generation and their subsequent target analysis by capillary zone electrophoresis with UV-absorption detection. Electrophoresis. 2010;31:3256–62.
Nowak PM, Woźniakiewicz M, Piwowarska M, Kościelniak P.
Determination of acid dissociation constant of 20 coumarin deriv- atives by capillary electrophoresis using the amine capillary and two different methodologies. J Chromatogr A. 2016;1446:149–57. Iwamuro Y, Iio-Ishimaru R, Chinaka S, Takayama N, Kodama S, Hayakawa K. Reproducible chiral capillary electrophoresis of meth- amphetamine and its related compounds using a chemically modified capillary having diol groups. Forensic Toxicol. 2010;28:19–24.
37.
32.
38.
33.
39.
34.
40.
35.
41.
42.
36.