Signature:©PolJRadiol,2010;75(2):73-81
REVIEWARTICLE
Analoganddigitalsystemsofimagingin
Received:2010.03.19
Accepted:2010.03.31
roentgenodiagnostics
DominikaOborska-Kumaszyńska1,2,SylwiaWiśniewska-Kubka3,4
1DepartmentofRoentgenodiagnosticsandUltrasonography,LowerSilesianOncologyCenterinWrocław,Wrocław,Poland2DepartmentofGeneral,InterventionalandNeuroradiology,AcademicClinicalHospitalinWrocław,Wrocław,Poland
3InstituteofBiomedicalEngineeringandInstrumentation,WrocławUniversityofTechnology,Wrocław,Wrocław,Poland4RadiologicalProtectionService,LowerSilesianOncologyCenter,Wrocław,Poland

Author’saddress:DominikaOborska-Kumaszyńska,DepartmentofRoentgenodiagnosticsandUltrasonography,LowerSilesianOncologyCenterinWrocław,PlacLudwikaHirszfelda12,53-413Wrocław,Poland,
e-mail:doborska-kumaszynska@wp.pl
Summary
Intherecentyears,wehavebeenwitnessingaverydynamicdevelopmentofdiagnosticmethodsofimaging.Incontemporaryradiology,thecarrierofthediagnosticinformationistheimage,obtainedasaresultofanX-raybeamtransmittedthroughthepatient’sbody,withmodulationofintensity,andprocessingofdatacollectedbythedetector.Dependingonthediagnosticmethodused,signalscanbedetectedwith analog(x-ray film)ordigitalsystems(CR, DR andDDR). Eachofthesemethodsofimageacquisition,duetoitsowntechnologicalsolutions,determinesadifferentqualityofimaging(diagnosticdata).Theintroductionofdigitalimagereceptors,insteadofconventionalSFsystems,increasedthepatientdose,asaresultofagraduallyincreasingexposure.Thisfollowedfrom thefactthat indigitalsystems, theincreasedradiationdosereducesimagenoiseandimprovesimagequality,andthatisowingtothedatacapacityofthesesystems(impossibleinSFsystemswithalimiteddatacapacityoftheimagedetector).Theavailabilityofthemultitudeofimagingsystems,eachcharacterizedbydisparatequalitativeandquantitativeparameters,impliestheproblemofevaluationandenforcementofaproperefficiencyfrommanufacturersofthesesystems.
Atthesametime,thereisalegalproblempresent inourcountry, i.e.thelackof lawsandregulationsregardingstandardsofthescopeofqualitycontrol(parameters)andmeasurementmethodologyforthesystemsofdigitalimageacquisition.IntheEuropeancountries,thescopeandstandardsofcontrolareregulatedbythemanufacturersandEuropeanGuidelines,whereasintheUnitedStates,AAPMReportshavebeenintroduced,thatspecificallydescribemethodsoftestsperformance,theirfrequency,aswellastargetvaluesandlimits.Thispaperisareviewofboth,thescopeofqualitycontrolparametersofimagedetectorsinanaloganddigitalsystemsofimaging,andthemeasurementmethodology.Theparametersdeterminingtheimagequalityareasfollows:detectionefficiency,dynamicrange,spatialsampling,contrastresolution,spatialresolution,noise,andquantitativedetectionefficiency.Validationofthemeasurementmethods,establishingstandardsofradiographictechniquesfortheperformedexaminations,andcreatingauniformsystemofsupervision,appearstobetheonlywaytoensureaneffectivecontrolofimagingsystemsandtoeliminateanincreasingexposure.
Keywords:
SFRadiography•CRRadiography•DigitalRadiography
PDFfile:
http://www.polradiol.com/fulltxt.php?ICID=878520
Background
withmodulationofintensity,andprocessingofdatacol-lected by thedetector.Depending on the diagnosticmethodused,signalscanbedetectedwithanalog(x-rayfilm)ordigitalsystems(CR,DRandDDR).Eachofthesemethodsofimageacquisition,duetoitsowntechnologicalsolutions,
Inthemoderndiagnosticradiology,thecarrierofthediagnosticinformationistheimage,obtainedasaresultofanX-raybeamtransmittedthroughthepatient’sbody,
73

ReviewArticle
©PolJRadiol,2010;75(2):73-81
Figure1.Cross-sectionaldiagramoftheX-rayfilm.
determinesadifferentqualityofimaging(diagnosticdata).Therefore,qualitycontrolproceduresandtheirscope,aswellasthetestedparametersandevaluationofthedetec-tionefficiencyvarybetweenthesetwodiagnosticsystems.Atthesametime,thereisalegalproblempresentinourcountry, i.e. thelackof lawsandregulationsregardingstandardsofthescopeofqualitycontrol(parameters)andmeasurementmethodologyforthesystemsofdigitalimageacquisition.
Figure2.Definitepointsonthecharacteristiccurve.
valencebandtotheconductionband,whereitcanmovefreelythroughoutthecrystal.Theelectronmovesuntilitreachestheplaceofcrystalstructuredistortion,whereitis‘trapped’.Thepresenceofthe‘trapped’electronwithanegativeelectricchargecausestheattractionofpositive-lychargedsilverions,whichfurtherleadstooriginofametallicsilveratominthatparticularplace.Thisplacebecomesapartofalatentimage.Asaresultofthispro-cess,latentimagesitesmayoccuronthecrystal,largeenough(severaltooveradozenofatomsofmetallicsil-ver)forthedevelopertoinitiatetheprocessofreductionofthewholecrystaltometallicsilver.Theamountofthedevelopedmetallicsilverdeterminesthelevelofblack-ening(opticaldensity),creatingarealimageonthefilm.“Revealing”oftheimageonanX-rayfilmproceedsinthecourseofphotochemicalprocessing.
ImagingSystemsinDiagnosticRadiology
Imagingindiagnosticradiologyisbasedonthreetechno-logicalsolutionsofdiagnosticdataacqusition:
- Analogsystems:cassettewithanintensifyingscreen,x-rayfilm,negatoscope;
- Computedradiography:cassetteswithimagingplates,CRreader,controlstationwithmonitorsfordescription;
- Direct andindirect digitalsystems(DR andDDR)- platesofamorphoussiliconcoatedwithalayerofscintillationmaterial,suchascesiumiodide(CSI)usedinimageinten-sifiersorTFTmatrix(electricallycontrolledswitches,thin-filmcounterpartsofMOSFET-typetransistors),con-trolstationwithmonitorsfordescription.
X-rayfilmhasacertainperformancecharacteristics,theparametersofwhichshould be takenintoaccountwhenset-tingtheexposureconditionsandtheproperdegreeoffilmutilization.Sensitometricpropertiesoflight-sensitivemate-rialsaredeterminedbythecharacteristiccurve(Figure2),
i.e.thegraphoffunctionoftheopticaldensityoverthelog-arithmofexposure,wheretheexposure(EorH)isthe prod-uctofilluminance(I)andexposuretime(t),andtheopti-caldensityisthelogarithmofthequotientofincidentlightintensitytolightintensitytransmittedthroughtheexposedX-rayfilm.Photosensitivematerialdoesnotreactinareasoflowexposure,andtheobtainedopticaldensityresultsfromopticaldensityofabasematerialcoatedwithphoto-sensitivelayers(Dp),andfromopticaldensityoffog(D0),whichaddedalltogether,constitutestheminimumopticaldensity(Dmin).Withacertainvalueofexposure,thevalueofopticaldensitystartsincreasingslowlyandconvertsintothelineardependency withaspecific angleofinclination.
Eachoftheabovementionedsystemsandtechnologicalsolutionsischaracterizedbypropertiesinfluencingimageparameters,i.e.improvingtheirquality,aswellascausingthelossordeteriorationofthediagnosticinformation.
Analogsystems
AnalogimagedetectorofanX-rayunitconsistsofthreecomponents:X-rayfilm,intensifyingscreenandlight-proofhousing(cassette).Anancillarypartofthisimagingsystemisthedevelopingdevicealongwithreagentsforphoto-chemicalprocessingofX-rayfilms(developer,fixingagent,water)andthedryer
X-rayfilmismadeofathin,transparent,flexiblepolyes-terfilm(basematerialofX-rayfilm)andathinlayerofaphotographicemulsion(Figure1).Theemulsioniscoatedwithprotectivepolymerlayersandwithadullingagent.Thephotosensitiveelementsaresilverhalidemicrocrys-tals.QuantaofX-raysorvisiblelightradiationenergyemittedbytheintensifyingscreenstransfertheirenergy,assoonastheyreachthesilverhalidecrystals.Iftheenergyissufficient,crystalelectroncanbetransferredfromthe
Increasingexposurereducestheincreaseofdensityandthecharacteristiccurvereachesplateau,paralleltotheaxisofexposurelogarithm,ataconstantdensity–themaximumdensity(Dmax),whenthelevelofthereachedexposureresultedinthedevelopmentofallcrystalscontainedinthelayer.
Sensitivityindexandcontrastindexarethemainsensito-metricmeasuresderivedfromcharacteristiccurve.
74

©PolJRadiol,2010;75(2):73-81
Oborska-KumaszyńskaDetal–Analoganddigitalsystemsofimaging…
Figure3.Principleofoperationoftheintensifyingscreen-impactonthesensitivityandresolutionoftheimagingsystem.(byAndrewP.Smith,FundamentalDigitalMammography,Physics,TechnologyandPracticalConsiderations.)
Sensitivityindex
Sensitivityofthephotographicmaterialisdefinedasthesmallestamountoflightrequiredtoproduceacertainopti-caldensity.ItisequaltothevalueoflogarithmofexposureneededtoobtainopticaldensitydefinedasDcr=Dmin+1.0,whichdeterminesacorrespondingvalueofthelogarithmofexposure.
Contrastindex
Contrastisdefinedasameasureofdifferencesinopti-caldensityintheimageanditcanbecalculatedfromtheinclinationoftherectilinearpartofthecharacteristiccurve.Itisdefinedasthegradientatapoint(e.g.,con-trastindexa,beingaslopeofatangenttoacharacteris-ticcurveatinflectionpointnearthemiddleoftherecti-linearpart)orasameangradient,definedastheslopeofthesegmentjoiningtwodefinitepointsofopticaldensityD1=Dmin+0.25andD2=Dmin+2.00(Figure2).
Figure4.Structureofanimagingplate.
X-rayfilmisadetectoroffinitedatacapacity.Thus,itisveryimportanttooptimizetheprocessofimagecreation,startingwiththepropersettingoftheunit(exposurecon-trol),throughphotochemicalprocessing(sensitivityofthesystem,artifactsintheimage,noiselevel),thedarkroomlightingconditions,rightselectionofimagingparametersofthewholeset(intensifyingscreenswithintherangeoftheemittedlightwavelength,correspondingtotheparametersofx-rayfilms).Properlysetelementsofdiagnosticdatacre-ationtranslateintothecreationofthemostfavorablecon-ditionsforachievinganimageofaproperquality.
Themainimagingparametersaretheopticaldensity,con-trastandresolution,where:
1.Opticaldensitymeanstheblackeningintheimage,andisdefinedasthevalueofthedecimallogarithmoftheinversetransmissioncoefficient.Thiscoefficientcanbebroughtdowntoaratiooftheintensityofthelighttransmittedthroughtheparticularpointontheimage,totheintensityoftheincidentlightatthatpoint:
Inanalogsystems,thequalitativeanddiagnosticassessmentofanimageisperformedusingnegatoscopesinreadingroomswhichshouldstrictlycomplywithparametersdefin-ingtheappropriatelightingconditions(nomorethan50lux)andconditionsofluminanceoftheemittedlight(cd/m2).
D=log(1/T)=log(Iincident/Itransmitted)
CRsystems
2.Contrastisameasureofdifferencesinblackeningoftheselectedimagefields,correspondingtothedifferencesindensityandthicknessoftissuesshownintheimage.Imagecontrastdependsonthequalityofradiation,thestructureofanexaminedtissueororgan,filmsensitivityandtypeofintensifyingscreen,aswellasthequantityofthescatteredradiationandthedegreeoffilmfogging.
Phosphorimagingplateconstitutesanimagingdetectorindigitizedcomputedradiography(CR).Anessentialcom-ponentofitsstructureisalayerofluminophore(PSP–Photostimulablephosphorimagingsystem)(Figure4).
Imagingplateisplacedinacassettesimilartotheoneusedforanalogradiography;geometryandimagingtechniquearesimilaraswell(Figure5).
3.Imageresolutionisdeterminedbythenumberoflinepairspermillimeter(lp/mm),whichcanbevisualizedandidentifiedasseparatestructures.Resolutiondefinesthesmallestsizeofanobjectpossibletovisualize,andthusdeterminesthesmallestdistinguishabledistancebetweentwoobjects(Figure3).
Patient,sourceofradiation,erasedplate,PSPdetector,exposedplate,imagingplatereutilization,plateerasingprocess,processofimagingplatereadoutanderasing,pre-processing(smoothing,defective(bad)pixelsrepair),pro-cessingoftheimage(filtration,algorithms),reconstructedimage
75

ReviewArticle
©PolJRadiol,2010;75(2):73-81
Figure5.Schemeofimageprocessingandrecordingonimagingplate.
Figure6.Theprocessofimagereadoutfromanimagingplate–point-scansystem.
Figure7.Theprocessofreadingtheimagefromtheimagingplate
–line-scansystem.Schematicdiagramofaline-scanPSPsystem.(AccordingtotheAAPMReportNo.93).
Inthesystemsbasingonphosphorousimagingplates(CR),X-rayradiationquantaareabsorbedbyaphosphorlayeroftheimagingplate(IP).Unliketheconventionalintensify-ingscreens,theenergyofX-rayradiationdepositedinthematerialofimagingplatesisstoredinaportionofenergy,locatedinmetastableregionscalledF-centers.During x-raybeamexposure,thelatentimageisformedinphosphorusbyaccumulation of energy in these centers and revealed bythereaderscanningtheimagingplates.ByscanningtheIPwithanarrowlaserbeamandprocessingthedata,adiag-nosticimageisgeneratedoutofthelatentform.Incaseofpoint-scanreadoutinthescanner(Figure6.),theimag-ingplateismovedinonedirection,whiletheconcentratedlaserbeammovesperpendicularlytothatdirection,fromonesideoftheimagingplatetotheoppositeone.
photomultiplier.Therecordedamountoflightgeneratedbyphotostimulation(PMT)staysinadequateproportiontothenumberofF-centers,andthusalsototheamountofx-rayradiationabsorbedinthatpoint.Photomultiplierconvertsthelightimageintotheanalogelectricsignal,which,ontheoutput,isconvertedintoadigitalsignalbyananalog-digi-talconverter.Beforedigitization,thePMTsignalisintensi-fied,usuallyinanon-linearmanner.Asanextstep,‘raw’signalvaluesareprocessedinsegmentation,rescalingandfilteringprocedures,usingmathematicalalgorithms.
Afterreadingiscompleted,theimagingplateisexposedtoavisiblelightemittedbytheerasinglampthat‘deletes’theX-rayimageandmakestheimagingplatereadyforreuse.
The entire surface ofthe plate is scannedby the laser beamandthelightgeneratedintheprocessofphotostimulationandemittedbyeachpointoftheimagingplate,iscollect-edbyanopticalfiber.Timeofplatereadoutdependsonthesizeofthedetectorandspeedofthereader(averagescanningtakesabout 60–70 s).In modern technology read-ers, a linearlaser source isused(Figure7),which increasesthespeedofdatareadout(averagescanningtimeisabout5–10s).Insuchreadingscanners,thelinearlaserbeamscansthestationaryimagingplate.
Inordertooptimizetheeffectivenessofimagingplateutilizationwithinawiderangeofexposure,thedigitizedsystemsprovidethepre-readingprocedure,whichallowsfortestingthesensitivityofthesignalreading.Initially,aweaklaserbeamisusedforreadinga“raw”imagedata.Appropriatereading,sensitivityandexposureconditionsaredeterminedbasingonanalysesofthedataobtainedinpre-readingprocedure.Theproperreadingfollowstheabovementionedprocedure.Thismethodenablesnormal-izationoftherangeofluminescence,inwhichthex-rayimageappears,inordertoallowtheconversiontodigitalsignals,irrespectivelyoftheobjectbeingtestedandtheX-rayradiationdose.
Atomsreturningtoequilibrium(groundstate)generatetheemissionofphotonsfromthespectrumofthevisiblelightwavelength,andthisphenomenonisrecordedbya
76

©PolJRadiol,2010;75(2):73-81
Oborska-KumaszyńskaDetal–Analoganddigitalsystemsofimaging…
Figure8.Thin-filmtransistorstructure(accordingtoMEDTECH).
Figure10.Directdigitalsystemdetector:(A)microphotography,
1. thephysicalstructureofasinglepixelofTFTmatrix
2. aschematicdiagramofthestructureoftwopixels(accordingtoMammographicdetectors,G.PANAGIOTAKISUNIVOFPATRAS).
InCRsystemsofdigitalradiographythephenomenonoffading(lossoftherecordedsignal)occurs,whichisdisad-vantageousforimageacquisition.Hence,thetimeframebetweentheexposureoftheimagingplateandthereadingprocessiscrucial.Anaverageimagerecorderlosesabout
Figure9.Constructionofdigitaldetectors–directandindirectdetection(accordingtoMEDTECHandPerformanceAssessmentofDRSystems,JASeibert,Ph.D.UCDavisMedicalCenterSacramento,CA,http://www.aapm.org/meetings/03AM/pdf/9787–48370.pdf).
25%ofthedepositedsignalbetween10minutesand8hoursperiodsubsequentlytotheexposure.Afterthattimeframe,thefadingslowsdown.
higherresolutionoftheimagingsystem.ThethicknessoftheCSIcrystalwithitsneedle-likestructurecanbeadjustedtothedesiredsensitivity ofthesystem(ensuringproperlevel ofabsorbanceofx-rayradiation),withthemaintenanceofhighspatialresolutionatthesametime.Photodiodes(Si:H),locat-edunderalayerofscintillationmaterial,converttheopticalsignalintoanelectricalsignal(charge),whichisaccumulatedinacapacitiveelementofapixel.
DRandDDRDigitalSystems
Inradiographywithdigitalimagedetectors,themostcom-monsolutionarepanelsmadeofamorphoussiliconorselenium(indirectdigitalsystems)andpanelsbasedonamatrixmadeofelectrodesseparatedbyalayerofinsula-torandtheactivecomponents,suchasthin-filmtransistor(Figures8,9).
Inthedirectdigitalimagingsystem,thedetectorismadeofphotoconductorscharacterizedbyahighatomicnumber(e.g.,SeorPbI2).Theycoveranactiveareaofthematrix.ThatkindofstructureformsalayerofphotoconductorswhichdirectlyconvertsX-rayradiationincidentsintochargecarriers,driftingtothecollectingelectrodes.Themainadvantageofdirectdigitalsystems,comparingtoCRsystemsandindirectDRsystems,intermsofimagequal-ity,isthelackofeffectsfromthelightphotonsscatteringatthedetectormaterial.Electriccharge,generatedastheeffectofx-raysradiation,iscollectedbyasingleelectrode,whichmakestheside-scatter(diffusion)effectnotsignifi-cantfortheprocessofimagecreation.Additionally,detec-torabsorption efficiency can be maximized by an appropri-ateselectionofthematerialofphotoconductors,calibra-tion,andaproperthicknessofthelayerofcapacitiveele-ments(Figure10).
Detectors,whichconsistofphotoconductors,suchasamor-phoussiliconorselenium,constitutethemainpartofindi-rectdigitalsystems.Layerofsilicondetectorcontainsamatrixofreceptors,eachequippedwithitsowncontrolcom-ponent(transistorordiode)andcorrespondingtoonepixeloftheimage.Regulating(control)systemsareresponsiblefortheprocessofdatareading:lineafterline,electricalsig-nalsareintensifiedandconvertedintoadigitalform.SiliconitselfisnotsufficientlysensitivetoenergyofX-raysradia-tionusedindiagnosticimaging.Therefore,siliconlayeriscoveredwithalayerofscintillationmaterial,suchascesiumiodide(CSI).CSIischaracterizedbyaneedle-likestructureofacrystal,causinglessside-scatteringoflightandensuring
77

A	B

C

ReviewArticle
©PolJRadiol,2010;75(2):73-81
betweenadjacentpositionsofthelaserbeamduringreadingprocessfromtheimagingplate.InDRsystems,pitchisthedistancebetweencentersofthespacessepa-ratingeachofdetectingelements.Thespatialfrequencyinsampling,determinesthedigitalsystem’sabilitytodisplayhigh-frequencyfluctuationsinX-raystream.IftheX-radiationstreaminteractswithreceptor,gener-atingthedataoffrequencieshigherthantheNyquistfrequency,andmodulationtransferfunction(MTF)forthesefrequencies,priortosampling,isnotfading,thenfalsesignal can appear inthe image at low frequencies.
Figure11.Thesensormatrixconstructionindisplaysandthemethodsofreadoutcontrol(accordingtoMedtech).
SpatialResolution–istheabilityofimagingsystemtovisualizetwoadjacentstructuresasseparateimageele-ments,oraclearedgemarkingintheimage(sharpness).Lossesinthespatialresolutionoccurduetoblurringcausedbygeometricfactors(e.g.,sizeofafocus,scat-teringoflightinthereceptor),theeffectiveareaofthedetectordeterminedbythesizeofaperture,patient’smovementsinrelationtothesourceofX-radiation,imagedetector,thethicknessofthedetectorelements,screen,CSIcrystalthicknessanddensityof datareading.
AnactivematrixconsistsofMxNnumberofpixels.Eachpixelhasthreebasicelements:theTFTswitch,pixelelec-trodeandcapacitor.Activematrixisdeterminedbythepixelwidth(a),widthofpixelcollection(acoll)andthedis-tancebetweenpixels(pitch)(d)(Figure10).
TFTelementsfunctionasswitches,foreachpixelindividu-ally,andcontrolthecharge.Eachlineofpixelsissimul-taneouslyelectronicallyactivatedduringthereadingpro-cess.Normally,allTFTelementsaredeactivated,allow-ingtheaccumulationofthechargesonpixelselectrodes.Datacanbeobtainedbyexternalelectronicsandcontrol-lingoftheTFTstatusbysoftware.EachTFTcontainsthreeelectricalcomponents:Gate(G)controlling“on”or“off”TFTstatus,Drain(D)connectingthepixelelectrodeandthepixelcapacitorandSource(S)connectedtoacol-lectivedatatransmissionline.Whenthegatelineisacti-vated,alltheelementsofTFTinaparticularroware‘on’andthechargecollectedontheelectrodesisreadfromthedataline.Paralleldataaremultiplexedintoserialdata,dis-cretizedandtransferredtoacomputertocreatetheimage(Figure11).
Noise–itcanbedefinedasfluctuationsintheimage,whichdonotcorrespondtodifferencesinX-radiationabsorptionbyobjects.Ameasureofnoisemaybedeter-minedbyestimatingthenoisepowerspectrum(NPS),whichdescribesthecorrelationofspatialfrequencyandnoise.Thenoiseintheimageisdominatedbyquantum(shot)noiseresultingfromquantumfluctuationsintheX-rayanddatadigitization(incaseofdigitalsystems).However,allimagereceptorscontaininternalsourcesofnoise,suchasnoisecomingfromthefilmgrainandelectronicnoiseintheCRandDRsystems.Theinternalnoiseofthedetector,withestablishedcorrelationcor-respondingtoplaceonthereceptor,resultsfromaspa-tialdifferenceintheintensifyingscreenthickness,inSFsystem,fromlightdetectionefficiencydependentonthelocationonplatereadersinCRsystems,anddifferencesinpreamplifyinginDRsystems.Oneoftheadvantagesofdigitalimagingisthattheinternalnoiseofdetectorcanbedigitallyeliminatedbypost-processing(obtainingvaluablediagnosticimage).Deteriorationoftheimageinradiographyisalsoconditionedbythescatteringofradiation,whichisanothersourceofnoiseandcontrib-utestodecreasingofimagecontrast.Thesolutiontothisproblemistheuseofanti-scattergridsplacedinfrontofthe imagedetector.Using thegrid is especially importantinCRsystems,duetoanincreasedsensitivitytoscat-teredradiationofbariumhalide(K-edgeisapproximately35keV),inrelationtotheSFscreenswhichcontaingado-liniumoxidesulphide(Kedgeisapproximately50keV).Inthescanningsystems(fieldslot-scan)ontheotherhand,DRdetectorsareableto“cut”thescatteredradia-tionoutfromrecordingandthereforedonotrequiretheuse of anti-scatter grid.
Anindisputablebenefitofdigitalimageacquisitionisthepossibilityofpost-processing.
FactorsDeterminingImageQuality
DetectionEfficiency–intheSF(“screenfilm”),CR(phosphorusimagingplates)andDR(digitalsystems)imagingsystems,quantumefficiencyisdeterminedbythickness,densityandstructure(composition)oftheabsorber(imagedetector).
DynamicRange–forhighqualityimagingindigitalradiography,imagedetectormusthaveagoodcontrastresolutioninawidespectrumofX-radiationexposureintensity.Dynamicrangeoftheimagingsystemisdeter-mined bytheratioofthelargestand smallestinputintensities,whichcanbevisualized.Thesmallestuse-fulvalueofintensityislimitedbythenoiselevelofthesystem,whilethehighestvalueofintensitydependsondetectorsaturationlevel.
Contrastresolution–referstothevalueofthesignaldifferencebetweentheexaminedstructureandthesur-rounding.ItistheresultofdifferencesinX-rayabsorp-tionintheexaminedtissues.Itisexpressedasarela-tivedifferenceinbrightnessbetweentherelevantareasinthedigitalimageshownonthemonitor.Radiographic
SpatialSampling–alldigitaldetectorssampletheper-manentlyfluctuatingstreamofX-raysattheinput,atdiscretelocations,separatedbygaps(pitch).InCRsys-tems,thespacingbetweensamplesisthedistance
78

©PolJRadiol,2010;75(2):73-81
Oborska-KumaszyńskaDetal–Analoganddigitalsystemsofimaging…
contrastisdeterminedbythecontrastoftheobjectandreceptorsensitivity.ItisstronglyreliantontheenergyspectrumofX-rayradiationandthescatteredradiationphenomenon.However,indigitalimaging,contrastintheimagecanbechangedbysettingthevisualizationparameters,independentoftheacquisitionconditions.
outbyplacing“loaded”cassettes,sidebyside,inafullyilluminatedroom.Theyallshouldbedevelopedafter15minutes.Thereshouldbenodark,indicatingwornseals,areasonthefilm.Faultycassettesshouldbereplaced.
2.Developingprocessinspection(darkroomqualitycontrol):thepurposeistocontroltherepeatabilityofphotochemicalprocessing,provedbytheconstancyofobtainedimageparameters.Threebasicparametersaretakenfromthecharacteristiccurve(Fig.2):themini-mumdensityDmin,sensitivityS,andameangradientg(g=(D1–D2)/(P1–P2)*0.15,Art.2a).MeasuredvaluesshouldbewithintherangeoftolerancespecifiedintheJournalofLaws,No.194,item1625.
Detectionquantumefficiency(DQE)isaparameterthatdescribestheimagereceptor,consideringitseffi-ciencyofradiationdetection,spatialresolutionandnoiselevel.DQEdescribestherelativeefficiencyinmain-tainingSNRlevel(signal–to–noiseratio),possibletoachieveintheimagingprocess.ThisparameterisdefinedasSNR2out/SNR2in, where SNR2inmeanssignal– to–noiseratioofexposureimpactonreceptorandnumeri-callyequivalenttotheinputstream.Thus,DQEcanbedescribedastheefficiencyofSNRtransmissionbythesystem,theeffectivenessofwhichtranslatesintothequalityofdetectionandimageacquisition.
3.Filmstorageconditions:thistestisperformedusingathermometerandhygrometertomeasurethetemper-atureandhumidityinthefilmstoragearea.Measuresshouldbeaccordingtothemanufacturer’sinstructions,placedonpackagingof the films.
Quality Assurance in Diagnostic Imaging
4.Film-screencontacttest:thistestiscarriedoutforallthecassettes,usingaspeciallydesignedphantom(filmsshouldbeloaded inacassette atleast15 minpriorto thetestandthecassettesshouldbeplacedinauprightposi-tion).CassettesshouldbeplacedwithintheX-raybeam,andcoveredwiththewiremesh(thewholesurfaceshouldbecovered).Conditionsoftheexposureareselect-edmanuallyinordertoachievetheopticaldensityvalueconsistentwithrecommendationofthemanufacturerofthetest.Thenumberandsizeofdarkareasshouldbechecked(placesofnofilm-screencontact).Acceptablesizeofthe“nocontactareas”isgiveninappropriateregulations(MinisterofHealthRegulations;JournalofLawsNo.194,item1625).
Assuringqualityinimagingisaprocessofdeliberateactinginordertoobtainthebestpossiblediagnosticimages,thatenableanappropriateclinicalevaluation.Itisimportantfortheimagetocontainenoughdatawiththeleastpatient’sexposuretoionizingradiation.Tofulfilltheseobjectives,acontinuousmonitoringofthediagnosticimagingprocessisnecessary,inaccordancewiththerequirementsoftheregu-lationsandqualitystandardsformedicalequipment.Forthispurpose,anappropriateacceptancetestingiscarriedout,aswellasroutineperformancetestsonabasicandspecialistlevel,controlling,analysesoftherejectedimages,andmonitoringoftheimageacquisitionprocess.
Acceptance testing,performedto verifywhether theequip-mentandroomsmeetthetechnicalrequirements,provideanutterandobjectiveassessmentoftheproceduresandequipmentcondition.However,themainpurposeofrou-tine performance evaluation (basic and specialist) is tocer-tifywhethertheequipmentandproceduresaremaintainedinan optimaland constant range.
5.Uniformsensitivity:inSFsystems,atestverifyingthesensitivityofintensifyingscreens(cassettes)shouldbeperformed.Thetestisbasedonimagingofahomo-geneousphantomwithautomaticallysetexposureinclinicalconditions.Opticaldensityvalueatthepointofreferenceismeasuredontherawimage.Evaluationofuniformityofthesensitivitybasesonacomparisonbetweenopticaldensityvaluesobtainedforeachcas-sette.DifferencesshouldnotexceedtherangespecifiedinMinisterofHealthRegulations(JournalofLawsNo.194,item1625).Atthesametime,differencesinval-uesoftheinputairkermaforeachcassetteshouldnotexceed5%.
Testsconnectedwithimagedetector,inallthreesystemspresentedabove,include:
Inanalogradiography:cassettetightness,inspectionsofthedevelopingprocess(darkroomqualitycontrol),filmstorageconditions,film-screencontacttest,anti-scattergrid,high-contrastresolution,low-contrastresolution,uniformsensitivityofintensifyingscreens.
6.Anti-scattergrid:attenuationcoefficientoftheanti-scattergridisassessedbymeasuringtheairkerma,plac-ingthedetectorinthefrontofthegrid(inrelationtothesource)andbehindthegrid,takingintoaccountacorrec-tionrelatedtoopticaldensityobtainedfortheexposedfilm placed inthe front andbehind thegrid, in relationtothesource.Theratioofthesemeasurements,takingintoaccountthecorrectionformeasurementdistanceinbothconfigurations,determinestheattenuationofx-raysradiationpassingthroughthegrid.Thecoefficientofthegridshouldbelessthanthe3.Thegridisalsotestedvisually,foranystructuraldamageorshiftdisturbance(incaseofmobilegrids),whichcancauseimageartifacts.
In computed radiography: noise (signal – to– noise ratio,SNR),uniformity(imagingplatetest),anti-scattergrid,erasurethoroughness,high-contrastresolution,low-con-trastresolution,defectivepixelsanalysis(badpixelmap),uniformsensitivityoftheplates,fading,imagingplatethroughput,linearityandautomaticresponse.
AnalogRadiography
1.Cassettetightnesstest:allclinicallyusedcassetteandthetestingcassettearesubjecttothistest.Itiscarried
79

ReviewArticle
©PolJRadiol,2010;75(2):73-81
7.Resolutionoftheimagingsystem:atestisperformedusingaresolutionphantom,whichisplaceddirectly(ascloseaspossible)ontheimagedetector.Whenassessingtheresolutionofthesystem,itisdesirabletouseseveralphantomsinthefollowingconfiguration:inthemiddleofthebeamaxisandontheedgesofthedetector.Itenablesustomeasuretheresolutionforanysizeofthedetector,consideringimagingwithinthebeamaxisandedgecondi-tions(offtheaxis).Thevalueoftheresolutionshouldbeclosetothevaluegiveninthemanufacturer’sspecification.
severalphantomsplacedinthemiddleofthebeamaxisandontheedgeofthedetectorenablesthemeasurementofresolution for any kind and size of the detectorin bothdirectionsofscanning).Themeasuredresolutionshouldbeapproximatetothevaluespecifiedbythemanufac-turer.Iftheresolutionvaluedeviatesby10%fromthespecifiedvalue,thesystemmustbecorrected.Resolutionof the system is affected not only by the size of the detec-torbutalsothealgorithmofhigh-contrastprocessing.Indigitalimagingsystems,theresolutiontestshouldbeperformedfortheclinicallyusedsettingsofexposure,andimageformatsusuallyapplied.Resolutionassess-mentisbasedontheMTFanalysis(objectivemethod)orvisualassessmentdefinedasthenumberofpairs/mm(subjectiveassessment).ResolutioninCTsystemsisalsodeterminedbythesizeofthelaserbeamcrosssection,aswellasoscillationsandfocusingofthelaser.
ComputedRadiographyandDigitalRadiography(CRandDRSystems)(jointtests):
1.Noisetest:basedonthemeasurementofameanpixelvalueandstandarddeviationinthereferenceROI(RegionofInterest)foranunexposeddetector.Incaseofsystemswithalinearresponse,SNRshouldbeesti-matedandplacedonagraphoffunctionofSNR2overinput surface kerma(it canbe determined bytests ofthedetectorresponsefunction).Obtained measures(points)allowtodrawalineofbestfit(linearregressiontrend)inordertocalculateacorrelationcoefficient.Noisetestshouldbecarriedoutfordifferentanode-filterconfigura-tionsusedinclinicalpractice.Thederogationfromlin-earityisameasureofquantumnoise(systemelectron-ics).Forlogarithmicsystems,thisgraphshalldefinetherelationshipbetweenstandarddeviationandinversedinputsurfacekermaoftheair.Forthemeasurementpoints,a line of best fit and correlation coefficient shouldbedetermined.Shiftofthepointintersectionwiththekermadomainaxisonthegraphisameasureofquan-tumnoiseinthesesystems.Theresultsofbothsystemsshouldbecomparablewithpreviousresults.
6.Low-contrastResolution:thetestisperformedforeachsizeofthedetector,usingphantomwithlow-con-trastelements.Threeexposuresforthreedifferenttimeperiodsshouldbeperformed.Visibilityofthelow-con-trastelementsistobeevaluatedandcomparedwiththeresultsofbaselinemeasurements.Intheprocessofimageacquisition,mathematicalhigh-contrastalgorithmoredgeimprovementalgorithm(softwarefilter)shouldbeapplied.Astheresultofthattestdependsonthetypeofphantomandobjectiveassessmentofthetester,theassessmentofseveralimagesorbyseveraltestersshouldbeimplementedasastandardroutine.
7.Defectivepixelstest(Bad pixelmaps):Anevalua-tionoftheseartifactsiscarriedoutusingraw,“pre-pro-cessed”image–thisimageisinspectedforthepresenceofuncorrectedbadpixels(hotandcold).Eachcountryenforcesitsownregulationsregardingtheamountofbadpixelsintheimage.Informationontheuncorrecteddefectivepixelsshouldbespecifiedbythemanufactur-er,andtheirnumberinthetestshouldnotchangewhencomparedtotheoriginal.Inthistest,imageisobtainedasaresultofexposureofahomogeneoustestphantomunderclinicalconditions.TheanalysisoftheresultscanbeconductedviavisualassessmentorsearchingtheROIofsignaldeviationexceeding20%.
2.Uniformity:inordertoassesstheuniformityofthedetec-torresponse,animageofahomogeneoustestphantomshouldbeperformed,coveringthewholeareaofimaging.Exposureparametersinthistestshouldcorrespondtoclin-icalconditions.Ontheobtainedrawimage,themeanpixelvalueandstandarddeviationinthereferenceROIshouldbecalculated.Theseparametersshouldbealsoevaluatedforareasofreadout localizedindifferentpartsofthedetec-tor.Themeanpixelvalueforthewholeimageshouldbecalculated,aswellasSNRforallROIs.ResultsobtainedforeachROIshouldbecomparedwithmeanvalues.Maximumacceptableabberationshouldnotexceed15%.
ComputedRadiography
1.Uniformsensitivity:inCRsystems,asensitivitytestofimagingplatesshouldbecarriedout.Thetestisbasedonhomogeneousphantomimagingunderautomaticallyselectedexposureinclinicalconditions.ThemeanpixelvalueandstandarddeviationinROIofreferenceismea-suredonthe‘raw’image.Theassessmentoftheuniformsensitivityofutilizedimagingplatesisbasedonacom-parisonofSNR.Differencesshouldnotexceed10%.Atthesametime,thevaluesoftheinputairkermaobtainedinthetestforeachplateshouldnotdifferbymorethan10%.
3.Anti-scattergridtest:performedasinanalogueradi-ography.Forsystemswithdirectdigitization,thetestisnotperformed.
4.Erasurethoroughness/DRdetectorreset:ThetestappliestoallavailableimagingplatesandDRdetectors,whichshouldbeerased(reset)inafull cycle,togetridofanysignalsthatcauseartifacts.Thedetectorpreparedinsuchawayshouldbescannedinordertocheckwhetherthereceivedimageis“clean”Thewholeproceduremustberepeated,ifartifactsor remnants ofthe previousimageappearonanyofthedetectors.
2.Evaluationoffadingphenomenon:Thistestisper-formedonimagesobtainedfromahomogeneousphan-tomexposureinthe clinical conditions.These imagesaresubjecttoacquisitionindifferenttimeintervalsafterplateexposure(1min,2min,5min,10min,30min;).
5.Resolutionoftheimagingsystem:resolutionphantomisused,whichisplaceddirectlyonthedetector(using
80

©PolJRadiol,2010;75(2):73-81
Oborska-KumaszyńskaDetal–Analoganddigitalsystemsofimaging…
Forimagesobtained inthismethod, themeanpixelvalueinROIofreferencecanbedetermined.Theresultsaretocomparewiththereferencevaluesoftheacceptancetests.
performed,withreadoutdelayof10minutes.Thewholeprocess shouldberepeated 3times.The calculatedvalueoftheexposuretimeshouldbewithin20%(foreachplate)inrelationtothevaluementionedinthespeci-ficationorwithin10%,ascomparedtothemeanvalueforalltested plates.Measuringrangeusedtoevalu-atethedetectorresponsefunctionfallsbetween0.1ofthestandardexposuretimeandquintupleofthattime.Whileanalyzingtheresults,agraphofinputsurfacekermafunctionoverameanpixelvalueofROIshouldbedrawn.Correlationcoefficientoftherelationbetweentheregressionlineandrealvaluepointsshouldbehigh-erthan0.99.Thepointofintersectionwithmeanpixelvalueaxis(Y)determinestheoffsetofthepixelvalue.
3.Phosphorimagingplatethroughput:inthistest,4platesofequalsizeshouldbeexposed,withastandardgeometryof thesystem. Plates,exposedasdescribed, aretobeputonebyonetothereader(withouttimedelay).Throughputrateshouldamounttoapprox.10%(foreachplate)of thevalue given inthe specification.
4.Thelinearityandautomaticresponse:thistestdeter-minestheresponseofthedetectorandinformationretrievalsystemtothechangesinexposure.Thetestisconductedbysetting3differenttimesofexposureandmeasuringtheairkerma,whichissubsequentlyrecal-culatedrelativelytothesurfaceofthedetector.Foreachexposure,threeindependentimagesshouldbe
Inaddition,inCRsystems,ananalysisoftrendsforthevaluesofthemeasuredparameters(especiallyforlog-linearconfiguration)mustbeperformed,referringtothevaluesofthebasictests.
References:
1. AAPMREPORTNO.93,Acceptancetestingandqualitycontrolofphotostimulablestoragephosphorimagingsystems,2006
2. AAPMREPORTNO.96,Themeasurement,reportingandmanagementofradiationdoseinCT,2008
3. AAPMREPORTNO.116,,Anexposureindicatorfordigitalradiography,2009
4. AAPMREPORTNO.74,Qualitycontrolindiagnosticradiology,2002
5. PruszyńskiB:,Diadnostykaobrazowa.Podstawyteoretyczneimetodykabadań.PZWL,Warszawa2001
1. KowskiR,KubasiewiczM:Mammografia–podręcznikzachowaniastandardówjakości.WydawnictwoLekarskie,ACR,Warszawa2001
2. PanagitakisG:Mammographicdetectors.http://www.hep.upatras.gr/class/download/bio_sim_eik/mammographic_detectors.pdf
3. Practiceguidelinefordigitalradiography,ACRpracticeguideline,2007(Res.42)*
4. SeibertJA:PerformanceAssessmentofDRSystems,UC
DavisMedicalCenterSacramento,CA.http://www.aapm.org/meetings/03AM/pdf/9787-48370.pdf
81