©PolJRadiol2018;83:e437-e445
DOI:https://doi.org/10.5114/pjr.2018.79651
Received:28.05.2018
Accepted:29.05.2018
Published:17.09.2018
http://www.polradiol.com
Originalpaper
Reproducibilityofintravoxelincoherentmotionofliver
ona3.0Tscanner:free-breathingandrespiratory-triggeredsequencesacquiredwithdifferentnumbersofexcitations
AndrzejCieszanowski1A,B,D,E,F,KatarzynaPasicz2A,B,C,D,E,F,JoannaPodgórska1,3A,B,C,D,E,F,EwaFabiszewska2A,B,D,
WitoldSkrzyński2A,C,D,E,F,JakubJasieniak1A,B,C,D,F,AgnieszkaAnysz-Grodzicka1A,D,IwonaGrabska2B,D,JakubPałucki1A,D,MartynaNaduk-Ostrowska1A,B,D,BeataJagielska4A,D,PawełKukołowicz2A,B,D,E
1DepartmentofRadiologyI,MariaSkłodowska-CurieMemorialCancerCenterandInstituteofOncology,Warsaw,Poland2MedicalPhysicsDepartment,MariaSkłodowska-CurieMemorialCancerCenterandInstituteofOncology,Warsaw,Poland32ndDepartment ofClinicalRadiology,MedicalUniversityofWarsaw,Warsaw,Poland
4DepartmentofOncologicalDiagnosisandCardio-Oncology,CurieMemorialCancerCenterandInstituteofOncology,Warsaw,Poland
Abstract
Purpose:Tooptimisetheintravoxelincoherentmotion(IVIM)imagingoftheliverona3.0Tscannerbyassessingparameterreproducibilityonfree-breathing(FB)andrespiratory-triggered(RT)sequencesacquiredwithdifferentnumbersofsignalaverages(NSA).
Materialandmethods:Inthisprospectivestudy20subjects(M/F:10/10;age:25-62years,mean:39years)underwentIVIMmagneticresonanceimaging(MRI)ona3.0Tscannerusingan18-channelphase-arrayedcoilandfourdiffer-entecho-planarsequences,eachwith10bvalues:0,10,30,50,75,100,150,200,500,and900s/mm2.ImageswereacquiredwithFBandRTwithNSA=1-4(FBNSA1-4,RTNSA1-4)andwithNSA=3-6(FBNSA3-6,RTNSA3-6).Subsequently,fortheassessmentofreproducibilityofIVIM-derivedparameters(f,D,D*),eachsubjectwasscannedagainwithanidenticalprotocolduringthesamesession.IVIMparameterswerecalculated.ThedistributionofIVIM-parametersforeachDWIsequenceweregivenasthemedianvaluewithfirstandthirdquartile.Inter-scanreproducibilityforeachIVIMparameterwasevaluatedusingcoefficientofvarianceandBland-Altmandifference.DifferencesbetweenFBsequenceandRTsequenceweretestedusingnon-parametricWilcoxonsigned-ranktest.
Results:Meancoefficientofvariance(%)forf,D,andD*rangedfrom60to64,from58to84,andfrom82to99forFBNSA1-4sequence;from50to69,from41to97,andfrom80to82forRTNSA1-4sequence;from22to27,15,and
from70to80forFBNSA3-6sequence;andfrom21to32,from12to,andfrom50to80forRTNSA3-6sequence,respectively.
Conclusions:IncreasingthenumberofsignalaveragesforIVIMacquisitionsallowsustoimprovethereproducibilityofIVIM-derivedparameters.Thesequenceacquiredduringfree-breathingwithNSA=3-6wasoptimalintermsofreproducibilityandacquisitiontime.
Keywords:diffusion-weightedimaging,b-value,pseudo-diffusion,liver,intravoxelincoherentmotion.
Correspondenceaddress:
JoannaPodgórska,MD,PhD,DepartmentofRadiologyI,MariaSkłodowska-CurieMemorialCancerCenterandInstituteofOncology,5RoentgenaSt.,02-781Warsaw,Poland,phone:+48225462280,fax:+48225462307,e-mail:jpodgo@gmail.com
Authors’contribution:
AStudydesign∙BDatacollection∙CStatisticalanalysis∙DDatainterpretation∙EManuscriptpreparation∙FLiteraturesearch∙GFundscollection
©PolJRadiol2018;83:e437-e445
e437

AndrzejCieszanowski,KatarzynaPasicz,JoannaPodgórskaetal.
DWIoftheliverona3.0Tscannercouldbebene-
ficialdueto improvedSNR comparedto1.5Tscanners;however,prominentmagneticsusceptibilityartefactsanddistortionrelatedtoeddycurrentsremainachallengeandmayleadtosignificantdegradationofacquiredimages.DespitesomepublicationsregardingimplementationofIVIM-derivedparametersforthediagnosisofliverdis-ease, therehavebeennoreports addressingthe reproduc-ibilityofthistechniquewithdifferentrespiratoryschemesat3.0TMRsystems.
Tothebestofourknowledge,therehasbeenonestudythataddressedtheeffectofthenumberofsignalaverages(NSA)ontheoverallqualityofIVIMimages,SNR,andstandarddeviation(SD)ofmeasurements[11],butnonethatfocusedontheeffectonthereproducibilityofthecalculations.WehypothesisedthattheuseofhigherNSAcouldimproveSNR,decreaseSDofmeasurements,andpossiblyincreasethereproducibilityofcalculatedIVIMparameters.Therefore,theaimofthisstudywasto opti-miseIVIMimagingoftheliverona3.0TMRplatformbyassessingthereproducibilityofIVIMparameters(f,D,D*)forFBandRTsequencesacquiredwithdifferentnumbersofsignalaverages.
Introduction
Therehavebeenseveralattemptstouseintravoxelin-
coherentmotion(IVIM)diffusion-weightedimaging(DWI)forthecharacterisationoffocalliverlesionsandforthediagnosisofdiffuseparenchymalliverdisease[1-5].However,duetosignificantdifferencesinreportedperfu-sion-relatedIVIMparameters indisease-freesubjects,itisstillnotpossibletoestablishnormalvaluesofperfusionfraction(f),whichreflectsthevolumeofflowingblood(rangeinpublishedstudies:from19.0%±5.5%to32.16%
±8.13%)andthepseudo-diffusioncoefficient(D*),whichispresumedtoshowthevelocityofcapillaryblood(range:from39.61×10–3mm2/s±12.34to110.6×10–3mm2/s±79)[2-6].Onlythevaluesofpuremoleculardiffusion(D)weresimilarinseveralpublications,allowingmorereliablecomparisonofthisparameterbetweenstudies.Moreover,IVIM-derivedparametersandmapsarehighlyvariableandinconsistent[7].ThisinconsistencyofIVIMtechniquecouldberelatedtodifferentfactors,includingvariousdis-tributionofappliedbvalues,useofdifferentdiffusiongra-dientpolarity(monopolarvs.bipolargradients),orseveralmotioncorrectionmethods,suchasfree-breathing(FB),breath-holding(BH),andrespiratory-triggering(RT),andpost-processingfittingmethods[4,8,9].Furthermore,ac-quiredimagesusuallyarelowquality,regardlessofthetypeofappliedsequence. The other important possiblecauseofdiverseresultsaswellaspoorqualityofIVIMimagesmaybelowsignal-to-noiseratio(SNR).PoorSNRmayleadtohigh,non-physiologicalheterogeneityamongvox-els,precludingreliablemeasurements.Eventhoughsev-eralstudies ontheuseofIVIMforassessmentofhepaticdiseasehavebeenpublished,onlyafewreportsaddressedthereproducibilityofIVIM-derived parameters[1,9,10].LackofstandardisationofIVIMtechniqueandsignificantvarianceincalculatedparametersamongstudies,thusfarhampertheapplicationofthismethodforroutineMRliv-erimagingandforapplicationofIVIM-derivedparame-tersaseligiblebiomarkers.
Materialandmethods
ThestudywasapprovedbytheInstitutionalReviewBoard,
andallvolunteersparticipatinginthestudysigned anin-formedconsentform.
Studypopulation
BetweenJanuary2017andMarch2017,20healthyvol-unteerswereenrolledforthissingle-centreprospectivestudy.Inclusioncriteriawereage≥18yearsandwilling-nessandabilitytoundergoMRIandparticipateinthestudy.ExclusioncriteriawerecontraindicationstoMRimagingsuchaspacemakers,metalimplants,andsevereclaustrophobia;age<18years.
Table1.Numbersofsignalaverages(NSA)ofdiffusion-weightedimagingsequences
Magneticresonanceimaging
AllMRIstudieswereperformedata3Tunit(MagnetomSkyra,SiemensMedicalSolutions,Erlangen,Germa-ny)usingexplorergradients(maximumamplitudeof45mT/mandslewrateof200mT/m/ms)andphased-ar-raymulticoilsystem(18elements).
Theexaminationprotocolbeganwithbreath-hold(BH)T1-weightedgradient-echo(GRE)DIXONsequenceandrespiratory-triggered(RT)T2-weightedHASTEsequence,bothperformedintheaxialplane.Thenfourecho-pla-nardiffusion-weightedimaging(DWI)sequencesintheaxialplanewereacquiredwithvariableNSA.Twoofthemwithsinglesignalaverageforthefivelowestbval-uesandfoursignalaveragesforthehighestb:FBNSA1-4,RTNSA1-4,andtwootherswiththreesignalaverages
FB–free-breathing,RT–respiratory-triggering
e438
©PolJRadiol2018;83:e437-e445
b[s/mm2]	FBNSA1-4	RTNSA1-4	FBNSA3-6	RTNSA3-6
0
1
1
3
3
10
1
1
3
3
30
1
1
3
3
50
1
1
3
3
75
1
1
3
3
100
2
2
4
4
150
2
2
4
4
200
2
2
4
4
500
2
2
4
4
900
4
4
6
6

ReproducibilityofIVIMofliverat3.0T
Table2.Acquisitionparametersofapplieddiffusion-weightedimagingsequences
NSA–numbersofsignalaverages,FB–free-breathing,RT–respiratory-triggering
Table3.Meanregionsofinterestsizesfordiffusion-weightedimagingsequences
NSA–numbersofsignalaverages,FB–free-breathing,RT–respiratory-triggering
forthefivelowestbvaluesandsixsignalaveragesforthe
highestb:FBNSA3-6,RTNSA3-6(Table1).Selectedpa-rametersofappliedDWIsequencesareshowninTable2.AllDWIsequenceswereacquiredwiththesame10bval-ues:0,10,30,50,75,100,150,200,500,900s/mm2but
differentNSA(NSAforeachbvalueforallDWIsequenc-esarepresentedinTable1).
Afterconclusion ofthewholeMRimagingprotocol,eachpatientwasmovedoutoftheMRIunitandinstantlypositionedagaininsidethemagnet.ThenfourDWIse-quenceswereacquiredforasecondtimewiththesameparameters,toevaluatethereproducibilityofIVIMpa-rameters.
Theacquisition/meanacquisitiontimefordifferentDWIsequencesrangedfrom1min58sec(forFBNSA1-4)to9min48sec(forRTNSA3-6).
liverlobeintheanalysisduetothepresenceofartefacts
relatedtocardiacmotion.Table3showsthemeanROIsizesfordifferentDWIsequences,separatelyforthefirst(No.1)andsecond acquisition(No.2).
Estimatedsignal-to-noiseratio(eSNR)wascalculatedforfourbvalues(0,30,200,900s/mm2)accordingtothefollowingformula:
SImeans
eSNR=
SDnoise
whereSImeansismeansignalintensity,andSDnoiseisstandarddeviationofbackgroundnoise(measuredon
ROIplacedinthebackground).
IVIMparameterswerecalculatedusingtheIVIMmodelequationsdescribedbyLeBihanetal.[12].
S
S0
Imageanalysis
Allmeasurementswereperformedatcommercialwork-stations(Syngovia,SiemensMedicalSolution,Erlangen)bytworadiologists(JPandAAG)withsixandfouryearsofexperienceinbodyMRimaging,respectively.ForthecalculationofIVIMparameters(f,D,andD*)regionsofinterest(ROIs)weredrawnonchosenslicesofb=0images,independentlyforeachDWIsequencefortherightliverlobe.ThenROIswerecopied andpastedfromtheseimagestocorrespondingscansacquiredwithhigh-erbvalues.Areascontainingbloodvesselsandartefactswereavoided.Measurementsofsignalintensity(SI)wereperformedtwiceandaveraged.Wedidnotincludetheleft
=(1–f).exp(–b.D)+f.exp(–b.(D*+D))
S–signalintensityinthepixelwithdiffusiongradientb,
S0–signalintensityinthepixelwithoutdiffusiongra-dient,
D–truediffusionasreflectedbypuremoleculardiffusion,f –fractionalperfusionrelatedtomicrocirculation,
D*–pseudo-diffusioncoefficientrepresentingperfu-sion-relateddiffusionorincoherentmicrocirculation.Forcalculationsweappliedourownprogramwrit-
ten inGnuplotprogramversion5.0, patchlevel4,utilis-ingthemethodofthenonlinearleast-squares(NLLS)Marquardt-Levenbergalgorithm.
©PolJRadiol2018;83:e437-e445
e439
Sequences	FBNSA1-4	RTNSA1-4	FBNSA3-6	RTNSA3-6
Examination
No.1
No.2
No.1
No.2
No.1
No.2
No.1
No.2
Meanvalue(cm2)
1.35
1.59
1.28
1.31
1.32
1.29
1.39
1.58
Standarddeviation(cm2)
0.62
0.52
0.56
0.53
0.54
0.47
0.56
0.58
Sequences	FBNSA1-4	RTNSA1-4	FBNSA3-6	RTNSA3-6
Acquisitiontime
1’58”
3’48”-6’27”(mean5’04”)
4’
7’45”-16’05”(mean9’48”)
Resolutionmatrix
128×84
128×84
128×84
128×84
FoVread
380
380
380
380
FoVphase
65.6%
65.6%
65.6%
65.6%
Slicethickness
5mm
5mm
5mm
5mm
Slicegap
35%
35%
35%
35%
Slices
33
33
33
33
TR
6100ms
4200ms
6100ms
4200ms
TE
56ms
56ms
56ms
56ms
Accelerationfactor
3
3
3
3

AndrzejCieszanowski,KatarzynaPasicz,JoannaPodgórskaetal.
Table4.Intravoxelincoherentmotion-derivedparametervaluesoftheliver
NSA–numbersofsignalaverages,FB–free-breathing,RT–respiratory-triggering
SimilarvalueswereobtainedwithRTNSA3-6sequence:
f=22%,D=1.07×10–3mm2/s,83×10–3mm2/s.Cor-
respondingvaluesdifferedforFBNSA1-4acquisition(f=27%,D =0.98×10–3mm2/s,D* =56×10–3mm2/s)
andforRTNSA1-4acquisition(f=27%,D=1.01×10–3mm2/s,D*=78×10–3mm2/s).
Statisticalanalysis
AllstatisticalanalysiswasperformedusingRversion
3.3.2.Thedistribution ofextractedIVIM-parametersforeachDWIsequencewasgivenasthemedianvaluewithfirstandthirdquartile.Thecalculatedvaluesf,D,andD*werecomparedbetweenspecificgrouppairs(FBinthefirstexaminationvs.FBinthesecondexamination;RTinthefirstexaminationvs.RTinthe secondexamination).Anon-parametricWilcoxonsigned-ranktestwasused.Statisticalsignificancewasdefinedasp<0.05.NormaldistributionofdifferencebetweenFB1vs.FB2andRT1vs.RT2wastestedusingShapiro-Wilktest.Inter-examinationrepeatabilitybetweenFBandRToff,D,andD*wereas-sessedbycoefficientofvariation(CV)andBland-Altmanmeandifferenceand95%limitsofagreements(BA-LA).
Reproducibility
Coefficientofvariance(CoV)forFBNSA1-4,RTNSA1-4,FBNSA3-6,andRTNSA3-6sequencesvariedfrom12%to99%(Table5).ThecoefficientofvarianceofIVIM-de-rivedparametersoftheliverwasbetterforNSA3-6 se-quences,comparedtoNSA1-4sequences(exceptforD*calculatedfromRTNSA3-6acquisition).Thecomparisonbetweenfree-breathingandrespiratory-triggeredacquisi-tionsdemonstratedthatrespiratory-triggeringgeneratedbetterinter-scanreproducibilityforNSA1-4sequences,whereasfree-breathingacquisitionwasmorerepeatableintermsofIVIM-derivedparametersforNSA3-6sequences.
Bland-Altmanplotsofinter-examinationrepeatabilityofIVIM parametersforNSA3-6 sequenceareshowninFigure2.Bland-AltmanplotsforNSA1-4sequenceswerenotpresentedduetothefactthatdatadidnotfulfiltherequirementofthenormaldistribution.ForNSA3-6FBsequenceallthedatapointsoff,D,andD*werewithinthelimitsofagreements.ForNSA3-6RTsequenceonlyDandD*werewithinthelimitsofagreement.Thelimitsofagreementforf,D,andD*parameterswerelowerforFBsequence(Table6).
Results
Intravoxelincoherentmotionparameters
Table4showsvaluesofcalculated IVIMparametersoftherightliverlobeforappliedDWIsequences.Calculat-edvaluesoff,D,andD*werenotsignificantlydifferentbetweenFBandRTsequencesforthefirstandsecondexaminations.
Figure1showsdistributionofIVIMparametersinhealthyvolunteers,obtainedwiththeuseofallappliedDWIsequencesforthefirst(No.1)andsecond(No.2)acquisition.Thehorizontallineineachboxrepresentsthemedian.Topandbottomofboxesrepresent25%and75%percentilesofdatavalues,respectively.Individualpointsplottedonthefigurerepresentoutliers.
MedianvaluesofIVIM-derivedparametersoftheliverThemeanIVIM-derivedparameterswerecalculated
asamedianvalueoftwoconsecutiveacquisitionsforeachtypeofsequence.UsingthemostreproducibleDWIse-quence(FBNSA3-6),thefollowingnumberswereobtained:f=23%,D=1.03×10–3mm2/s,D*=72×10–3mm2/s.
Signal-to-noiseratio
Signal-to-noiseratiosoftheliverforselectedbvalueim-ages(0,30,200,900)forallappliedDWIsequencesareshowninTable7.InallacquisitionsandforallanalysedbvaluestheaverageeSNRwasabove30.Inbothtypesofacquisitions(NSA1-4vs.NSA3-6)eSNRswerehigh-er(forallbvalues)inrespiratory-triggeredsequences.
e440
©PolJRadiol2018;83:e437-e445
Parameter	Firstexamination	Secondexamination	P
FB,median	RT,median	FB,median	RT,median	FB	RT(quartiles)	(quartiles)	(quartiles)	(quartiles)
NSA1-4
f(%)
28(25-34)
28(25-31)
26(0.22-0.30)
25(0.19-0.30)
0.50
0.28
D(×10-3mm2/s)
0.90(0.76-1.06)
1.00(0.89-1.16)
0.99(0.90-1.21)
1.02(0.92-1.15)
0.05
0.42
D*(×10-3mm2/s)
44(8-120)
78(29-115)
128(23-185)
76(38-152)
0.06
0.23
NSA3-6
f(%)
24(18-28)
21(16-24)
22(19-26)
24(20-26)
0.66
0.16
D(×10-3mm2/s)
1.01(0.97-1.13)
1.03(0.99-1.14)
1.07(0.99-1.14)
1.06(0.98-1.28)
0.39
0.45
D*(×10-3mm2/s)
78(49-117)
89(49-114)
66(42-123)
67(44-111)
0.97
0.62

ReproducibilityofIVIMofliverat3.0T
A
B
100
40
80
30
60
20
40
10
20
0
0
FBNo1
FBNo2
RTNo1
RTNo2
FBNo1
FBNo2
RTNo1
RTNo2
NSA1-4:parametrf
NSA3-6:parametrf
C
D
2
2.0
1
1.5
0
1.0
–1
0.5
–2
0.0
FBNo1
FBNo2
RTNo1
RTNo2
FBNo1
FBNo2
RTNo1
RTNo2
NSA1-4:parametrD
NSA3-6:parametrD
E
F
300
400
250
300
200
150
200
100
100
50
0
0
FBNo1
FBNo2
RTNo1
RTNo2
FBNo1
FBNo2
RTNo1
RTNo2
NSA1-4:parametrD*
NSA3-6:parametrD*
Figure1.Box-plotsoff,D,andD*values
©PolJRadiol2018;83:e437-e445
e441
10–3
10–3
[%]
10–3
10–3
[%]

AndrzejCieszanowski,KatarzynaPasicz,JoannaPodgórskaetal.
Table5.Coefficientofvariation(expressedin%)ofintravoxelincoherentmotionparameters
TheeSNRswerehigherforNSA3-6sequencescompared
toNSA1-4sequences,exceptforRTacquisitionwithabvalueof200(meaneSNRwashigherforNSA1-4se-quence)andFBacquisitionwithabvalueof900(meaneSNRwasequalforbothacquisitions).
Discussion
IVIMDWmodel,introducedinthelate1980sbyLeBihan
etal.[12],isan imagingtechnique employingmultiplebvalues andbi-exponentialfittingforthesimultaneouses-timationofthe puremolecularwater diffusionand micro-circulationofbloodinthecapillaries(perfusion).InterestintheapplicationofIVIMtechniquefortheimagingof
100
80
60
40
20
0
–20
–40
–60
80
A
B
60
40
20
0
35
–20
–40
Average
FB:fparameterMean
Average
FB:DparameterMean
Difference
Upper
Lower
Difference
Upper
Lower
C1200
1000
800
600
400
200
0
D
100
80
60
40
20
0
–20
–40
–60

0–2005
35
–400
–600
–800
Average
FB:D*parameterMean
Average
RT:fparameterMean
Difference
Upper
Lower
Difference
Upper
Lower
80
E
F1200
1000
800
600
400
200
0
60
40

20

0
0–2005
–400
–600
–800
–20
–40
Average
RT:DparameterMean
Average
RT:D*parameterMean
Difference
Upper
Lower
Difference
Upper
Lower
Figure2.Bland-Altmanplotofpercentagedifferenceagainstmeanmeasurements,withmeanpercentagedifferenceandupperandlowerlimitsofagree-ments
e442
©PolJRadiol2018;83:e437-e445
Difference[%]
Difference[%]
Difference[%]
Difference[%]
Difference[%]
Difference[%]

0	100	150	200	250

0	0.2	0.4	0.6	0.8	1.0	1.2

0	0.05	0.1	0.15	0.2	0.25	0.30	0.

0	100	150	200	250

0	0.2	0.4	0.6	0.8	1.0	1.2

0	0.05	0.1	0.15	0.2	0.25	0.30	0.

Firstexamination	Secondexamination
FB	RT	FB	RT

NSA1-4
f[%]
60
50
64
69
D[×10–3mm2/s]
84
41
58
97
D*[×10–3mm2/s]
99
82
99
88

NSA3-6
f[%]
27
32
22
21
D[×10–3mm2/s]
15
12
15
17
D*[×10–3mm2/s]
70
50
80
80

ReproducibilityofIVIMofliverat3.0T
Table6.95%confidenceintervalsofpercentagedifferencebetweenthepairedmeasurementsforNSA3-6
abdominalorganshasrecentlyincreased,includingdiag-
nosisofliver,pancreatic,andrenaldiseases.Thepossibleadvantageofthistechniqueisquantitativeassessment ofbothtumour(ortissue)cellularity andvascularity,withouttheneedforintravenouscontrastmaterialadministration.However,thereareseveralimportantlimitationsprevent-ingtheapplicationofIVIManalysisinroutinepractice.Amongthemostimportantonesarethelackofstandard-isationofIVIMsequences,prolongedacquisitiontime(duetoapplicationofmultiple bvalues),relativelypoorqualityofobtainedimages,andpredilectionofsomeoftheIVIM-derivedparameters(particularlyD*)toerrors.
BecauseIVIMimagingreliesonquantitativeanal-ysisoff,D,andD*,reproducibilityofcalculatedIVIMparametersremainsacrucialissue.Forthisreason,itisnecessarytostandardiseboththeimagingprotocolandanalysisofacquireddata.Previousoptimisationstudiesmostlyconcentratedonseveralissuesincludingchoiceofbvalues,differentrespiratorycorrectionschemes,andpostprocessingmethods.Tothebestofourknowledge,ourstudyisthefirstattempttosystematicallyevaluatetheeffectofthenumberofsignalaveragesonthereproduc-ibilityofIVIM-derived parameters.Forthispurpose,wecomparedconsecutiveacquisitionsobtainedduringthesamesession,withtheshortestpossibleintervalbetweenthem.Inpreviousstudiesevaluatingthereproducibili-tyofIVIMacquisitionstheaveragedelaysbetweentwoMRexaminations wereninedays(range:0-45days)and13days(range:5-45days)[6,10].Webelievethatthemethodutilisedinourstudyenabledustodiminishpos-siblesubstantialchangesrelatedtoalteredphysiologicalperfusioninthehumanbody.
Inpublishedreports,NSAforlowbvaluesvariedfrom1to3andinsomereportsthisfigurewasnotavail-able[6,10,13,14].ThenewgenerationofMRIscannerspermitsseparateadjustmentofNSAforeachbvalue,thusenablingfurthercompensation ofsignal loss onhighbvalueimagesbyincreasingtheNSA.Weutilised thisapproachbyemployingincreasingNSA:from1to4forFBandfrom3to6forRTsequences(Table2).Ourre-sultsshowedthatutilisationofsinglesignalaverage (forlowbvalues)forIVIMimagingisnotsufficienttoobtainreliableresults.IncreasingNSAallowsustoimprovetheglobalinter-scanreproducibilityofIVIM-derivedparam-etersforbothfree-breathingandrespiratory-triggeredse-quences.ThisimprovementwasmostnoticeableforFBtechniqueforparameterD*–decreaseofmeancoefficientofvariancefrom99%forFBNSA1-4to70%(inthefirstexamination)and80%(inthesecondexamination)forFBNSA3-6,whichisparticularlypronetomiscalculation.Interestingly,intermsofinter-scanreproducibilitythe
FBNSA3-6sequenceperformedbetterthanRTNSA3-6sequence.Thelimitsofagreementsbetweenthepairedmeasurementsforf,D,andD*parameterswerelowerforFBsequence(Table6).ThisdifferencewasagainmostprominentforD*.Thisfindingisinoppositiontoaprevi-
Widerintervalsindicate poorermeasurementreproducibility.
NSA–numbersofsignalaverages,FB–free-breathing,RT–respiratory-triggering
Table7.Estimatedsignal-to-noiseratiooftheliverforselectedbvalueimages(shownasmeanvalue±standarddeviation)
NSA–numbersofsignalaverages,FB–free-breathing,RT–respiratory-triggering
ouspublicationofDyvorneetal.,whoconcludedthatRT
sequenceyieldedthebestimagequality,reproducibility,andabilitytodiscriminatebetweenfibroticandhealthyliver,whereasstrongerimageblurringwasnotedonFBacquisitions[10].TheyreportednosignificantdifferenceinD*betweenhealthyvolunteersandHCVpatients,at-tributingittohighuncertaintyandpoorreproducibilityofD*.Inourgroupofvolunteers,thereproducibilityofD*wasbetterusingFBNSA3-6acquisitionthanusingRTNSA3-6acquisition.Wedidnotperform visual, qualita-tiveevaluationoftheofIVIMimagesbecausewebelievethatquantitativeassessmentofIVIM-derivedparame-tersisthemostimportant.ThepossiblereasonforbetterperformanceofFBNSA3-6overRTNSA3-6acquisitionmaybetheirregularpatternofrespirationencounteredinsomevolunteers.Itcouldcause,alongwithpossiblyimperfectsystemofnavigatorechogating,suboptimalac-quisitionofrespiratory-triggeredimages.Moreover,themeanacquisitiontimeofRTNSA3-6sequence(9’48”;range:7’45”-16’05”)wasoverdoublethatoftheequiv-alentfree-breathingsequence(4’).Wecannotexcludethatitmayfurtherincreasepatients’discomfortduringexaminationanddiminishtheircooperation.Inaddition,prolongedacquisitiontimeofIVIMsequenceleadstoex-tensionofoverallimagingtimeanddisturbstheworkflowindailyclinicalpractice.
InpublishedstudiesregardingIVIMMR imagingoftheliver,differentrespiratoryschemeswereemployed,in-cludingfree-breathing[5,6,10,13,14],breath-holding[9],andrespiratory-triggering[6,10,15].Thesestudieswereperformedonboth1.5T[1,6,7]and3.0Tsystems[3,6,15].Dyvorneetal.,whocomparedRTandFBIVIMDWIse-quencesona1.5Tsystem,concludedthatRTsequence
©PolJRadiol2018;83:e437-e445
e443
b-value	FBNSA1-4	RTNSA1-4	FBNSA3-6	RTNSA3-6
0
65±44
94±76
78±59
99±77
30
55±44
63±48
62±40
74±53
200
43±34
66±59
51±37
61±43
900
35±41
37±35
35±36
40±34

NSA3-6
FB(%)	RT(%)
f[%]
–39;41
–49;83
D[×10–3mm2/s]
–22;25
–27;36
D*[×10–3mm2/s]
–137;162
–762;1039

AndrzejCieszanowski,KatarzynaPasicz,JoannaPodgórskaetal.
withbipolardiffusion-encodingschemesdemonstrated
thehighestimagequalityandreproducibility[10].Theseresultsarecontrarytoourfindings,althoughthisdiffer-encemaybeattributed todifferentfieldstrengthsofMRscannersusedinthesestudies.Tothe bestofourknowl-edge,ourpublicationisthefirstaddressingreproduci-bilityofIVIMDWIsequenceswithdifferentrespiratoryschemesforliverimagingona3.0TMRsystem.
3.0TDWIoftheliverappearstobeaninterestingop-tion,duetoimprovedSNRcomparedto1.5Tscanners.However,itremainschallengingbecauseofprominentmagneticsusceptibilityartefactsanddistortionrelatedtoeddycurrents[16],andmostresearchersutilise1.5Tplat-forms.IVIMimagingoftheliverisevenmoredemandingbecauseitisamovingorgan,pronetomotion-inducedmiscalculations.RegardingADC values, Rosenkrantzetal.[17]reported similar resultsforboth1.5Tand3.0T
systems,whereasDaleetal.[18]foundsignificantdiffer-
encesinADCvaluesforthesetwofieldstrengths.Theoret-ically,IVIM-derivedparametersmayprovidemoreaccu-ratedataontruemolecularmovementofwater(diffusion),alongwithmicrocirculationofbloodinthecapillaries(perfusion).However,theknowledgeoftheaveragevaluesofIVIM-derivedparametersofnormalliverforboth1.5Tand3.0Tunits iscriticalforthe diagnosis andassessmentofseverityofhepaticdisease.Otherwise,itcannotbeusedasaneligible,quantitativebiomarkerfortheestimationofresponsetotherapyinprospectivedrug trials.
Inourstudyof20healthyvolunteersthemedianDvalue(1.03±0.18×10–3mm2/s)calculatedfromopti-malsequence(FBNSA3-6)wassimilartovaluesobtainedbyotherresearchers[6,13,15]utilising3.0Tsystems.Thedifferenceinmedianfobtainedinourstudy(23%±7%)andinpreviouspublications(26.9%±8.8%;19.0%±5.5%
–on3.0Tand30.8%±4.95%–on1.5T)wasmoderate,whereasthediscrepancyinaverageD*betweendifferentstudies wasmoreprominent(72±61×10–3mm2/s inourstudyand110.6±79,57.2±40.2–on3.0Tand59.67±
12.34–on1.5T,inpreviousreports)[6,13,15].Thisdata,alongwithconsiderableinter-scanreproducibilityforD*,confirmsthatthepseudo-diffusioncoefficientisburdenedwithalargererror.Consequently,preciseassessmentofthisparameterrequiresfurtherresearchintermsofthestandardisationandoptimisationofimageacquisitionaswellaspostprocessing.Moreover,alackofstandardisationofIVIMtechniqueandsignificantvarianceincalculatedparametersamongstudiesthusfarprecludesapplicationofthismethodforroutineMRliverimagingforappli-cationasaneligiblebiomarker.Inaddition,IVIM-de-rivedparameters,similarlytoADCvalues,arepossiblybothvendorandfieldstrength dependent(1.5Tvs.3.0T)[19,20].AmongmanyfactorsthatmayinfluenceDWIcalculationsareecho time, gradientamplitude,andsepa-ration ofdiffusiongradient [20].Accuratecalculationof
IVIMparametersfroma3.0TMR scannerrequireshigh
SNRforallacquiredbvalues.Ourstudysuggeststhatapplication ofminimum NSA= 3forlowb values(withprogressiveincreaseupto6forthehighestbvalue)isre-quired togetreproducibleIVIMcalculations.
DWIsequencesarenowadaysincorporatedintorou-tineMRIprotocolsoftheliverandutilisedforthediag-nosisofbothfocalanddiffusehepaticdisease.Quantita-tiveassessmentofapparentdiffusioncoefficient(ADC)facilitateddiscriminationofmalignantfrombenignfo-calliverlesions(FLLs)andbysomeresearchersisadvo-catedforevaluationofdiffuseliverdisease.Supposedly,IVIM-derived parameters,such asperfusionfraction(f),pseudodiffusioncoefficient(D*),andtruediffusionco-efficient(D)enablegreaterinsightintothephysiologyoflivertissuecomparedtoADCalone.Thishypothesiswasconfirmedbyseveralreports[1-4,10,13,15].Yoonetal.,whoevaluatedFLLsona3.0TsystemusingIVIM-derivedparameters,concludedthatDprovidedbetterdiagnos-ticperformancethanADCindiscriminationofbenignandmalignantFLLs,whereasD*andfweresignificantlyhigherinhypervascularlesions[13].ThedecreaseofD*wasreportedmostfrequentlyinliversteatosis[5],fibrosis,andcirrhosis[2,4,13].ThedeclineofD[2,5]andf[2,10]wasalsoreportedasavaluableindicatorofdiffuseliverdiseaseinhumanstudies.
Ourstudyhassomelimitations.First,thenumberofstudiedsubjectswaslimitedto20andthemeanageofthisgroupwasrelativelylow.Second,weappliedonly10bvaluesforIVIMacquisitions.TheutilisationofmorebvaluescouldresultinmoreaccuratecalculationofIVIM-derivedparameters butattheexpenseofconsider-ableprolongationofimagingtime.Ourimagingprotocolwasatrade-offbetweenaccuracyandreasonablescanningtime.Similarlytosomeotherresearchers,weassumedthatthenumberoflessthaninitiallysuggested16bvaluesenablesacquisitionofreliabledata[4,13].Applicationofmorebvalueswouldprolongexaminationtimeandcouldbedifficulttoacceptinclinicalpractice.Third,wedidnotincludetheleftliverlobeinourmeasurementsandanaly-sisduetosignificantcardiacmotionartefactsinthisarea.Inconclusion,increasingthenumberofsignalaver-agesforIVIMacquisitionsallowsadecreasetherepro-ducibilityerrorofIVIM-derivedparameterson3.0Tsys-tems. Thesequenceacquiredduringfree-breathingwithNSArisingfrom3(forlowbvalues)to6(forthehighestbvalue)was optimalinterms ofbothreproducibilityand
acquisitiontime.
Conflictofinterest
Theauthorsdeclarethattheyhavenoconflictofinterest.
e444
©PolJRadiol2018;83:e437-e445

ReproducibilityofIVIMofliverat3.0T
References
1.
DyvorneH,JajamovichG,KakiteS,etal.Intravoxelincoherentmo-
tiondiffusionimagingofthe liver:Optimalb-valuesubsamplingandimpactonparameterprecision andreproducibility.EurJRadiol2014;83:2109-2113.
PatelJ,SigmundEE,RusinekH,etal.DiagnosisofcirrhosiswithintravoxelincoherentmotiondiffusionMRIanddynamiccon-trast-enhancedMRIaloneandincombination:preliminaryexpe-rience.JMagnResonImaging2010;31:589-600.
YoonJH,LeeJM,BaekJH,etal.Evaluationofhepaticfibrosisus-ingintravoxelincoherentmotionindiffusion-weightedliverMRI.JComputAssistTomogr2014;38:110-116.
LucianiA,VignaudA,CavetM,etal.Livercirrhosis:intravoxelincoherentmotionMRimaging–pilotstudy.Radiology2008;249:891-899.
GuiuB,PetitJ-M,CapitanV,etal.Intravoxel incoherentmotiondif-fusion-weightedimaginginnonalcoholicfattyliverdisease:A3.0-TMRstudy.Radiology2012;265:96-103.
CuiY,DyvorneH,BesaC,etal.IVIMdiffusion-weightedimagingoftheliverat3.0T:Comparisonwith1.5T.EurJRadiolOpen2015;2:123-128.
LemkeA,StieltjesB,SchadLR,etal.Towardanoptimaldistributionofbvaluesforintravoxelincoherentmotion imaging.MagnResonImaging2011;29:766-776.
YamadaI,AungW,HimenoY,etal.Diffusioncoefficientsinab-dominalorgansandhepaticlesions:evaluationwithintravoxelin-coherentmotionecho-planarMRimaging.Radiology1999;210:617-623.
LemkeA,LaunFB,KlaußM,etal.Differentiationofpancreascarci-noma fromhealthypancreatictissueusingmultipleb-values:Com-parisonofapparentdiffusioncoefficientandintravoxelincoherentmotionderivedparameters.InvestRadiol2009;44:769-775.
DyvorneHA,GaleaN,NeversT,etal.Diffusion-weightedimagingoftheliverwithmultiplebvalues:effectofdiffusiongradientpolarityandbreathingacquisitiononimagequalityandintravoxelincoherent
motionparameters–apilotstudy.Radiology2013;266:920-929.
11.
LeporqB,Saint-Jalmes H, RabraitC, et al. Optimizationofintra-
voxelincoherentmotionimagingat3.0Teslaforfastliverexamina-tion.JMagnResonImaging2015;41:1209-1217.
LeBihanD,BretonE,LallemandD,etal.SeparationofdiffusionandperfusioninintravoxelincoherentmotionMRimaging.Radiology1988;168:497-505.
YoonJH,LeeJM,YuMH,etal.Evaluationofhepaticfocallesionsusingdiffusion-weightedMRimaging:Comparisonofapparentdiffusioncoefficientandintravoxelincoherentmotion-derivedpa-rameters.JMagnResonImaging2014;39:276-285.
YoonJH,LeeJM,BaekJH,etal.Evaluationofhepaticfibrosisus-ingintravoxelincoherentmotionindiffusion-weightedliverMRI.JMagnResonImaging2014;38:110-116.
GuiuB,PetitJ-M,CapitanV,etal.Intravoxel incoherentmotiondif-fusion-weightedimaginginnonalcoholicfattyliverdisease:A3.0-TMRstudy.Radiology2012;265:96-103.
SpringerF,MartirosianP,BossA,etal.Currentproblemsandfutureopportunitiesofabdominalmagneticresonanceimagingathigherfieldstrengths.TopMagnResonImaging2010;21:141-148.
RosenkrantzAB,OeiM,BabbJS,etal.Diffusion-weightedimagingofthe abdomenat3.0Tesla:Imagequalityandapparentdiffusioncoefficientreproducibilitycomparedwith1.5Tesla.JMagnResonImaging2011;33:128-135.
DaleBM,BraithwaiteAC,BollDT,etal.Fieldstrengthanddiffusionencodingtechniqueaffecttheapparentdiffusioncoefficientmeas-urementsindiffusion-weightedimagingoftheabdomen.InvestRadiol2010;45:104-108.
BarbieriS,DonatiOF,FroehlichJM,etal.Comparisonofintra-voxelincoherentmotionparametersacrossMRimagersandfieldstrengths:evaluationinupperabdominalorgans.Radiology2016;279:784-794.
ZhuJ,ZhangJ,GaoJ,etal.Apparentdiffusioncoefficientnormali-zation ofnormalliver:Willitimprovethereproducibilityofdiffu-sion-weightedimagingatdifferentMRscannersasanewbiomarker?
Medicine(Baltimore)2017;96:e5910.
12.
2.
13.
3.
14.
4.
15.
5.
16.
6.
17.
7.
18.
8.
19.
9.
10.
20.
©PolJRadiol2018;83:e437-e445
e445