©PolJRadiol2018;83:e421-e425
DOI:https://doi.org/10.5114/pjr.2018.79588
Received:8.04.2018
Accepted:30.05.2018
Published:10.09.2018
http://www.polradiol.com
Originalpaper
Invitroanalysisofurinarystonecompositionindual-energy
computedtomography
MonikaStępień1A,B,C,D,E,F,G,RobertChrzan1A,B,C,D,E,WojciechGawlas2A,B,F,G
1DepartmentofRadiology,JagiellonianUniversityinKraków,CollegiumMedicum,Poland
2DepartmentandClinicofUrology,JagiellonianUniversityinKraków,CollegiumMedicum,Poland
Abstract
Purpose:Dualenergycomputedtomography(DECT)isanewmethodofcomputedtomography(CT)imaging,allowingtheassessmentofnotonlytheobject’smorphology, butalsoitscomposition.TheaimofthestudywastoevaluatethepotentialofinvitroDECTevaluationofurinarystones’chemicalcomposition.
Materialandmethods:SixsamplesofsurgicallyremovedrenalstoneswerescannedusingDECTandanalyzedbyscannervendorsoftware.Uricacidstonesweremarkedredandcalciumstoneswhitebythesoftware.Therealcom-positionofthestoneswasfinallyverifiedusingphysicochemicallaboratoryanalysis.
Results:In5outof6samples,thecompositionofstonesinDECT(3samplesidentifiedasuricacidand2samplesascalcium)wasconsistentwiththephysicochemicalanalysis(3samplesidentifiedasuricacid,1ascalciumphosphate,1ascalciumoxalate).InDECTitwasnotpossibletodeterminemorepreciselythetypeofcalciumcompounds(calciumphosphatevs.calciumoxalate)asestablishedinthephysicochemicalanalysis.
Inonestoneidentifiedinphysicochemicalanalysisasuricacid,DECTdetectedacompositelayered structurecon-tainingbothuricacidandcalciumcompounds.
Conclusions:DECTallowsuricacidtobedistinguishedfromcalciumurinarytractstones,whichiscrucialinthechoiceofappropriatetherapy.Usingtheavailablehardwareandsoftware,itwasnotpossibletomoreaccuratelydistinguishtypesofcalcifiedstones.EvaluationofthestonetypeinDECTmaybelimitedinthecaseofmixedchemicalcomposition.
Keywords:urinarystone,dualenergycomputedtomography.
ureterorbladderwheretheycangroworbeexcretedin
theurine.
Sometimesthedepositsreachaconsiderablesize,andfilltheentirerenalpelvisandcalyces(staghorncalculi),leadingtokidneydamage.
Themostcommondepositsarecomposedofcalci-umoxalate,lessfrequentlycalciumphosphate,includingcalciumphosphatedihydrate(brushite),uricacid,mag-nesiumammoniumphosphate(struvite)andcystine.Inspecialcases,othertypesofurinarystonesarise:xanthine
Introduction
Urolithiasisisacommondiseaseassociatedwiththepres-
enceofinsolubledepositsintheurinarytract,whichariseasaresultoftheprecipitationofchemicalsintheurine,whentheirconcentrationexceedsthesolubilitythresh-old.Kidneystonesarefoundin10-12%ofmenandabout5-6%ofwomenthroughoutlife[1].Depositscanariseinvarioussectionsoftheurinarytract,mostcommon-lyinthecalycesandrenalpelvis,andthenmovetothe
Correspondenceaddress:
MonikaStępień,DepartmentofRadiology,JagiellonianUniversityinKraków,CollegiumMedicum,19KopernikaSt.,31-501Kraków,Poland,e-mail:mdmonikas@gmail.com
Authors’contribution:
AStudydesign∙BDatacollection∙CStatisticalanalysis∙DDatainterpretation∙EManuscriptpreparation∙FLiteraturesearch∙GFundscollection
©PolJRadiol2018;83:e421-e425
e421

MonikaStępień,RobertChrzan,WojciechGawlas
deposits(ininheriteddeficiencyofxanthineoxidase),acid
ammoniumurate(inchronicdiarrheaandhypokalemia),oxypurinolandxanthine(inchronictreatmentwithallop-urinolortriamterene)[2].
Treatmentofurolithiasisdependsonthechemicalcomposition,locationandsizeofthedepositsandmaybeconservativeorsurgical.Knowledgeofthechemicalcompositionmaybecrucialinchoosingtherighttherapyandinforecastingthetreatmentresults.
Chemolysisofdepositsispossibleonlyinsometypesofnephrolithiasis.
Inthe caseofstruvite,brushite,cystineanduricaciddeposits,percutaneousrinsingwithsolubilizingagents(hemiacridrin,Suby’sG,THAM,N-acetylcysteine)maybeused.
Foruricaciddeposits,oralmedicationsmayalsobeef-fectivetoalkalinizetheurinetopHof6.5-7.2usingalkalinecitratesorsodiumbicarbonate;invasivetreatmentisim-plementedonlywhenconservativetreatmentisineffective.CystinedepositsarebestdissolvedatpH=7.0-7.5us-ingmainlyD-penicillamine,alpha-MPG,captopriland
tiopronin[3].
Themostcommonmethodofurolithiasistreatmentislithotripsywithshockwavesgeneratedextracorporeal-ly(extracorporealshockwavelithotripsy–ESWL).Thismethoduseselectromagnetic,electrohydraulicorpiezo-electriclithotripters,locatingthestonewithfluoroscopyand/orultrasound.TheoptimallocationforESWListhepelvisandcalyces(especiallyupperandmiddle)andtheproximalpartoftheureter(abovethesacroiliacjointline);inthedistalpartoftheureter,itisbettertoperformure-teroscopy.UsuallyESWLisusedinkidneydepositsupto20mminsizeandureterdepositsupto10mm[3].Calci-umoxalateanduricacidstonesdisintegraterelativelyeasi-ly.Calciumphosphates,especiallythosecontainingbrush-ite,crumbleworseandcystinedepositsdefinitelybadly[4].Othermethodsoftreatmentinclude:ureterorenosco-
py(URS)involvingremovalofureterstonesundervisualcontrolusingaureterorenoscopeandpercutaneousneph-rolithotomy(PCNL)basedonremovalofthedepositfromthekidneyor upperureter,wholeor crushed,usingan en-doscopeinserteddirectlyintothepelvico-calycealsystem.Determinationofthechemicalcompositionofureter-
alconcrementsinvivocouldsignificantlyinfluencethe
selectionofoptimaltherapeutictreatment.
Untilrecently,wedidnothaveadiagnosticmethodtoobtainsuchdatabeforespontaneousexpellingorsurgicalremovalofconcrement.
Thissituationchangedonlyaftertheintroductionofdualenergycomputedtomography(DECT).
TheaimofDECTistoobtaintwosetsofdata(im-ages)duringtheacquisitionofthesameanatomicalareaatdifferentvaluesofX-rayenergy(usually80keVand140keV)[5].
ThetheoreticalprinciplesofDECTweredevelopedinthe1970s,butthefirstcomputedtomographyscanners
withthepracticalimplementationofthismethod,using
varioustechnicalsolutions,appearedonlyafewyearsago.DECTallowsdepositscontaininguricacidtobedis-tinguishedfromdepositswhichdonotcontainuricacid,intestsperformedbothinvivoandinvitro[6].Duetothedifferentchemicalstructure,uricaciddepositsshowhigherradiodensity Hounsfieldunitvaluesusinghigh-en-ergyradiation(athigherkilovoltage),whileothertypesofdeposits(includingcalciumdeposits)showhigherra-diodensityHounsfieldunitvaluesusinglow-energyra-diation(lowkilovoltage)[7].DataobtainedinDECTarepostprocessedusingdedicatedsoftware,inwhichvoxelscontaininguricacidarecodedinadifferentcolorcom-
paredtovoxelsthatdonotcontainuricacid.
TheassessmentoftheusefulnessofDECTindetermin-ingthechemicalcompositionofurinarystonesisanin-creasinglycommonsubjectofresearch[7-11].Mostofthemareinvivostudies,butsomefocusoninvitroevaluation.
TheaimofourstudywastoevaluatethepossibilitiesofinvitroDECTassessmentofurinarystones’chemicalcomposition.
Materialandmethods
Six samplesofsurgicallyremovedurinarystoneswere
analyzed.Theindividualsampleswereplacedinclosedplasticcontainers.Therewerebothsinglelargeandnu-meroussmallconcrementswithinthesamples.
ThesampleswerescannedintheDECTtechniqueus-ingaToshibaAquilionPrime80computedtomographyscanner,withthefollowingparameters:voltageandcur-rent135kVand100 mAor80kVand570 mA,rotationtime0.5s,configurationofdetectors80×0.5mm,pitch0.171,reconstructionintervalandslicethickness1mm,FOV(fieldofview)65mm.
TheCTDIvol(CTdoseindex)fortheaboveprotocolwas33.2mGy.
AfterprocessingthedatafromscanningintheDECTtechnique,theaxialimagesandthesecondary2Dand3Dreconstructionswereobtained.
TheanalysisoftheaboveimageswasperformedusingToshibaDEStoneAnalysissoftware.
Theelementsofsamplescontaininguricacidweremarkedbytheprograminred,whilethosecontainingcal-ciumcompounds weremarkedinwhite(Figures1and2).Thecompositionofstonesamplesdeterminedintheabovewaywascomparedwiththeircompositionevaluated
inthelaboratorytestsduringphysico-chemicalanalysis.
Results
AfterDECTexaminationofsixsamplesanddataprocess-
ing,axialimagesandsecondary2Dand3Dreconstruc-tionswereobtained,withelementscontaininguricacidmarkedinred,andthosecontainingcalciumcompoundsinwhite(Figures1-8).
e422
©PolJRadiol2018;83:e421-e425

Invitroanalysisofurinarystonecompositionindual-energycomputedtomography
Figure1.DECT–6samplesofsurgicallyremovedurinarystones–2Dre-constructioninthesagittalplane;respectivelyfromabove:uricacid,calciumoxalate,uricacid,uricacid,calciumphosphate,layeredstonecontainingbothuricacidandcalciumcompounds;uricacidmarkedinred,calciumcompoundsinwhite
Figure2.DECT–6samplesofsurgicallyremovedurinarystones–3Dre-construction;respectivelyfromabove:uricacid,calciumoxalate,uricacid,uricacid,calciumphosphate,layeredstonecontainingbothuricacidandcalciumcompounds;uricacidmarkedinred,calciumcompoundsinwhite
Figure3.DECT–asampleofsurgicallyremovedurinarystone–axialcrosssection–uricacid
Figure4.DECT–asampleofsurgicallyremovedurinarystone–axialcrosssection–calciumoxalate
Infiveoutofsix samples(Figures3-7)the assess-
mentofthestones’compositioninDECT(threesam-plesdeterminedasuricacidandtwosamplesascalciumcompounds)agreedwiththeresultofphysico-chemicalanalysis(threesamplesdefinedasuricacid,oneascalci-umphosphate,oneascalciumoxalate).InDECTtech-nique,itwasnotpossibletodeterminemorepreciselythetypeofcalciumcompounds(calciumphosphatevs.calciumoxalate),asverifiedinthephysico-chemicalstudy.
Inoneofthesixsamples(Figure8),thestonedeter-minedinthephysico-chemicalanalysisasauricacidde-
posit,inDECTturnedouttohaveacomplexlayeredstruc-
ture,containingbothuricacidandcalciumcompounds.
Discussion
Commonlyavailablemulti-rowCTallowstheassessment
ofonlythesizeandlocationofstones.Duetotheover-lappingrangesofCTattenuationvaluesamongdifferenttypesofcalculi,itisimpossibletodeterminetheirchem-icalcompositioninatypicalCTscan[12].
TheintroductionoftheDECTmethodchangedit–dualenergycomputedtomographyallowsonetodiffer-
©PolJRadiol2018;83:e421-e425
e423

MonikaStępień,RobertChrzan,WojciechGawlas
Figure5.DECT–asampleofsurgicallyremovedurinarystone–axialcrosssection–uricacid
Figure6.DECT–asampleofsurgicallyremovedurinarystone–axialcrosssection–uricacid
Figure8.DECT–asampleofsurgicallyremovedurinarystone–axialcrosssection–layeredconcrementcontainingbothuricacidandcalciumcompounds
Figure7.DECT–asampleofsurgicallyremovedurinarystone–axialcrosssection–calciumphosphate
entiatewithhighaccuracyuricacidfromnon-uricacid
stones,bothinvivo[7,8]andinvitro[9-11].
OurinvitroDECTstudyofthestones’composition,withanalysisperformedbythescannermanufacturer’ssoftware, confirmedthepossibility ofdistinguishinguricacidfromcalcium-containingconcrements;however,itwasnotpossibletodeterminemorepreciselythetypeofcalciumcompounds.
Similarly,Graseretal.[9]foundthatinDECTitispossi-
bletoreliablydistinguishbetweenuricacidandothertypesofstones,buttherangeofattenuationvaluesforstruvitestoneslargelyoverlappedwiththe rangeformixedstonesandpartiallywiththerangeforcalcifiedconcrements.
Accordingtodatafromtheliterature,moregroupsofdepositscanbedistinguishedbymeansofmoreadvancedsoftwareandmodificationofthebeamfiltration.
Hidasetal.[8],inastudyusingaCTscannerwith
asinglex-raytubeandadualarrayofhighandlowenergyphotondetectors,assessedtheurinarystonesinpatientsinvivopreoperativelyandcomparedtheresultswithpost-operativeX-raydiffractionanalysisofthestones.
BeforeinvivoCTscanning,theresearcherscreatedacatalogofstoneswith knownchemicalcomposition,inwhichtheyassigneddepositstosevengroups(includinguricacid,cystine,struvite,apatite,weddellite,whew-elliteandbrushite)andthentheyevaluatedtheratiooflow-energyattenuationtohigh-energyattenuationforeveryabove-mentionedstone.Usingdataobtainedfromanin vitroCTstudy,researcherswereabletodistinguishbetweenuricacid,cystineandcalciumstones,inaCTstudyperformedinvivo.Duetosomesimilarityofchem-icalstructure,technicallimitationsoftheCTscanner
e424
©PolJRadiol2018;83:e421-e425

Invitroanalysisofurinarystonecompositionindual-energycomputedtomography
anddifferencesinradiationbeamabsorptioninpatients
withdifferentbodystructures,itwasnotpossibletodis-tinguishbetweenstruvitestonesandsubtypes ofcalciumdeposits.
Fungetal.[10],usingadual-sourceCTscannerwith
atinfilterattachedtotheX-raytubeemittinghigh-energyradiation,wereabletodifferentiateinvitrostonescontain-ing uric acid,calciumoxalatesandcalciumphosphates.Applicationofatinfilterresultedinbetterseparationoftheattenuationvaluesofoxalatesandcalciumphosphates,enablingtheirdifferentiation.
Similarly,Quetal.[11]evaluatedthecompositionofstonesinvitro,usingadualenergy,dual-sourceCTscannerwithanadditionaltinfilter.Inthisstudyitwas
possibletodistinguishfivegroupsofconcrements:(group1:uricacid,uricaciddihydrate,ammoniumacidurate;group2:cystine;group3:struvite;group4:calciumox-alatemonohydrate,calciumoxalatedihydrate,brushite;group5:hydroxyapatiteandcarbonateapatite).However,theauthorsnotedpartialoverlapofthefeaturesofsomeconcrementtypes.Moreover,theadditionaltinfilterisanoptionavailableonlyfromsomemanufacturers.
InourdualenergyCTstudy,inmostcasestheassess-mentofthechemicalcompositionofstonesagreedwiththeresultofthephysico-chemicalanalysis(uricacidvs.calciumdeposits);inonecase,however,wefounddis-cordance.Thereasonwasamixed-layeredstructureofthestone,built ofboth uricacidandcalciumcompounds.Similarly,Manglavitietal.[7],investigatingthecon-crementswithadual-sourceCTapparatus,infourcasesfoundadiscrepancybetweentheCTassessmentofthechemicalcompositionofconcrementsandtheresultofthecrystallographicstudy.Alltheerroneouslyevaluatedconcrementshadamixedchemicalstructure(theycon-
taineduricacidandhydroxyapatite)andhadadiameter
below1cm.Probablythemixedchemicalstructureoftheconcrement,combinedwithitssmallsize,contributetotheaccuracyreductionintheDECTtechniqueanalysis.
OurstudyconcernedtheuseofDECTtechniquetoassessurinarystonecompositioninvitro,whereasinthecaseofsuchinvivoassessment,theradiationdoseusedduringthestudybecomesanimportantfactor.Thecom-putedtomographydoseindex(CTDIvol)forourprotocolwas33.2mGy,whichisasignificantlyhighervaluethan
thatusedinstandardabdominalandpelvicexaminations(about7-18mGy)andmuchhigherthanthevalueinthededicatedlow-doseprotocolfortheassessmentofurinarytractstones(approx.4mGy).Fortunately,inthecaseofinvivoevaluation,theDECT protocolwiththehigherdoseisusedonlyintherangeincludingthelocationofthecon-crements.Moreover,Chaytoretal.[13]foundthatreliableinvivoDECTassessmentofstonecompositionispossibleusingalowerCTDIvolvalueof11.7mGy.
Conclusions
Dual-energyCTtechniqueenablesdifferentiationof
uricacidfromcalcium-containingurinarystones.
1. Usingtheavailablehardwareandsoftware,moredetaileddistinctionofcalcifiedstonesubtypeswasnotpossible.
2. Evaluationofurinarystones’chemicalcompositionindual-energycomputedtomographymaybelimitedinthe caseofconcrementswith mixed chemicalcom-position.
Disclosure
Theauthorsdeclarenoconflictofinterests.
References
1.
EdvardssonV,IndridasonO,HaraldssonG,etal.Temporaltrendsin
theincidenceofkidneystonedisease.KidneyInt2013;83:146-152.SułowiczW,StompórT,DrabczykR.Kamicanerkowa.In:GajewskiP(ed.).InternaSzczeklika2016.MedycynaPraktyczna,Liszki2016;1567.ŚwiniarskiP.Kamicanerkowa–rodzaje,objawy,leczenie.PrzeglądUrologiczny2014;2:13.
KoźmińskaE.Kamicamoczowa.In:BorkowskiA(ed.).Urologia.WydawnictwoLekarskiePZWL,Warszawa2015;207-210.
JohnsonTRC.Dual-energyCT:generalprinciples.AJRAmJRoent-genol2012;199:S3-S8.
DaleJ,GuptaRajanT,MarinD,etal.Imagingadvancesinurolithi-asis.JEndourol2017;31:623-629.
ManglavitiG,TresoldiS,GuerrerCS,etal.Invivoevaluationofthechemicalcompositionofurinarystonesusingdual-energyCT.AJRAmJRoentgenol2011;197:W76-83.
HidasG,EliahouR,DuvdevaniM,etal.Determinationofrenalstonecompositionwith dual-energyCT:invivo analysisand com-
parisonwithx-raydiffraction.Radiology2010;257:394-401.
9.
GraserA,JohnsonTR,BaderM,etal.DualenergyCTcharacteriza-
tionofurinarycalculi:initialinvitroandclinicalexperience.InvestRadiol2008;43:112-119.
FungGS,KawamotoS,MatlagaBR,etal.Differentiationofkidneystonesusingdual-energyCTwithandwithoutatinfilter.AJRAmJRoentgenol2012;198:1380-1386.
QuM,Ramirez-GiraldoJC,LengS,etal.Dual-energydual-sourceCTwithadditionalspectralfiltrationcanimprovethedifferentiationofnon-uricacidrenalstones:anexvivophantomstudy.AJR AmJRoentgenol2011;196:1279-1287.
PlattJF,KazaRK.Renalapplicationsofdual-energyCT.AbdomRadiol(NY)2016;41:1122-1132.
ChaytorRJ,RajbabuK,JonesPA,etal.Determiningthecomposi-tionofurinarytractcalculiusingstone-targeteddual-energyCT:evaluationofalow-dosescanningprotocolinaclinicalenviron-
ment.BrJRadiol2016;89:20160408.
2.
10.
3.
4.
11.
5.
6.
12.
13.
7.
8.
©PolJRadiol2018;83:e421-e425
e425