Cardiology Journal
n
u
l
2007, Vol. 14, No. 2, pp. 155-159
M 2 Ï Ï
Copyright © 2007 Via Medica
O R IG IN A L A R TIC LE
v i a m e d ic a
is s n 1507-4145
Retrospective analysis of reasons for failure
of DDD pacemaker implantation in patients
operated on between 1993 and 2005
Jacek Lelakowski, Jacek Majewski, Barbara Małecka,
Jacek Bednarek, Paweł Stypuła and Marcin Szeglowski
Collegium Medicum
Department of Electrocardiology, Institute of Cardiology,
Jagiellonian University, John Paul II Hospital, Cracow, Poland
Abstract
Background: During implantation of a DDD pacemaker the following difficulties may be
encountered: venous anomalies (the absence of vessels of adequate calibre or difficulty in
subclavian vein puncture), arrhythmias during implantation (episodes of atrial flutter/fibril­
lation while the atrial leads are being positioned), lack of mechanical stability of the electrode
in the heart chamber and inability to achieve an acceptable pacing and sensing threshold
during implantation. The purpose of the study was to analyse retrospectively the reasons for
DDD pacemaker failure in patients operated on between 1993 and 2005.
Methods: We reviewed retrospectively all implantation data from 1988 to 2005 to identify
patients with primary failure of DDD pacemaker implantation. Further analysis included
patients who had received a DDD pacemaker between 1993 and 2005, when this type of
pacemaker made up between 9 and 40% of all pacemaker implantations. We implanted 7469
pacemakers, including 1958 (26.2%) dual-chamber pacemakers, in 783 patients with atriov­
entricular block (AVB), 392 with sick sinus syndrome (SSS), 450 with AVB ± SSS and
333 with tachy-brady syndrome (TBS). The mean age of the patients was 65.5±17.3 years.
DDD pacing was unsuccessful in 108 (1.4%) patients, including 32 with AVB, 22 with SSS,
16 with SSS ± AVB and 38 with TBS. The mean age of these patients was 78.5 ± 19.4 years.
Results: The reasons for failed implantation were venous anomalies in 12%, an arrhythmia
episode in 27.8%, a high pacing threshold in the atrium in 17.6%, low atrial potential ampli­
tude in 25.9% and lack
of mechanical stability of the electrode in 16.7% of patients.
The difficulties were encountered in elderly patients (p < 0.01), most frequently in patients
with SSS and TBS (71). Between 2004 and 2005 venous anomalies and a high pacing
threshold were the main causes of failure.
Currently the main difficulties encountered during pacemaker implantation
Conclusions:
are venous anomalies and a high pacing threshold. Arrhythmia episodes, low atrial potential
Address for correspondence: Jacek Lelakowski, MD, PhD
Department of Electrocardiology, Institute of Cardiology
CMUJ, John Paul II Hospital
Prądnicka 80, 3 1 -2 0 2 Cracow, Poland
T el: + 4 8 12 614 22 77, fax: + 4 8 12 633 23 99
e-mail: jlelakow@szpitaljp2.krakow.pl
Received: 12.09.2006
Accepted: 27.01.2007
w w w .cardiologyjournal.org
155

---
Cardiology Journal 2007, Vol. 14, No. 2
amplitude and lack of mechanical stability are of minor importance. Elderly patients with sick sinus
syndrome and tachy-brady syndrome have the highest failure rate. (Cardiol J 2007; 14: 155–159)
Key words: arrhythmia episode, high pacing threshold, low atrial potential
amplitude, lack of mechanical stability, venous anomalies, DDD pacing
with this problem. There are, however, separate
Introduction
reports of electrode insertion during atrial flutter/
In most cases permanent pacing is performed
/fibrillation and at low atrial potential amplitude.
with an endocavitary electrode. The electrodes are
inserted via the veins of the upper thorax and neck
Methods
[1–6] and, exceptionally, via the femoral vein (in
According to the guidelines for electrotherapy
patients with superior vena cava syndrome) [7].
of the heart [1], apart from achieving appropriate
Exposure of the cephalic vein in the brachiot-
electrode placement, it is necessary to optimise
horacic groove is regarded as equivalent to subcla-
sensing and pacing parameters and to achieve at
vian venipuncture technique [2–5]. When these two
least the threshold values of:
main access sites are not available, it is recommend-
— ventricular potential amplitude > 4 mV;
ed to use the veins lying between the pectoral mus-
— rate of rise of ventricular potential > 0.5 V/s;
cles or superficial jugular veins, side branches of the
— atrial potential amplitude > 2 mV;
internal jugular vein (the inferior thyroid vein or
— rate of rise of atrial potential > 0.2 V/s;
transverse vein of the neck), the internal jugular
— pacing threshold < 1.0 V/0.5 ms [1].
vein itself and, finally, venipuncture or exposure of
Not all episodes of supraventricular arrhythmia
the internal jugular vein [7].
were the cause of a change in pacing mode. Cardiac
Implantation of a dual-chamber pacemaker in-
arrhythmia that did not resolve spontaneously and
volves placement of a ventricular electrode followed
was resistant to anti-arrhythmic drugs forced us to
by an atrial electrode inserted via the same or an-
abandon implantation of a DDD pacemaker. Cardio-
other vein. Appropriate fixation of the atrial elec-
version was not performed during the procedure.
trode (similar to the ventricular electrode) is con-
Mechanical stability of the electrode was con-
figured by a slight elevation of up to 0.2 mV of the
firmed during the Valsalva manoeuvre and while the
AV segment in the endocardial lead. Where there
is difficulty in placing atrial (ventricular) electrodes,
patient was coughing. The myocardial pacing
threshold was checked during these manoeuvres.
it is preferable to use active fixation electrodes [2].
In patients with venous anomalies (cephalic
To sum up, during implantation of a DDD pace-
maker the following difficulties may be encoun-
vein, other tributaries of the subclavian vein and
difficulty in puncture of the latter) we made an at-
tered:
tempt to implant the pacemaker on the contralat-
— venous anomalies (absence of vessels of ade-
quate calibre or difficulty in subclavian vein
eral side. In most cases the procedure was success-
ful. The implantation procedure was always per-
puncture);
formed by the most experienced operator. Contrast
— arrhythmias during implantation (episodes of
atrial flutter/fibrillation while the atrial leads
medium was injected to a peripheral vein to visual-
ise abnormalities. Occlusion of the subclavian vein
are being positioned);
and absence of a tributary of adequate calibre were
— lack of mechanical stability of the electrode in
the heart chamber;
the most frequent anomalies.
Passive fixation electrodes were applied and in
— inability to achieve an acceptable pacing and
the event of mechanical instability the screw-in type
sensing threshold during implantation (mainly
in the atrium).
was used. Atrial electrodes were placed in the right
atrial appendage, whereas, until 2004 ventricular
These difficulties may be the reason for chang-
electrodes were placed at the right ventricular apex.
ing the primary decision regarding DDD pacemak-
A pacing threshold in the atrium above 2 V was re-
er implantation. The purpose of the study was to
garded as unacceptably high.
analyse retrospectively the reasons for DDD pace-
We retrospectively reviewed all implantation
maker failures in patients operated on between 1993
data from 1988 to 2005 to identify patients with
and 2005. A review of the available literature shows
primary failure of DDD pacemaker implantation.
that there are no publications dealing extensively
156
www.cardiologyjournal.org

---
Jacek Lelakowski et al., Analysis of reasons for failure of DDD pacemaker implantation
Table 1. The number of pacemaker implanta-
32 with AVB, 22 with SSS, 16 with SSS ± AVB and
tions in consecutive years.
38 with TBS. The mean age of these patients was
78.5 ± 19.4 years and there were 46 men (43%).
Years
No. of all implantation/
/no. of DDD pacemakers
Statistical analysis
1993
465/43 (9%)
Data are expressed as mean ± 1 standard de-
1994
469/88 (18%)
viation and as a percentage. Student’s t-test was
1995
465/81 (17%)
used for analysis. A p value below 0.05 was consid-
1996
506/103 (20%)
ered significant.
1997
524/132 (25%)
1998
500/131 (26%)
Results
1999
479/135 (28%)
2000
532/103 (19%)
The reasons for abandoning implantation of DDD
2001
587/131 (22%)
pacemakers were venous anomalies in 13 (12%), an
2002
707/174 (24%)
episode of atrial flutter/fibrillation during the proce-
2003
806/271 (33%)
dure in 30 (27.8%), a high pacing threshold in the atri-
2004
646/264 (40%)
um in 19 (17.6%), low atrial potential amplitude in
2005
783/302 (38%)
28 (25.9%) and lack of mechanical stability of the elec-
trode in 18 (16.7%) patients (Table 2).
We were most frequently forced to abandon
the implantation of a DDD pacemaker in SSS and TBS
(71 cases) as a result of an arrhythmia episode,
These patients finally received a VVI pacemaker.
low amplitude and lack of mechanical stability
As the number of DDD pacemakers implanted be-
(Table 3).
tween 1988 and 1992 was too small (below 8%),
Between 1993 and 2003 excessively low atri-
further analysis included patients who had received
al potential amplitude and arrhythmia episodes
a DDD pacemaker between 1993 and 2005, when
(atrial flutter/fibrillation) during the procedure
this type of pacemaker made up between 9 and 40%
were the main causes of failure. Of note is the
of all pacemaker implantations. At this time we
decrease in the rate of unsuccessful insertion of
implanted 7469 pacemakers, including 1958 (26.2%)
atrial electrodes with increasing experience and
dual-chamber pacemakers, in 783 patients with atri-
technological progress in electrodes and pace-
oventricular block (AVB), 392 with sick sinus syn-
makers and between 2004 and 2005 the absence
drome (SSS), 450 with AVB ± SSS and 333 with
of failure as a result of arrhythmia episodes, low
tachy-brady syndrome (TBS). The mean age of the
atrial potential amplitude and inappropriate me-
patients was 65.5 ± 17.3 years. DDD pacing was
chanical stability of the electrode. Only single
unsuccessful in 108 (1.4%) patients, including
cases of venous anomaly and a high pacing thresh-
32 with AVB, 22 with SSS, 16 with SSS±AVB and
old were encountered. Patients in whom the pri-
38 with TBS (Table 1). The mean age of the patients
mary procedure was abandoned were older than
at the time of implantation was 78.5±19.4 years
those with successful implantation of a DDD pace-
(16–94) and 3361 (45%) were men. DDD pacing was
maker (p < 0.01).
unsuccessful in 108 (1.4%) patients, including
Table 2. Reasons for failure of DDD pacemaker implantation in consecutive years.
Reasons for failure of DDD
1993 1994 1995 1996 1997 1998 1999 2000 2001 2002 2003 2004 2005
pacemaker implantation
Venous anomalies 13 (12%)
0
1
2
1
0
1
1
2
1
3
0
0
1
Episode of atrial flutter
1
1
3
5
3
5
4
5
1
0
2
 0
0
or fibrillation 30 (27.8%)
High pacing threshold 19 (17.6%)
2
2
1
0
0
1
0
2
2
0
1
5
3
Low atrial potential amplitude
4
2
11
2
2
2
1
1
2
1
0
0
 0
28 (25.9%)
Lack of mechanical stability
3
4
5
2
1
1
1
1
0
0
0
0
 0
18 (16.7%)
157

---
Cardiology Journal 2007, Vol. 14, No. 2
Table 3. Reasons for failure of DDD pacemaker implantation according to indication. Patients with indi-
cations to pacemaker implantation (N = 108).
Reasons for failure of DDD
AVB
SSS
SSS + AVB
TBS
pacemaker implantation
 (n = 32)
 (n = 22)
(n = 16)
(n = 38)
Venous anomalies — 13
4
3
4
2
Episode of atrial flutter or fibrillation — 30
2
7
5
16
High pacing threshold — 19
5
4
5
5
Low atrial potential amplitude — 28
4
8
4
12
Lack of mechanical stability — 18
1
5
3
9
AVB — atrioventricular block, SSS — sick sinus syndrome, SSS + AVB — double node disease, TBS — tachy-brady syndrome
Unfortunately it is not possible to provide fluor-
The rate of failure decreased with increasing
oscopy times during the implantation of DDD pace-
experience and progress in pacemaker technology
makers in consecutive years. It is now 10 minutes
[9–11]. In 1997 the introduction of bipolar active
on average; prior to 1998 it was much longer, at
fixation electrodes (screw-in endocavitary elec-
about 25 minutes.
trodes) and multiprogrammable pacemakers with
special anti-tachycardia features and atrial sensitiv-
ity programmed to 0.1 mV allowed for better me-
Discussion
chanical stability of the electrode and implantation
On 4 October 1984 the first dual-chamber pace-
at low atrial potential amplitude (< 2 mV). Further-
maker (multiprogrammable Medtronic Symbios
more, atrial flutter/fibrillation during insertion of the
7006) was implanted in a 57-year-od man with sec-
atrial electrodes is no longer a significant problem.
ond-degree AVB and transiently third-degree AVB
According to Kindermann et al. [12], in the event
with Morgani-Adams-Stokes syndrome. We used
of unexpected arrhythmia atrial electrodes should
the original bipolar electrodes (a straight ventricu-
be placed in locations with potential amplitudes
lar electrode and a J-shaped atrial electrode). The
above 1.0 mV [13–15]. This is a useful approach to
ventricular electrode was fixed as usual at the right
achieving sufficient atrial sensing after restoration
ventricular apex. The atrial electrode was inserted
of sinus rhythm and the only remaining problems
via venipuncture of the subclavian vein under fluor-
between 2004 and 2005 were venous anomalies and
oscopic control and fixed in the right atrial append-
a high pacing threshold.
age. By the end of 1985, probably for financial rea-
The high rate of arrhythmias (27.8%) during the
sons, only three more dual-chamber pacemakers
procedure was found mainly in patients with the TBS.
had been implanted [8]. In 1993 the number of im-
Several investigators [12, 13, 16] report rates rang-
planted DDD pacemakers began to increase stead-
ing from 1 to 5.5%, but mainly in patients with
ily (> 9%) and as a result later analysis included
pure SSS, in whom, under normal conditions, the
patients who had received a DDD pacemaker be-
rate of arrhythmias is much lower. Anomalies of the
tween 1993 and 2005. A total of 7469 pacemakers
cephalic vein (absent vein or hypoplasia) occur in
were implanted, of which 26.2% were dual-cham-
about 10% of patients undergoing pacemaker im-
ber pacemakers. In 2006 dual-chamber pacemakers
plantation (about 0.6% in the present study) and,
made up 49% of all implantations.
according to Wiegand et al. [17], in about 2% of pa-
In 108 patients selected for dual-chamber pac-
tients undergoing pacemaker implantation it was
ing the procedure was abandoned and VVI pacing
impossible to achieve correct electrical parameters
was used. These patients were older (78.5 years),
(in 2.4% in the present study).
probably with greater myocardial injury and 55% of
Despite the limitations of the present study
them had SSS and TBS. The reasons for failure were
related to its retrospective character some impor-
low atrial potential amplitude (< 2 mV), an arrhyth-
tant conclusions can be drawn.
mia episode (atrial flutter/fibrillation) during manip-
ulation of the atrial electrode, lack of mechanical
Conclusions
stability and a high atrial pacing threshold (> 2 V).
1.
Currently the main reasons for abandoning im-
The first three causes were most frequently en-
plantation of dual-chamber pacemakers are
countered in patients with SSS and TBS.
158

---
Jacek Lelakowski et al., Analysis of reasons for failure of DDD pacemaker implantation
venous anomalies and a high pacing threshold
7.
Bellot PH. Unusual access sites for permanent cardiac
(> 2 V). Low atrial potential amplitude, ar-
pacing. W: Barold SS (ed) Recent advances in cardiac
rhythmia episodes and lack of mechanical sta-
pacing. Futura Publishing Company, Inc. Armonk,
bility are of minor importance.
New York 1998: 137–180.
2.
Elderly patients with sick sinus syndrome and
8.
Książka Główna Kliniki Elektrokardiologii w Kra-
tachy-brady syndrome have the highest failure rate.
kowskim Szpitalu Specjalistycznym im. Jana Pawła II.
9.
Hayes DL, Furman S. Cardiac pacing: how it started,
where we are, where we are going. PACE, 2004; 27:
References
693–704.
1.
Świątecka G (ed) Standardy postępowania w ele-
10.
Furman S. The early history of cardiac pacing. PACE,
ktroterapii serca. Folia Cardiol, 1999; 1 (suppl. 1).
2003; 26: 2023–2032.
2.
Holmes DR, Hayes DL. Pacemaker implantation
11.
Aubert AE, Ector H. New lead technology for the
st
techniques. W: Saksena S, Goldschlager N (eds)
21
 century. In: Antonioli GE (ed) Pacemaker leads.
Electrical
therapy
for
cardiac
arrhythmias.
Monduzzi Bologna, Italy 1997: 17–24.
W.B. Saunders Company, Philadelphia 1990: 173–
12.
Kindermann
M,
Frohlig
G,
Berg
M,
Lawall
P,
–190.
Schieffer H. Atrial lead implantation during atrial flut-
3.
Bellot PH, Byrt CL. Recent developments in pace-
ter or fibrillation? PACE, 1998; 21: 1531–1538.
maker implantation and lead retrieval. In: Barold SS,
13.
Fyfe T, MacIntyre PD, Robinson JF, Clark AL. Blind
Mugica J (eds) New perspectives in cardiac pacing.
atrial pacing for patients with sinus node disease who
Futura Publishing Company, Inc. Mount Kisco, New
develop atrial fibrillation during permanent pacemak-
York 1991: 105–131.
er implantation. Inter J Cardiol, 1997; 58: 188–191.
4.
Midei M, Brinker J. Pacemaker implantation. In:
14.
Kerr CR, Mason MA. Amplitude of atrial electrical
Ellenbogen KA (ed) Cardiac pacing. Blackwell Scien-
activity during sinus rhythm and during atrial flutter-
tific Publications, Boston 1992: 211–262.
fibrillation. PACE, 1985; 8: 348–355.
5.
Hayes DL, Holmes DR, Furman S. Permanent pace-
15.
Marshall H, Griffith M, Grammage M. Atrial lead
maker
implantation.
In:
Furman
S,
Hayes
DL,
placement during atrial fibrillation. Is it safe? PACE,
Holmes DR (eds) A Practice of Cardiac Pacing. Futura
1997; 20 (4): A484 (abstract).
Publishing Company, Inc. Mount Kisco, New York
16.
Sutton R, Kenny RA. The natural history of sick
1993: 261–307.
sinus syndrome. PACE, 1986; 9: 1110–1114.
6.
Bellot PH, Reynolds DW. Permanent pacemaker im-
17.
Wiegand UKH, Bode F, Bonnemeier H, Eberhard F,
plantation. In: Ellenbogen KA, Neal Kay B, Wilkoff BL
Schlei M, Peters W. Long-term complication rates in
(eds) Clinical cardiac pacing. W.B. Saunders Company,
ventricular, single lead VDD, and dual chamber pac-
Philadelphia 1995: 447–490.
ing. PACE, 2003; 26: 1961–1969.
159