© Pol J Radiol 2022; 87: e24-e29
DOI: https://doi.org/10.5114/pjr.2022.112466
Received: 02.09.2020
Accepted: 26.10.2020
Published: 07.01.2022
http://www.polradiol.com
Original paper
Efficacy of low dose and ultra-low dose on the visibility of peri-implant
fenestration and dehiscences: a computed tomography study
Ceren Aktuna-Belgin1,A,B,D,E,F, Gozde Serindere1,A,D,E,F, Huseyin Berkay Belgin2,A,B,C,D, Mehmet Serindere3,A,C,D,F,
Kaan Orhan4,A,C,D,E
1Department of Dentomaxillofacial Radiology, Faculty of Dentistry, Hatay Mustafa Kemal University, Hatay, Turkey
2Private Dental Clinic, Turkey
3Hatay Education and Research Hospital, Hatay, Turkey
4Department of Dentomaxillofacial Radiology, Faculty of Dentistry, Ankara University, Ankara, Turkey
Abstract
Purpose: This study aimed to evaluate the visibility of peri-implant fenestration and dehiscences on computed tomo- graphy (CT) images taken with 2 different doses.
Material and methods: The defects were created on the apical of 6 implants randomly selected from 20 titanium im- plants placed in the ribs, and dehiscences were created on the cervical of 8 implants. No defects were created around 6 implants. Macroscopic analysis of the implanted ribs was accepted as the gold standard. From the samples, images were taken by using both ultra-low dose (80 kVp, 50 mA, 1.25 mm slice thickness) and low dose (100 kVp, 50 mA,
1.25 mm slice thickness) protocols in CT. The images obtained were evaluated using a 5-point scale.
Results: No significant difference was found between the area under the receiver operating characteristic of ultra-low dose protocol and low dose protocol in both defects based on the Wilcoxon test (p > 0.05).
Conclusions: The ultra-low dose protocol could be applied by adhering to the “as low as reasonably achievable” prin- ciple in the diagnosis of peri-implant defects.
Key words: computed tomography, radiation dose, implant, fenestration, dehiscences.
placed [3]. Insufficient bone in the buccolingual direction
can lead to bone fenestrations and dehiscences formation, which compromise the long-term prognosis of dental im- plants [4]. Dehiscences is a buccal or facial alveolar bone- limited defect that exposes the roots of the dental implant or teeth from the alveolar crest to the apical. Fenestration is a window-shaped defect of 3 mm or more when the im- plant or tooth root is covered with gingival tissues rather than bone [5].
The determination of alveolar bone defects plays a role in determining the need for grafts before surgical procedures [6]. However, insufficient bone thickness can lead to negative
Introduction
For dental implants to be considered clinically success-
ful, should absent pain and tenderness during function, mobility, or abscesses. It is also stated that stable, rigid implants with a pocket depth of 2 to 6 mm can be consid- ered successful [1]. When examined radiologically, verti- cal marginal bone loss at the peri-implant surface should not exceed 1-2 mm in the first year after implant loading and 0.2 mm in later years [2]. Another factor that deter- mines the success of dental implants is whether there is sufficient bone volume in the area where the implant is
Correspondence address:
Dr. Ceren Aktuna-Belgin, Department of Dentomaxillofacial Radiology, Faculty of Dentistry, Hatay Mustafa Kemal University, Hatay, Turkey, e-mail: dtcaktuna@gmail.com
Authors’ contribution:
A Study design ∙ B Data collection ∙ C Statistical analysis ∙ D Data interpretation ∙ E Manuscript preparation ∙ F Literature search ∙ G Funds collection
This is an Open Access journal, all articles are distributed under the terms of the Creative Commons Attribution-Noncommercial-No Derivatives 4.0
International (CC BY-NC-ND 4.0). License (https://creativecommons.org/licenses/by-nc-nd/4.0/).
e24

Different dose efficacy in the visibility of peri-implantitis
consequences such as inaccurate implant placement during
surgery, inflammatory biofilm, and occlusion errors. More- over, it can lead to gingival recession, loss of alveolar bone, failing osseointegration, and implant failures [3,5,7,8].
The most common method for assessing marginal bone loss is conventional periapical radiography [1]. In im- plants, bone loss after the 1st placement occurs mostly on the facial or buccal side of the implant because the bone is thinner in these areas. However, periapical radiographs can detect bone loss in the mesial or distal. Because the buccal or facial area is superposed on 2-dimensional periapical radiographs, it is inadequate to show early bone destruc- tion [9,10]. In computed tomography (CT), a 3-dimen- sional evaluation without superposition can be detected buccolingual or mesiodistal alveolar bone destruction [11]. Although CT has been shown to yield accurate and reliable assessments for orthodontic and maxillofacial applications, its effective dose is much higher than that of conventional radiography or cone beam CT (CBCT) [12].
As a result of technological advances in the field of dentistry, both digital imaging techniques and 3-dimen- sional imaging techniques have been introduced in the routine. Especially since the introduction of CBCT de- vices, the interest in 3D imaging has increased rapidly in the last decade [13,14].
Although CBCT is routinely using for implant plan- ning, CT is also used in some instances where CBCT is not available. Moreover, with the introduction of modern multiple-detector computed tomography (MDCT) tech- nology and continuing technological developments, es- pecially iterative reconstruction algorithms, low dose CT examinations are currently employed, such as in lung and maxillofacial imaging. Various kV and mAs values have been used in MDCT with different estimated effective doses. A recent study, aimed at determining the distribu- tion of effective dose in paranasal sinus imaging associated with a single low dose CT examination concluded that ac- ceptable CT screening can be accomplished at an overall average computer tomography dose index (CTDI) of ap- proximately 2.9 mGy [15].
Also, the ALARA (as low as reasonably achievable) principle, which is an underlined view in radiation safety, emphasizes that the radiation dose given to living struc- tures is the lowest possible value without affecting diagnosis and treatment. Following this principle, the most appropri- ate dose protocol should be selected based on the reliability and accuracy of the clinical diagnostic result and radiation dose [16].
Moreover, no extensive data on the use of low dose CT are available, and further studies are needed to assess the usefulness of this technique for maxillofacial and dental imaging.
The purpose of this study was to evaluate the visibility of peri-implant fenestrations and dehiscences for dental implant exposures from ultra-low dose and low dose CT set-ups. The null hypothesis indicated that there was no
significant difference in the detection of peri-implant de-
fects at the 2 different dose set-ups on the detection of peri-implant defects.
Material and methods
Study design
The sample size was calculated with the GPower 3.1 soft- ware (http://www.gpower.hhu.de/); the total mean of 2 groups compared based on Student’s t-test with the power of 80% and 0.05 type 1 error was found to be at least 13 patients.
Based on the study of de-Azevedo-Vaz et al. [4], soft
tissue-free ribs were used to represent the alveolar bone. Fresh beef ribs were obtained from a local butcher from slaughtered animals. A total of 20 dental implants (NTA, Pilatus Swiss Dental, Egolzwil, Switzerland) were used in diameter range, 3.7-5.1 mm; length range, 8.5-10 mm. After placement of 20 titanium implants, similarly to the literature [17], 3-mm diameter fenestrations and de- hiscences were created on the rib. In the apical of 6 ran- domly selected implants, fenestrations were formed in elliptical form 10 mm apical to the upper edge of the rib. In the cervical of 8 implants, half-elliptical form dehis- cences were formed on the alveolar crest. No defects were created around the other 6 implants, and they were evalu- ated as a control group.
Radiographic imaging
CTs were taken after the ribs were placed in a water-filled container to resemble soft-tissue attenuation. CT images were performed with a 256-slice CT scanner (Revolution CT; GE Healthcare). To evaluate the visibility of the created defects in 2 different doses, the samples were firstly taken with the ultra-low dose protocol (80 kVp, 50 mA, 1.25 mm slice thickness). Then the same samples were scanned with the low dose protocol (100 kVp, 50 mA, 1.25 mm slice thickness). The size-specific dose estimates in CT (SSDE) of the ultra-low dose protocol and low dose protocol were
6.66 mGy and 12.14 mGy, respectively. The dose length protocol (DLP) of the ultra-low dose and low dose protocol was 77.32 mGy cm and 171.57 mGy cm, respectively. The CTDI of the ultra-low dose protocol and low dose protocol was 5.95 mGy and 10.90 mGy, respectively (Figure 1).
Sample analysis
CT images were evaluated by 2 oral radiologists (C.A.B. and G.S., 5 and 6 years of clinical experience), 1 radiolo- gist (M.S., 6 years of clinical experience), who did not know whether there were defects. A 5-point scale was used to evaluate the images (1 – defect absent; 2 – de- fect probably absent; 3 – uncertain; 4 – defect probably present; 5 – defect present). Macroscopic analysis of the
© Pol J Radiol 2022; 87: e24-e29
e25

Ceren Aktuna-Belgin, Gozde Serindere, Huseyin Berkay Belgin et al.
A
B
C
D
Figure 1. Sagittal slice on computed tomography images of titanium implant with simulated peri-implant defects with 2 different dose protocols.
A) Dehiscences on low dose, B) fenestration on low dose, C) dehiscences on the ultra-low dose, D) fenestration on the ultra-low dose
implanted ribs was accepted as the gold standard (de-
fect present or absent) [4]. Evaluations were made with a RadiAnt Dicom Viewer (2020.1, Medixant, Poznan, Poland). Observers were allowed to view the images in 3 planes (sagittal, axial, coronal), change the contrast or brightness of the images, and use the rotate or zoom tool.
ultra-low dose protocol was good, and the inter-observer
agreement of the low dose protocol was excellent.
The Az values and p-values for the diagnosis of peri- implant defects are shown in Table 2. The diagnosis of observers was equal at both dose protocol for fenestration and dehiscence diagnosis without a significant difference. Also, there was also no significant difference between Az values of the ultra-low dose and low dose protocol in both defects.
Statistical analysis
Kappa’s was obtained regarding compatibility among the observers’ agreement on the correct diagnosis of the defect types between the 2 protocols using SPSS 26.0. (SPSS Inc., Chicago, IL). Based on Landis and Koch [18], low agreement – 0.40; moderate agreement – 0.40-0.59; good agreement – 0.60-0.74; and excellent agreement
– 0.75-1.00 were considered. The areas under receiver operating characteristic (ROC) curves (Az values) for the 2 protocols for each observer were computed from the ROC-curve and compared with the Wilcoxon-test at a 5% significance level. Also, the sensitivity (the correct identification of a defect presence), specificity (the correct identification of a defect absence), accuracy (proportion of correctness), positive predictive value (PPV: the prob- ability that a positive evaluation is a true-positive result), and negative predictive value (NPV: the probability that a negative evaluation is a true-negative result) were cal- culated for each protocol. In the study, assessment of the compatibility of diagnosis between observers revealed a group of individuals in the group of people who have the case group and in the control group.
Table 1. Values of kappa’s for inter-observer agreement
Table 2. Value of the area under the receiver operating characteristic (ROC) curve (Az values) and p-value for each observer and protocols
 Results
The average and standard deviations of the kappa  values
are shown in Table 1. The inter-observer agreement of the
Obs – observer
e26
© Pol J Radiol 2022; 87: e24-e29
Ultra-low dose	Low dose
Az values (p-value) Az values (p-value)
Fenestration

Obs 1
1 (0.004)
1 (0.004)
Obs 2
1 (0.004)
1 (0.004)
Obs 3
1 (0.004)
1 (0.004)
Mean (SD)
1
1
p-value
1.00
Dehiscences

Obs 1
1 (0.004)
1 (0.002)
Obs 2
0.969 (0.004)
1 (0.002)
Obs 3
0.958 (0.005)
0.937 (007)
Mean (SD)
0.975 ± 0.02
0.979 ± 0.03
p-value
0.15
Fenestration, mean (SD) Dehiscences, mean (SD)
Ultra-low dose
0.64 (0.03)
0.62 (0.11)
Low dose
0.72 (0.17)
0.75 (0.14)

Different dose efficacy in the visibility of peri-implantitis
Table 3. Diagnostic values for fenestration and dehiscences using the tested protocols
ultra-low dose and low dose protocols. Although in the
literature there are studies visibility of peri-implant defects according to voxel size, the field of view, and the use of various algorithms and filters, there are few articles about dose-related [6,17,29-31]. Values from Misch et al. [30] (47.7 mA, 120 kVp) and Pinsky et al. [31] (98 mA, 120 kVp) were used as the scan protocol for the detection of fenestra- tions and dehiscences. In both studies, it was stated that ar- tificially created fenestrations and dehiscences were identi- fied and the measurement accuracy was high. In our study, 2 different scan protocols, 50 mA 100 kVp and 50 mA 80 kVp, were used. Although a lower rate of mA and kVp was used than the work of both Misch et al. [30] and Pin- sky et al. [31], defects were detected and there was no sta- tistically significant difference between them. Bayrak et al. [17] evaluated the visibility of peri-implant defects with the metal artifact reduction algorithm and the adaptive image noise optimizer filter. As a result, they stated that the use of both filters together would be useful in detecting defects. In another study [29], 2 different voxel sizes (0.2 mm3 and
0.12 mm3), and two different scan modes (180° and 360°), were used to detect defects. It was stated that the voxel size did not affect the visibility of defects, but dehiscences were seen more clearly with a 360° scan mode. Kamburoglu et al. [6] investigated the effectiveness of the field of view in detecting peri-implant defects and stated that there was no relationship between the field of view and defect vis- ibility. Interobserver agreement was good to excellent for the 2 different doses protocols in the present study. In this study, the inter-observer agreement of the ultra-low dose protocol was good, and the inter-observer agreement of the low dose protocol was excellent. However, de-Azevedo-Vaz et al. [4] found that inter-observer agreement was moderate to excellent in their study. The reason for this may be that the imaging techniques used in both studies are different.
PPV – positive predictive value, NPV – negative predictive value
Of all implants, 6 had fenestrations, 8 had dehiscences,
and 6 had no defect. The sensitivity (the correct identifi- cation of a defect presence), specificity (the correct iden- tification of a defect absence), accuracy (proportion of correctness), PPV, and NPV for the studied protocols are shown in Table 3. Comparison of both defects showed that values were higher for fenestrations than for dehiscences, but they were similar across protocols.
Discussion
Peri-implantitis is defects that can develop due to poor oral
hygiene, history of periodontitis, bacteria, systemic diseases such as diabetes, smoking, or technical issues such as im- plant surface, type, structure, and osseointegration degree [19,20]. Peri-implant defects such as fenestrations and de- hiscences are factors that decrease the success of the implant [1]. In addition to decreasing implant success, undiagnosed alveolar bone defects can also lead to a relapse of orthodon- tic treatment [21] or gingival recession that will result in an unaesthetic ending of orthodontic treatment [22]. Also, cases of peri-implantitis associated with endodontic problems have been reported in the literature [23].
Jeffcoat et al. [24] reported that 2-dimensional imaging
is strongly affected by vertical and horizontal angulation errors during film exposure. Therefore, 3-dimensional im- aging is recommended for the visualization of periodontal defects [25]. Radiation can cause carcinogenesis genetic damage to irradiated tissues [26]. Radiation is known to cause DNA damage with a low probability of ionization ef- fect [27]. Therefore, radiography should be taken with the lowest radiation dose that will not affect the diagnosis and treatment, without damaging the biological structure in the area of interest [28].
In the present study, the effectiveness of the 2 differ- ent dose protocols invisibility of peri-implant defects were compared. As a result of this study, it was observed that there was no statistically significant difference between
Limitations of the study
This study has some limitations. The first is that the
number of samples is low. It is thought that studies car- ried out on larger samples will yield more realistic results. Another is that the bone structure in which defects are formed did not belong to human bone. Although it is sup- ported by the literature that cow ribs are similar to the human mandible bone in terms of bone density and cor- tical-cancellous bone relationship [32-34], the ribs do not mimic the human bone. Another limitation is the use of implants of different sizes and diameters. Whether differ- ent sizes and diameters of implants will change the result of the study may be the subject of another study.
Conclusions
In this study, it was found that there was no difference
between the ultra-low dose protocol and the low dose pro- tocol in imaging bone defects. In the radiological exami-
© Pol J Radiol 2022; 87: e24-e29
e27
Parameter	Ultra-low dose	Low dose
Fenestration

Sensitivity
1
0.7
Specificity
1
1
PPV
1
1
NPV
1
1
Accuracy
1
0.85
Dehiscences

Sensitivity
1
0.91
Specificity
1
0.62
PPV
1
1
NPV
1
1
Accuracy
1
0.76

Ceren Aktuna-Belgin, Gozde Serindere, Huseyin Berkay Belgin et al.
nation of cases with peri-implantitis defects, it is thought
that a lower dose protocol can be applied to patients fol- lowing the ALARA principle.
Conflict of interest
The authors report no conflict of interest.
References
Misch CE, Perel ML, Wang HL, et al. Implant success, survival, and
failure: the International Congress of Oral Implantologists (ICOI) Pisa Consensus Conference. Implant Dent 2008; 17: 5-15.
1. Smith DE, Zarb GA. Criteria for success of osseointegrated endos- seous implants. J Prosthet Dent 1989; 62: 567-572.
2. Blanco J, Alonso A, Sanz M. Long-term results and survival rate of implants treated with guided bone regeneration: a 5-year case series prospective study. Clin Oral Impl Res 2005; 16: 294-301.
3. de-Azevedo-Vaz SL, Peyneau PD, Ramirez-Sotelo LR, et al. Effica- cy of a cone-beam computed tomography metal artifact reduction algorithm for the detection of periimplant fenestrations and de- hiscences. Oral Surg Oral Med Oral Pathol Oral Radiol 2016; 121: 550-556.
4. Leung CC, Palomo L, Griffith R, et al. Accuracy and reliability of cone- beam computed tomography for measuring alveolar bone height and detecting bony dehiscences and fenestrations. Am J Orthod Dentofacial Orthop 2010; 137: 109-119.
5. Kamburoğlu K, Murat S, Kılıç C, et al. Accuracy of CBCT images in the assessment of buccal marginal alveolar peri-implant defects: effect of field of view. Dentomaxillofac Radiol 2014; 43: 20130332.
6. Mengel R, Kruse B, Flores-de-Jacoby L. Digital volume tomography in the diagnosis of peri-implant defects: an in vitro study on native pig mandibles. J Periodontol 2006; 77: 1234-1241.
7. Sancho-Puchades M, Hammerle CH, Benic GI. In vitro assessment of artifacts induced by titanium, titanium-zirconium, and zirconium dioxide implants in cone-beam computed tomography. Clin Oral Implants Res 2015; 26: 1222-1228.
8. Kamburoglu K, Gulsahi A, Genç Y, et al. A comparison of peripheral marginal bone loss at dental implants measured with conventional intraoral film and digitized radiographs. J Oral Implantol 2012; 38: 211-219.
9. Schliephake H, Wichmann M, Donnerstag F, et al. Imaging of periimplant bone levels of implants with buccal bone defects. Clin Oral Implants Res 2003; 14: 193-200.
10. Pompa V, Galasso S, Cassetta M, et al. A comparative study of Mag- netic Resonance (MR) and Computed Tomography (CT) in the pre-implant evaluation. Ann Stomatol (Roma) 2010; 1: 33-38.
11. Oz U, Orhan K, Abe N. Comparison of linear and angular measure- ments using two-dimensional conventional methods and three-di- mensional cone beam CT images reconstructed from a volumetric rendering program in vivo. Dentomaxillofac Radiol 2011; 40: 492-500.
12. Shah N, Bansal N, Logani A. Recent advances in imaging technolo- gies in dentistry. World J Radiol 2014; 6: 794-807.
13. Suomalainen A, Pakbaznejad Esmaeili E, Robinson S. Dentomaxil- lofacial imaging with panoramic views and cone beam CT. Insights Imaging 2015; 6: 1-16.
14. Hoxworth JM, Lal D, Fletcher GP, et al. Radiation dose reduction in paranasal sinus CT using model-based iterative reconstruction.
AJNR Am J Neuroradiol 2014; 35: 644-649.
Kaplan DJ, Patel JN, Liporace FA, et al. Intraoperative radiation safe-
ty in orthopaedics: a review of the ALARA (as low as reasonably achievable) principle. Patient Saf Surg 2016; 10: 27.
1. Bayrak S, Orhan K, Kursun Çakmak ES, et al. Evaluation of a metal artifact reduction algorithm and an optimization filter in the estima- tion of peri-implant dehiscence defects by using cone beam comput- ed tomography: an in-vitro study. Oral Surg Oral Med Oral Pathol Oral Radiol 2020; 130: 209-216.
2. Landis JR, Koch GC. The measurement of observer agreement for categorical data. Biometrics. 19 Am J Orthod Dentofacial Orthop 2006; 130: 476-484.
3. Lindhe J, Meyle J, Group D of European Workshop on Periodonto- logy. Peri-implant diseases: consensus report of the sixth European workshop on periodontology. J Clin Periodontol 2008; 35: 282-285.
4. Smeets R, Henningsen A, Jung O, et al. Definition, etiology, preven- tion and treatment of peri-implantitis – a review. Head Face Med 2014; 10: 34.
5. Rothe LE, Bollen AM, Little RM, et al. Trabecular and cortical bone as risk factors for orthodontic relapse. Am J Orthod Dentofacial Orthop 2006; 130: 476-484.
6. Yared KF, Zenobio EG, Pacheco W. Periodontal status of mandibular central incisors after orthodontic proclination in adults. Am J Orthod Dentofacial Orthop 2006; 130: 6.e1-8.
7. Daubert D, Black RM, Chrepa V, et al. Endodontic peri-implant de- fects: a new disease entity. J Endod 2020; 46: 444-448.
8. Jeffcoat MK, Wang IC, Reddy MS. Radiographic diagnosis in perio- dontics. Periodontol 2000 1995; 7: 54-68.
9. Noujeim M, Prihoda T, Langlais R, et al. Evaluation of high-reso- lution cone beam computed tomography in the detection of simu- lated interradicular bone lesions. Dentomaxillofac Radiol 2009; 38: 156-162.
10. Garib DG, Calil LR, Leal CR, et al. Is there a consensus for CBCT use in orthodontics? Dental Press J Orthod 2014; 19: 136-149.
11. Ludlow JB, Timothy R, Walker C, et al. Effective dose of dental CBCT-a meta-analysis of published data and additional data for nine CBCT units. Dentomaxillofac Radiol 2015; 44: 20140197.
12. American Academy of Oral and Maxillofacial Radiology. Clinical recommendations regarding use of cone beam computed tomogra- phy in orthodontics. [corrected]. Position statement by the American Academy of Oral and Maxillofacial Radiology. Oral Surg Oral Med Oral Pathol Oral Radiol 2013; 116: 238-257.
13. de-Azevedo-Vaz SL, Vasconcelos Kde F, Neves FS, et al. Detection of peri implant fenestration and dehiscences with the use of two scan modes and the smallest voxel sizes of a cone-beam computed to- mography device. Oral Surg Oral Med Oral Pathol Oral Radiol 2013; 115: 121-127.
14. Misch KA, Yi ES, Sarment DP. Accuracy of cone beam computed to- mography for periodontal defect measurements. J Periodontol 2006;
77: 1261-1266.
e28
© Pol J Radiol 2022; 87: e24-e29

Different dose efficacy in the visibility of peri-implantitis
Pinsky HM, Dyda S, Pinsky RW, et al. Accuracy of three-dimensional
measurements using cone-beam CT. Dentomaxillofac Radiol 2006; 35: 410-416.
1. May KB, Edge MJ, Russell MM, et al. The precision of fit at the im- plant prosthodontic interface. J Prosthet Dent 1997; 77: 497-502.
2. Chiodo TA, Ziccardi VB, Janal M, et al. Failure strength of 2.0 lock- ing versus 2.0 conventional Synthes mandibular plates: a laboratory model. Oral Maxillofac Surg 2006; 64: 1475-1479.
3. Choi BH, Huh JY, Suh CH, et al. An in vitro evaluation of miniplate fixation techniques for fractures of the atrophic edentulous mandible.
Int J Oral Maxillofac Surg 2005; 34: 174-177.
© Pol J Radiol 2022; 87: e24-e29
e29