©PolJRadiol,2010;75(3):81-86
Varia
Numberofactivedetectorrowsandqualityofimageincomputedtomography
DariuszChruściak1,MariaA.Staniszewska2
1RegionalSanitaryEpidemiologicalStation,Kielce,Poland
2DepartmentMedicalImagingTechniques,MedicalUniversity,Łódź,Poland
Summary
RapidtechnicaldevelopmentofComputedTomography(CT)meansusagethemodern,fastscannerswhichpossibilitiesarenotalwaysknownbyusers.Acommonconvictionisthehighnumberofdetectorsisequivalenttoanaccurateinvestigationofsub-millimetreslicesofpatientbodyanddetectionofsuchsmall-sizelesions.
AimofthepaperwastocompareanimagingpossibilitiesofthetwoCTscannersofthesamevendorswhichdifferonlytheconfigurationofdetectorsinuse.(Thegroupof48rowsofdetectorisphysicallythesameforthebothscanners.)
QualityofimageswascontrolledusingCatphanseries500.Thehigh-contrastresolution,low-contrastresolutionandsignal-to-noiseratiowereevaluatedforimagesrecordedattheconstantreconstructionthicknessof5mm.
Aprecision of imagingwas evaluated onthe basis signal-to-noiseratio (SNR) inrelation to CTDoseIndex.ThelowestSNRatsatisfyingvisualisationofthelow-contrastobjectswasslightlyabove5.ThecorrespondingCTDoseIndexwasover5mGy.Care-doseoptiondecreasesSNRtonearly2whatfollowsunacceptablevisualisationoflow-contrastobjects.
onthebasistheresultsforthebothscannerswasfoundthathighnumberofactivedetectorrowshelptoachieveabettervisualisationoflow-contrastobjectsunderconditionofhigherCTDoseIndexvalues.
Thespaceresolutiondoesnotchangewiththedetectorconfiguration.
Background:
Material/Methods:
Results:
Conclusions:
Background
ofavailabilityofverythinslices(forexample0.5mm).Formostofpeoplesthatmeansapossibilitytopictureinpatientbodyofverysmalldetailsofsizeequaltominimalcollimation.Suchthinkingisfrequentlyaddedtothefan-tastic“demo”pictures,especiallyfromdynamicproceduresplusinformationaboutdose-savingtoolsinsoftwareofthescanner.Inconsequencethepurchaserexpectskindof“wizardtool”foreverydaypractice.
AnintensivetechnicaldevelopmentofCTiscontinuedfrom2000.Thevendorscompeteinconstructiveandsoft-waresolutionsaimedtomakeatimeofscanningshorterandtoreducethedosestopatients.Shorteningthetimeisachievedthroughthehighernumberofactivedetectorsandshortertimeofrotation.Thescannerswith256oreven320activedetectorrows(orratherbands)areavailabletoday,togetherwiththeothershaving64rowsofdetectorsandrotationtimelowerthanhalfofsecondor–havingtwox-raytubes workingsimultaneously.
However,thetruepossibilitiesofCTscannerinsignificantlevelaredeterminedbytheinstalledversionofitssoft-ware,andagreatnumberofdetectorrowsisnotusedandnecessaryinalltheprocedures.
Sohighnumberofdetectorsneedthegreatamountofinformationwhatfollowsthehighintensityofradia-tionandhighdosestopatients.Softwaremethodsofdosereductionreliedonthemodulationofx-rayintensityinscanningline(i.e.OZ-axis)andtheplaneofslice(i.e.OXYplane).Anothersolutionappliedinthedynamicpro-ceduresisthegatingofsignal(forexampleinthecardiacprocedures).
AimofthepaperistocompareimagingpossibilitiesofCTscannersthesamevendor,whichdifferoftheconfigurationofdetectorsinuse.
MaterialandMethods
EvaluationofpossibilitiesofCTscanneriscompletedfromtwoparts:aqualityofimageandagenerallyunderstanddose.
KnowledgeofthepotentialpurchasersofnewCTscan-nersaboutrealpossibilitiesoftheequipmentofferedtothemisrathersmall.Thegreatnumberofdetectorrowsisunderstandastheruleashighlypositivefeature,because
QualityofCTimageisdeterminedbythefactorsaccept-edforeverydigitaltechniques,i.e.:levelofnoise–oreven
81

Varia
©PolJRadiol,2010;75(3):81-86
more-signal-to-noiseratio,high-contrastresolutionandlow-contrastresolution.
The image created in CT techniqueis the space distributionofattenuationcoefficientsoftheminiaturevolumeele-mentsofthescannedobject(voxels),onwhichananalysedslaboftheobjectisdivided.Thatallowstopictureneigh-bouringstructuresindetails,althoughtheirabsorptionfea-turesdifferminimally.Themethodtogetherwithcomputa-tionanalysebythesoftwareallowstoachieve3D-imagesathighlevelofcontrast.
Thevaluesofattenuationcoefficientsarerelatedtowater,asabasecomponentofhumantissues,andgiveninHounsfieldunits(HU).ThatmeansthedependenceaswellqualityandformalcontentofCTimageonx-rayenergy,i.e.voltage(kV)andfiltrationonx-raytube.(Thisisthediffer-encebetweenconventionalx-rayradiography,wheretheimageisanegativemapofdensitydistribution.Moreover,inconventionalradiographynumberofgreylevelsisnearly100,whereastherangeoftheHounsfieldscaleis4096lev-els.Thatmakepossibletodifferentiatethegreylevelsforthemediaofverysimilarabsorptionfeatures,likethepar-ticularsofttissues.
Figure1.CATPHANarrangedtoexposure.
andstandarddeviationofthevaluewastreatedas“noiselevel”.
ThesecondpartofevaluationofCTexposureisdosimetry,meansaslevelofpatientirradiation.ThecommonlyusedquantityistheComputedTomographyDoseIndex(CTDI),definedas

CTDI1D(z)dz
ThebasicindicatorsofCTimagequality(i.e.low-contrastresolution,signal-to-noiseratioandhigh-contrastresolu-tion)areestimatedusingthededicatedphantoms.Themanufacturersofthescannersofferedownconstructionsof the phantom. Although these are similar to eachother,asobjectivearetreatedresultsobtainedwithuse“aneutral”versioni.e.CATPHANphantoms(madebyThePhantomLaboratoryIncorporated,Salem,NY)preparedproperlyforinvestigatedtypeofscanner.
w

whereD(z)–dosedistributionalongOZaxis,w-slicethickness.
Inpractice,fromthemeasurementsisdeterminedso-called“volumeCTDI”:
DLPważ
InthepaperimagequalitywasanalysedusingCATPHANseries500,designedtocontrolofhelicalmodeofwork(Figure1).
CTDI
vol
NΔzp
whereN–thenumberoftheactivedetectorrows,z–thethicknessofsingledetector,p–pitchfactor.
Low-contrastresolutionwasestimatedonthebasisoftheimagesoftheCATPHANsection515.Boththe“supra-slice”and“sub-slice”elementswere analysed.Thefirsthavethicknessequaltothatofthesection,butthepartsofdif-ferentcontrastlevel.Thesecondhaveconstantcontrastof1%,but contain thepartsof differentthickness. The resultsobtainedfromanalyseofthesection515allowtoevaluateofdetectionleveloflow-contrastandsmall-sizedobjectsduringthescanninginhelical modeofwork.
DLP–theweightedmeanvalueofDoseLengthProduct:
1	2
DLPważ=3DLPc+3DLPp
DLPCmeansthevalueinthecentreofdosimetricalphantom,
Fromtheimagesof“supra-slice”objectsthelow-contrastresolutionwasevaluatedfordetailsofnominalcontrastlevelof1%.Visualisationof“sub-slice”objectsrelaysoncheckingthediameterofthesmallestdetailof5mmthick-ness(i.e.thinnerthenthesection).
DLPPmeansthevaluesforperipheryofthephantom.
CTDIvolhavetobeunderstandasanaveragedvalueofdoseabsorbedinascannedvolume.Thus,itcanbetreat-edasameasureofpatientirradiationatthegivenexpo-sureparameters,andthevalueofCTDIvolisdisplayedonscannerconsole(aftertheselectionparametersoftheprocedure).
High-contrastresolutionwasevaluatedonthebasisimagesofCATPHANsection528andgiveninlinepairsperlengthunit.
Signal-to-noiseratio(SNR)wasevaluatedforregionofinterest(ROI)of100mm2beinglocalizedaroundtheobjects ofnominal contrastlevelof 1%.CT numbers(inHUunits)forthegivenROIwereinterpretedas“signallevel”
TakingtheaboveintoconsiderationthebestCTscannershouldallowtoproducedetailimagewithouthighdosestopatients.Howmuchthehighnumberofdetectorrowscanhelpinthatwasthegoalofthepresentedpaper.
82

©PolJRadiol,2010;75(3):81-86
Varia
Table1.TheindicatorsofimagequalityforthebothCTscannersatthicknessofreconstructionslab5mm.
Low-contrastobjects
High-contrastresolution[lp/cm]
CT
scanner
mAs/scan
Activedetectors
CTDI
[mGy]
kV
p
SNR
Supra-slicedmin(1%)[mm]
Sub-slicedmin(5mm)[mm]
24×1.2
(24)
7
7**
144
120
1.2
7.87
10.75
3
3
24×1.2
(24)
7
7**
66
120
0.55
6.99
10.79
3
3
24×1.2
(24)
7
7**
144
100
1.2
4.84
6.75
4
7
24×1.2
(24)
7
7**
66
100
0.55
4.05
6.77
4
5
24×1.2
(24)
7
7**
72
120
1.2
5.59
5.38
4
5
24×1.2
(24)
7
7**
33
120
0.55
5.26
5.40
5
5
40×0.6
(40)
7
7**
144
120
1.2
7.28
13.03
3
3
40×0.6
(40)
7
7**
66
120
0.55
7.88
13.03
3
5
40×0.6
(40)
7
7**
144
100
1.2
5.10
7.98
4
5
40×0.6
(40)
7
7**
66
100
0.55
4.74
7.98
5
5
40×0.6
(40)
7
7**
72
120
1.2
5.23
6.52
5
5
40×0.6
(40)
7
7**
33
120
0.55
6.10
6.52
5
5
27
(Care-dose)
24×1.2
(24)
7
7**
120
1.2
3.43
2.43
9
–
29
(Care-dose)
24×1.2
(24)
7
7**
120
0.55
3.95
2.70
9
–
26
(Care-dose)
40×0.6
(40)
7
7**
120
1.2
3.75
2.85
9
–
29
(Care-dose)
40×0.6
(40)
7
7**
100
1.2
2.28
1.77
15
–
34
(Care-dose)
40×0.6
(40)
7
7**
100
0.55
2.38
2.30
–
–
27
(Care-dose)
40×0.6
(40)
7
7**
100
1.2
1.98
1.52
8
–
34
(Care-dose)
40×0.6
(40)
7
7**
100
0.55
2.55
1.95
8
–
Results
ForthispurposewereperformedtheexperimentsontwoCTscannersofthesamevendor,whichdifferonlyinthenumberofactivedetectortowsandremainingconstructivesolutionsandsoftwarewerenearlythesame.Thequalityofimageswascheckedforfindingtheindicatorsmostlydependentonconfigurationofthedetectors.
TheinvestigationswereperformedfortwoversionsofCTscannerbySiemens:SensationOpenandCardiac.
Physically,thebothversionsofscannersareequippedintothesamecollectionof48rowsofdetectors,whichindivid-ualwidthisz=0.6mm.
83
SensationOpenSIEMENS,∆z=0.6mm

Varia
©PolJRadiol,2010;75(3):81-86
Table1continued.TheindicatorsofimagequalityforthebothCTscannersatthicknessofreconstructionslab5mm.
Low-contrastobjects
High-contrastresolution[lp/cm]
CT	mAs/
scanner	scan
Activedetectors
CTDI
[mGy]
kV
p
SNR
Supra-slicedmin(1%)[mm]
Sub-slicedmin(5mm)[mm]
64×0.6
(32)
7
8*
224
120
1.4
9.94
12.25
2
3
24×1.2
(24)
7
8*
224
120
1.4
8.57
10.79
2
5
67
(Care-dose)
64×0.6
(32)
7
8*
120
1.4
5.69
5.19
4
5
67
(Care-dose)
64×0.6
(32)
7
8*
120
1.4
7.10
5.19
3
5
70
(Care-dose)
24×1.2
(24)
7
8*
120
1.4
5.55
4.72
4
5
70
(Care-dose)
24×1.2
(24)
7
8*
120
1.4
6.96
4.72
3
5
64×0.6
(32)
7
8*
224
100
1.4
6.02
7.11
4
3
24×1.2
(24)
7
8*
224
100
1.4
6.35
6.28
4
5
*highcontrastresolutionforimagestransformed“bone”function;**highcontrastresolutionforimagestransformed“osteo”function.
Numberofactivedetectorrows(i.e.participatingintoimageregistration)fortheparticularprocedureisdeter-minedbythewayconsolidationofsignalswhichispre-scribedinthesoftwareofthescannerversion.
InSensationOpentwocombinationsareusedalternative-ly:24×1.2mmand40×0.6mm.
Cardiacversionalsousesthecombination24×1.2mmandthesecond64×0.6mm.Thelastisrealizedastwo-foldexposureinto32rowsofdetectors(andconsequentlythesignalsrecording)of0.6mmwidth(so-calledFlyingFocalSpot).
Thecombination24×1.2mmisrealizedthroughthecou-plingofsignalsfromthepairsofcontiguousdetectors.Theremainingcombinationsmeansinpracticeaneliminationthesignalsfromthedetectorssituatedontheedgeofarchfromtherecording(i.e.fromimageregistration).(Thisdoesnotmeantomakenarrowthedoseprofile.)
Figure2.ThescannerSensationOpen:relationshipbetweenSNRandCTDIvolatdetectorconfiguration24×1.2mmandU=120kV.
thatmeansonlytheuniformloweringvaluemAsperscantothemeanvaluegivenintheTable1.
Astart-pointforinvestigationsofthebothscannerswasprotocolofroutineabdomen.Itsparameterswerethenmodifiedthroughmakingloweranodecurrentandtubevoltageandchangedpitchforcheckinganeffectonimagequality.
AscanbeseeninTable1theimageevaluationwaspossibleonlyforselectedassemblyofparameters:puttingon“Caredose”algorithmfollowedbydramaticloweringofvisibilityoflow-contrastobjects.
TheresultsobtainedforevaluatedimagesaregiveninTable1.TherearetheCTDIvolvaluesdisplayedon thescan-nerconsolefortheparametersofCATPHANexposure.TherowsinTable1denoted“Caredose”containtheresultsobtainedatactivatedalgorithmofautomaticdosereduc-tion(i.e.modulationofx-rayintensity).FortheCATPHAN
ThemaininterestininterpretationoftheobtainedresultswasputonrelationshipbetweenSNRandCTDIvol,asthefirstis“avisual”measureofimagequalityandthesecondisameasureofpatientdose.
TheevaluateddependencesaregraphicallypresentedontheFigures2–6.
84
SensationCardiac64SIEMENS,∆z=0.6mm
SNR

9
8
7
6
5
4
3
2
1
00	1	2	3	4	5	6	7	8	9101112
CTDI[mGy]

y=−0.0311x2+0.8716x+1.6475R2=0.9652

©PolJRadiol,2010;75(3):81-86
Varia
Figure3.ThescannerSensationOpen:relationshipbetweenSNRandCTDIvolatdetectorconfiguration40×0.6mmandU=120kV.
Figure5.ThescannerCardiac:relationshipbetweenSNRandCTDIvolatdetectorconfiguration24×1.2mm.TheresultsobtainedatU=120kVandatU=100kVareshowntogether.
Figure4.ThescannerSensationOpen:relationshipbetweenSNRandCTDIvolatdetectorconfiguration40×0.6mmandU=100kV.
Figure6.ThescannerCardiac:relationshipbetweenSNRandCTDIvolatdetectorconfiguration40×0.6mm.TheresultsobtainedatU=120kVandatU=100kVareshowntogether.
Table2.ParametersofexposureandindicatorsofimagequalitycorrespondingtoSNR=6.Low-contrast(LC)andHigh-contrast(HC)resolutionwereevaluatedforimagesrecordedwiththethicknessofreconstructionslab5mm.
HC[lp/cm](osteo/bone)
Detectors
U
CTDI(displayed)
LC-supra(1%)
LC-sub(5mm)
SensationOpen
24×1.2
120kV
7mGy
4mm
5mm
7
40×0.6
120kV
7.5mGy
5mm
5mm
7
max.SNR=5atCTDI=8mGy:
thenLC=5mmbothforthesupra-sliceobjectsandforthesub-sliceobjects
40×0.6
100kV
Cardiac
24×1.2
120+100(*)
5.5mGy
4mm
5mm
8
64×0.6
120+100(*)
7mGy
4mm
3mm
8
(*)–fortheCardiacscannertheresultsforthebothvoltagevaluesarejointedastheSNRvs.CTDIrelationshiparenotdifferfortheir.
TheresultsforthescannerCardiacwerecollectedatpitchvalue1.4andtubevoltageU=120kVandU=100kV.AstheseresultsdifferedonlyslightlytheyaretakentogetherforevaluationthedependenceSNRonCTDIvol.
OnthebasisoftheobtainedresultswasconcludedthelowestlevelofSNRatwhichlow-contrastdetailsarewelldistinguishedforCardiacscannerwaslittleabove
5.ThereforewasdecidedtocheckthevalueofCTDIvol
85
SNR
SNR
SNR
SNR

12
10
8
6
4
2
0
01	234	56	7	8910111213
CTDI[mGy]

y=−0.0837x2−0.8574x+7.8858R2=1

6
5
4
3
2
1
0
0	1	2	3	4	5	6	7	8	9
CTDI[mGy]

y=−0.0185x2+0.6128+1.2101R2=0.9844

9
8
7
6
5
4
3
2
1
00	1	2	3	4	5	6	7	8	9101112
CTDI[mGy]

y=−0.0034x2+0.5501x+3.029R2=1

9
8
7
6
5
4
3
2
1
001234567891011121314
CTDI[mGy]

y=−0.0224x2+0.7313x+1.8474R2=0.9484

Varia
©PolJRadiol,2010;75(3):81-86
necessarytoobtainSNR6forimagesfromthebothscan-nersandatdifferentconfigurationofdetectors.
24×1.2mmonthehighernumberofactiverows(40or64–respectively).(Thisisdespiteconstancyoftheallremainderparameters.)
1. Thevisualisationoflow-contrastobjectsplacedinamiddleofthescannedslab(i.e.alevelofLC-subslice)forthescannerCardiacbecomesbetterwithchangefrom24×1.2mmon64×0.6mm.
2. Atsignificantreductionofx-rayintensity(Caredoseoption)thescannerSensationOpenwithconfigurationofdetectors40×0.6mmgivesverylowlevelofSNR(2,i.e.athresholdofimageacceptability).LCresolutionhasjustsolowlevelthatimagesarenotapplicabletoevaluation,evenatthereconstructionlayerof5mm.
Moreoverthelevelsoflow-contrastresolution(LC)andhigh-contrastresolution(HC)correspondingtoSNR6werealsoevaluated.
ThisallowstoestimateanimagingprecisionatconstantnominallevelofSNRand“adoseexpense”correspondingtosuchqualityofimage(Table2).
Conclusions
Analysisoftheobtainedresultsallowstodrawthefollow-ingconclusions.
C.Tosummarise
A.High-contrastresolution(HC)
Abenefittoimagesresultingfromthehighnumberofactivedetectorrowsisvisibleforthelow-contrastobjects,butneedthehigherx-rayintensities.
1. ConfigurationofactivedetectorrowsnotaffectonHClevel.
2. ForthescannerSensationOpenthelevelofHCisnotbetterdespitetransformationofimagesthrough“bone”filteringfunction(acontentthesameasfor“standard”function).ForthescannerCardiacthelevelofHCtrans-formationthrough“bone”filteringfunctionmakestheimagesalittlebetter(8lp/cmat“bone”insteadof7lp/cmatstandard).
3. HC leveldo notchange forreduced x-rayintensity.
Thegoodimagequalityisnotachievablewhenthethick-nessofrecordedslicesaresmallandthedoseisreduced.
High-contrastobjectsarewellvisibleevenatlownumberofactivedetectorrowsandrelativelylowx-rayintensi-ties.Consequently,thehighexposureparametersformulti-detectorrowsscannersisnotjustified.
Arealbenefit fromhigh numberof activerowsof detectorsis alsodependent onthe software:eventhe sameconstruc-tionofthescannercangiveanotherresultsforimages,whentheusedversionsofsoftwarearedifferent.
B.Visualisationoflow-contrastobjects
1.ForthebothinvestigatedscannersCTDIvolincrease
when
the
configuration
of
detectors
changes
from
86