Signature:©PolJRadiol,2009;74(4):43-46
ORIGINALARTICLE
Useofcomputertomographyand3DPRapidPrototyping
Received:2009.07.20
Accepted:2009.08.19
techniqueincranioplastyplanning–analysisofaccuracyofbonedefectmodelling
OlimpiaMarkowska1,AdriannaGardzińska1,RobertChrzan2,AndrzejUrbanik2,SławomirMiechowicz1
1InstituteofMachineConstruction,FacultyofMechanicalEngineeringandAeronautics,RzeszówUniversityofTechnology,
Rzeszów,Poland
2DepartmentofRadiology,JagiellonianUniversity,CollegiumMedicum,Cracow,Poland

Author’saddress:RobertChrzan,RadiologyDepartmentoftheJagiellonianUniversity,CollegiumMedicum,Kopernika19Str.,31-501Cracow,Poland,e-mail:rchrzan@mp.pl
Summary
Theaccuracyconsiderationsofalignmentofskullbonelossandartificialmodelofimplantarepresented.Instandardsurgicaltreatmenttheapplicationofprefabricatedalloplasticimplantsrequirescomplicatedproceduresduringsurgery,especiallyadditionalgeometryprocessingtoprovidebetteradjustmentofimplant.RapidPrototypingcanbeusedasaneffectivetooltogeneratecomplex3Dmedicalmodelsandtoimproveandsimplifysurgicaltreatmentplanning.Theoperationtimecanalsobesignificantlyreduced.Theaimofthestudyisadjustmentaccuracyanalysisbymeasurementsoffissurebetweenthebonelossandimplant.
The3DnumericalmodelwasobtainedfromCTimagingwithSiemensSensation10CTscanner.Thephysicalmodelswerefabricatedwith3DPRapidPrototypingtechnology.Themeasurementswereperformedindeterminedpointsofthebonelossandimplantborders.
Maximalwidthoffissurebetweenbonelossandimplantwas1.8mmandminimal0mm.Averagewidthwas0.714mm,standarddeviation0.663mm.
Accuracyof 3DP techniqueis enough tocreate medicalmodels in selectedfield of medicine.ModelscreatedusingRPmethodsmaybethenusedtoproduceimplantsofbiocompatiblematerial,forexamplebyvacuumcasting.Usingofmethodsuggestedmayallowshorteningofpresurgeryandsurgerytime.
Background:
Material/Methods:
Results:
Conclusions:
Keywords:
computedtomography(CT)•rapidprototyping(RP)•craniectomy
PDFfile:
http://www.polradiol.com/fulltxt.php?ICID=900376
Background
numericalmodelsofspecifichumanbodyparts,withanaccuracyappropriateformedicalmodelling[1,2].Thesemodels,togetherwiththemethodsofrapidprototyping,enablethescientiststoconstructfunctionalprototypesofallkindofprosthesesandimplantsformedicalpurposes.
Rapid prototypingisaconstructionofa prototypicalphysi-calobject(single)withtheuseofa3Ddigitalmodelcon-sistingofaspatialgridcomposedoftrianglesrepresentingthesurfaceoftheobject.ThesavingformatofsuchafileiscalledSTL(STereoLitography)andiscommonlyknownintechnologicalenvironmentsasaninputformatofdataforthemajority ofRP machines.
OneofthepotentialapplicationsofRPisdesigningtheprostheticrestorationofcranialbonedefects,forthepur-posesofcranioplasty.Tocreatesuchamodel,oneoftheRPtechniquescanbeused–3DP(3DPrinting)[3].Thisisfinanciallyprofitable[4].Averyimportantissueisalsotheaccurateperformanceofasuitableprostheticelement[5,6].
Computedtomography,beinganimagingtechniqueofahighspatialresolution,allowsforaconstructionof
43

OriginalArticle
©PolJRadiol,2009;74(4):43-46
Figure1.GeneratedSTLsurfacemodel:(A)headshapereproductionwithavisibledefect
(B)skullfragmentdefectandimplantprojectmodels.
Figure2.Physicalmodelsofdefectandimplant.
Theaimoftheworkwastoanalysetheaccuracyoftheadjustmentoftheprojectedprostheticelements(implants),forthepurposesofcranioplasty,performedwitha3DPtechnology,andtoanalysethesurfaceofcontactofcranialbonedefect.
technique(Figure2).Thetechnologyofthisprocessbasedonthecreationoflayersformingthemodel.TheCAD(STL)modelmaybesegmented;eachsegmentincludesspecificinformationonthegeometryofonelayer(2D).Thesubse-quentlayersofthemodelarecreatedbylayingdownthelayersof apowder.3DPtechnology issimilar to theprocessofinkprinting;thebindercementsthepowderwiththepreviouslayeronlyinthepointinwhichthemodelistobeformed.Then,theplatformdescends,allowingforanapplicationofthenextlayerofthepowderwhichwillgetboundwiththepreviousone.Theprocessrepeatsuntilthecreationofthemodelofthewholeobject.Afterremovingtheremainingreinforcingmaterial–theunboundpowder–theobjectgetsinfiltrated.Thetypeofinfiltratorinfluencesthemechanicalpropertiesofthemodel[3,4].
Materialandmethods
Dataacquisition–preparationofanumericalmodel
Thedatasetintheformofarastergraphicsimageisabasisfora3Dvectormodel.Wecollectedthesedataintheprocessofacquisition,withtheuseofamulti-row-detectorspiralCTscannerbySiemensSensation10.Theregionofscanningincludedheadwithcranialbonedefect(Figure1A).Wepreparedthenumericalmodelandpro-cesseditwiththeuseof3D-DOCTORsoftwarebyAbleSoftwareCorp.
Themodelwascreatedwiththeuseofa3DPZ510appara-tusbyZCorp.
Theobtained2Dimagesweresubjectedtosegmentation,
i.e.selectionofelementsconstitutingthebasisforthecon-structionofamodelofthecraniumpartwiththedefect.Contourimagesgeneratedonthebasisofsuchsegmenta-tionincludedinformationabouttheouterandinneredgesoftheselectedelements.Thecollectionofcross-sectionswasusedtocreatea3Dpresentationintheformofaspa-tialgridcomposedoftrianglesrepresentingthesurfaceoftheobject,intheSTLformat(STereoLitography).
Theanalysisofaccuracyinalignmentoftheimplantedges
ThemodelofbonedefectprosthesiswasadjustedwiththeuseofMaterialiseMagicsprogram.Thepartofthemodeloftheundamagedsideofthecraniumwassymmetricallymirror-reflectedagainstthemedialbodyaxis;next,theareaofthedefectwasdeducted.Asaresult,weobtainedadigitalmodeloftheprostheticelement(implant)recon-structingthecranialbonedefect.Theanalysisoftheaccu-racyofthealignmentbetweentheimplantandthemodelofthedefectwascarriedoutin18selectedpoints,situ-atedalongtheedgeofthedefect(Figures3,4).Themea-surementswereperformedwithanopticalreflection-lightmicroscope,Dino-LiteDigitalMicroscopePROAM413T,using60×zoom(Figure5).
The conversionofdatatoasurface modelallowedforafullmodelmanipulationinCADsystems,additionofsomeextrageometricalfeaturesorconstructionofamodelofprosthet-icrestorationofabonedefect(Figure1B).
Constructionofaphysicalmodelofthecranialbonedefectwiththe3DPmethod
Results
Theobtainednumericalmodelconstitutedabasisforthephysicalmodeloftheimplant,performedwith3DP
Table1presentsresultsonmodeladjustment.
44

A	B

©PolJRadiol,2009;74(4):43-46
MarkowskaOetal.–Useofcomputertomographyand3DP…
Figure3.Prototypeofdefectwithimplant,measurementpointsmarked(implantimagebeforefixation, forbetterpresentation).
Figure4.Prototypeofdefectwithamountedimplant.
Figure5.Edgeimagezoomedbymicroscope.
Table1.Resultsofthemeasurementsofthefissurebetweenthedefectanditsimplant;measurementsmadewiththeopticalreflection-lightmicroscope(Dino-LiteDigitalMicroscopePROAM413T).
Measurement
Fissurewidth(mm)
Measurement
Fissurewidth(mm)
Measurement
Fissurewidth(mm)
1.
2.4
7.
0.2
13.
0.9
2.
1.8
8.
0.25
14.
0.65
3.
0.6
9.
0.4
15.
1.8
4.
0.6
10.
0.65
16.
0.15
5.
0.45
11.
0
17.
0.1
6.
0.2
12.
0.7
18.
1.0
Thewidthofthe fissurerepresentsthefailure inprosthesisadjustment.Themaximalwidthofthefissureamountedto
1.8mm,andtheminimal:0mm.Meanwidthofthefissure,
i.e.meanerror,wasapprox.0.714mm;standarddeviation0.663mm.
withtheuseofthatprosthesis.Suchaphysicalmodelofbonedefectisthentheonlybasis–templateforafutureimplant.Thatiswhythereappearsaneedforadevelop-mentofamethodwhichwouldevaluatetheaccuracyoftheadjustmentofthosetwoelements.Theanalysedprob-lem results fromthe acceptedtechnological assumptionsofboneprosthesiscreation.
Discussion
Theissueofmedicalmodelconstructionismoreandmorefrequentlytoucheduponinthecurrentliterature[5–12].However,themajorityofthedescribedcasesfocusonthemethodofphysicalmodelconstruction,whichmaybethenusedforthepurposesofsurgeryplanning,creationof abiocompatible prostheses,and performanceof surgery
Normally,thesurgicalproceduresareperformedwiththeuseofprefabricatedprosthesesofstandarddimensionsandcurvatures,whicharethenadjustedduringtheprocedure[7].Thistechnologymakesitimpossibletoadjusttwoele-mentsprecisely,especiallywhen thegeometry ofthe defecthasadifficultshape.
45

OriginalArticle
©PolJRadiol,2009;74(4):43-46
Whenanalysingtheaccuracyoftheaforementionedadjust-ments,weshouldalsotakeintoconsiderationthemethodoffixation.Theprojectofprosthesispresupposedimplantfixationalongoneofthechosenaxes,againstthesurfaceofprostheticelement.Theanalysisofedgecollisionbasedonthatidealpresupposition.Inpractice,oneshouldexpectdifficultiesinprosthesisfixation,whichmayresultfromthepresenceofedgefragmentsthatcollidewithacompli-catedlateralsurfaceofthedefect.Thatiswhy,whenpro-jectingthemodel,weshouldtakeintoconsiderationthedegreeofcomplexityofthelateralsurfaceandincreasethefreespacewithinthefissureinthenumericalmodel,andapplyatleasttheconicalanalysisofcollisionoftheadjustedsurfaces.Lateralinnersurfaceoftheprosthesisedgeshouldundergosomeadditionalphasing,tomakeitsfixationeasier.Phasingdegreedependsonthecomplexityoftheinnersurfaceoutlineoftheprosthesis.
tothedifficultieswithprosthesisfixation,itisnecessarytoleavealargerfissure.
Easierfixationandanoptimaladjustmentsignificant-lyshortenthetimeoftheprocedure.Anidealsituationwouldbesuchapreparationoftheprocedurethatwouldnotrequirethesurgeontoperformanysupplementaryactivitiesapartfromthestandardmedicalprocedures;hewouldthenfocusonimplementingtheoptimallyadjustedimplantofthebonedefect.Toobtainthisaim,itisnec-essarytosolveawholerangeofproblemsconnectedwiththemodelsandthetechnologyitself.Asalreadymentionedabove,everytechnologyrequiresconsiderationofdiffer-entmodelingconditions.Theyincludei.a.theaccuracyofagivenmethod,appliedmaterials,dimensionalandtechno-logicallimitations.Whenprojectingthemodelwiththeuseofagiventechnology,itisnecessarytotakeintoconsid-erationthedifferencesinperformanceaccuracy,depend-ingonthemachineused(althoughtheyareallofthesametechnology).
Anotherproblemissharpedgesoftheoutersurfaceoutline.Inthenumericalmodel,thesharpnessofthoseedgesistheresultofapreciseadjustmentoftheoutersurfaceofthedefectandtheprosthesis(switchbetweenthesurfaces,alongthetangent).Thesharpnessbecomesamajorprobleminthephysicalmodel.Thisresultsfromtheabsoluteminimalvalueofthelayerthatcanbeproducedwithagivenmethodofrapidprototyping.Inmostcasesthisvaluedoesnotexceed
0.1mm[8,9].Structuresofalesservaluearenotcreated.
Theadvantageoftheappliedmethodofnumericalmodel-ingisitshigheffectiveness.Itallowsforafastandeasydeductionofthedefectareafromthemodeloftheundam-agedpartofthecranium,reflectedsymmetricallyagainstthemedialaxisofthebody(‘mirror-reflection’).Thismeth-odcannotbeusedforextensive,asymmetricaldamages.
Conclusions
Thenextproblemconnectedwithsharpedgesisthephysi-calpropertiesofthematerialthattheimplantismadeof.Infragileandbrittlematerialstheedgesmaybecomedamagedwhenwetrytopreciselyadjusttwoverycom-plexlateralsurfacesofthe prosthesisandbonedefect. Thatiswhy,whenchoosingtheRPmethod,itisnecessarytoconsiderthetypeofmaterialappliedandthusthemethodofedgeprotection(e.g.bymakingtheedgesmoreroundedandblunt).Inmostcases,thesatisfactoryaestheticeffectisobtained,evenwiththeedgesbeing1mmthick.Suchaninaccuracywillbenearlyinvisibleduetotheprosthesislocationanditscoveragebyskinandhair.Thatiswhy,due
Theaccuracyofthe3DPtechniqueissufficienttocreatemedicalmodelsineveryfieldofmedicine.
ModelscreatedwiththeRP(3DP)methodmaybethenusedtoformimplantsmadeofbiocompatiblematerials,
e.g.withvacuumcastingmethod.
Theapplicationofthesuggestedmethodmayshortenthepreoperativetime,aswellasthetimeofthesurgicalpro-cedureitself.
References:
1. UrbanikA,MiechowiczS,ChrzanR:Applicationofstereolitographyincreatingmedicalmodels.PolJRadiol,2005;(2):42–46
2. UrbanikA,MiechowiczS:MedicalmodelingwithCT,MRIandstereolitography.PolJRadiol,2004;69(Suppl.1):368
3. ZCorp:3DPrintingTechnology:ZCorporation,2009,http://www.zcorp.com
4. MarkowskaO,MiechowiczS:Ekonomiczneuwarunkowaniawykonaniamodelimedycznychwybranymitechnikamiszybkiegoprototypowania.SympozjonPKM,Przemyśl,2007
5. ChrzanR,UrbanikA,KarbowskiKetal:Wytwarzaniemodelimedycznychdlapotrzebzabiegukranioplastyki.IVKrakowskieWarsztatyInżynieriiMedycznej,Kraków,2007
6. MullerA,KrishnanKG,UhlEetal:Theapplicationofrapidprototypingtechniquesincranialreconstructionandpreoperativeplanninginneurosurgery.JCraniofacSurg,2003;(14):899–914
1. AndrzejakS,FortuniakJ,JaskólskiDetal:Supplementationofcranialdefectsbyanautologousboneflapstoredintheabdominalwall. NeurolNeurochirur,2005; (39): 220–24
2. RaicK,JansenT,vonRymon-LipinskiBetal:Fastgenerationofstereolithographicmodels.BiomedTech,2002;47(Suppl.):83–85
3. TanikawaY,Imai D,TanakaKet al:FabricationofRealisticandDynamicHumanHeadPhantoms.JSMEInternationalJournal,1999;423:487–91
4. ChiariniL,FigurelliS,PollastriGetal:Cranioplastyusingacrylicmaterial:anewtechnicalprocedure.JCraniomaxillofacSurg,2004;32:5–9
5. StoodleyMA,AbbottJR,SimpsonDA:Titaniumcranioplastyusing3-D computermodelingofskulldefects.JClinNeurosci.,1996;3:149–55
6. UedaK,ObaS,OmiyaYetal:Cranial-bonedefectswithdepressiondeformitytreatedwithceramicimplantsandfree-flaptransfers.BrJPlastSurg,2001;54:403–8
46