Acta Poloniae Pharmaceutica ñ Drug Research, Vol. 72 No. 6 pp. 1089ñ1099, 2015
ISSN 0001-6837
Polish Pharmaceutical Society
VALIDATION OF LC/MS/MS METHOD FOR ASSESSMENT OF THE
IN VITRO ACTIVITY OF SELECTED RAT CYTOCHROME P450 ISOENZYMES - APPLICATION TO EARLY DRUG METABOLISM SCREENING
KAMIL KUå1,2*, AGNIESZKA ZAKRZEWSKA1, MA£GORZATA SZAFARZ2,1, MARIA WALCZAK1, 3,
ANNA GONCIARZ1,2, AGNIESZKA KIJ1, JOANNA SURAJ1,2 and
2,1
1Jagiellonian Centre for Experimental Therapeutics (JCET), Jagiellonian University, BobrzyÒskiego 14, 30-348 KrakÛw, Poland
2Department of Pharmacokinetics and Physical Pharmacy, Jagiellonian University, Collegium Medicum, Medyczna 9, 30-688 KrakÛw, Poland
3Department of Toxicology, Jagiellonian University, Collegium Medicum, Medyczna 9, 30-688 Krakow, Poland
Abstract: A sensitive and specific liquid chromatography/tandem mass spectrometry (LC/MS/MS) method for simultaneous determination of seven metabolites of CYP450 model substrates (acetaminophen, 4-hydroxy- tolbutamide, 4í-hydroxymephenytoin, 1-hydroxybufuralol, 6-hydroxychlorzoxazone, 1í- and 4í-hydroxymida- zolam) in rat liver microsomes was developed. The assay used Kinetex analytical column and a gradient mobile phase consistent of acetonitrile and water with addition of 0.1% formic acid. The analysis was performed in selected reaction monitoring (SRM) mode both in positive and negative (for 6-hydroxychlorzoxazone) mode. The method was validated over the concentration ranges of 10-2000 ng/mL for 4í-hydroxymephenytoin and 4- hydroxytolbutamide, 50-2000 ng/mL for 1-hydroxybufuralol and 25-2000 ng/mL for the rest of the analytes. The intra- and inter-day precision (2-12%) and accuracy (93-119%) were within the limits set by the FDA and EMA guidelines. The developed method was successfully applied to assess the activity of selected CYP450 isoenzymes in rat liver microsomes after addition of ketoconazole.
Keywords: rat liver microsomes, CYP450 isoenzymes, cocktail method, LC/MS/MS, ketoconazole
The development of new drugs is determined largely by their ADME properties of which metabo- lism is an important factor affecting e.g., bioavail- ability, detoxification, drug-drug interactions and individual pharmacokinetics and pharmacodynam- ics variability. Drugs and xenobiotics are mainly metabolized by different oxidation and conjugation enzymes to more polar and better soluble metabo- lites to facilitate their excretion (1).
Cytochrome P450 system represents drug- metabolizing enzymes from a superfamily of hemo- proteins, engaged in phase I metabolism. Among the CYP isoforms, family 1 to 4 are involved in metab- olism of different xenobiotics, and in human more than 90% of drugs that undergo biotransformation are metabolized by CYP 1A2, 2A6, 2B6, 2C9, 2C19, 2D6, 2E1 and 3A4 (2, 3). In the development of new chemical entity, drug candidate, preclinical screening of metabolism rates and assessment of elimination routes are considered to be essential for further development.
In the field of drug discovery, a demand has risen for efficient and rapid drug metabolism screen- ing techniques. To speed up the selection of new drug candidates, to identify the metabolic profile of new molecules, to indicate likely drug interactions or to assess the role of polymorphic enzymes, dif- ferent in vitro systems have been developed.
LC/MS/MS technique provides a powerful solution for investigating the biotransformation of xenobiotics and is used at various stages of drug dis- covery and development (4). Because of the need to test a large number of compounds, recently a high- throughput screening methods were created. One of the application includes the use of the cocktail of substrates added to the microsomal incubations for simultaneous evaluation of the activities of CYP450 isoenzymes (5). Due to this approach multiple enzymes may be assessed in the same study and the effect of intra-individual variability is minimized. Recently, numerous of drug metabolism cocktails have been proposed and developed (6ñ15).
* Corresponding authors: e-mail: kamil.kus@uj.edu.pl; maria.walczak@jcet.eu; phone: +48 12 6645481, fax: +48 12 2974615
1089
JOANNA SZYMURA-OLEKSIAK

1090
KAMIL KUå et al.
According to FDA guidelines, preclinical metabolism studies of new drug candidates should be conducted using human microsomes (16), while there are numerous in vitro works for the assessment of preliminary metabolic pathway of new com- pounds using rat liver microsomes, cytosol and pri- mary hepatocytes, because rats are the most com- monly used model in the field of preclinical studies (17ñ20). Among the frequent examples, work of Xiao Hu et al. (21) demonstrated the variance analy- sis of Vmax, Km and Clint values from the study of enzyme kinetics of osthole and imperatorin, and showed no significant difference between these val- ues in rat and human liver microsomes. The study demonstrated that the rat liver microsomes were reliable to analyze drug metabolism in vitro. The study of De Bock et al. (22) described the validation of the method for determination of metabolites of model substrates in rat microsomal matrix using ear- lier derivatization of metabolites. The metabolic fate of sublethal exposure of different herbicides in rat microsomes was studied in the work of Larsen et al. (23). On the basis of the results of Xue-Jin Zhao et al. (24) one may assume that from a qualitative point of view, the dog, rat and even mouse could be an appropriate species for conducting an in vitro study on the metabolism of quinine. However, using rodent data to estimate human risk should be aware of interspecies variability both in expression and catalytic activity of relevant CYP450 isoforms (25).
The present study aims to evaluate a simple and an accurate LC/MS/MS method for simultane- ous determination of specific metabolites of six model substrates metabolized by CYP450 isoen- zymes after incubation with rat liver microsomes. In the next step, the optimized method was used to establish the strength of enzyme inhibition by keto- conazole on selected CYP450 isoenzymes.
Formic acid was obtained from Fluka (Buchs, Switzerland). Potassium dihydrogen phosphate, sodium chloride, acetone and acetonitrile were pur- chased from J.T. Baker (Phillipsburg, PA, USA). Purified water (18.2 M) used throughout the study was delivered by a Milli-Q water system (Millipore, Billerica MA, USA).
Isolation of hepatic microsomes
Rat liver microsomes were prepared using a differrential centrifugation method. Liver fragments were minced with scissors, washed with TRIS/KCl (pH 7.4) buffer and homogenized. The homogenate was centrifuged (Sorval WX Ultra Series, Thermo Scientific) at approximately 11 500 ◊ g for 20 min at 4OC. The supernatant (S9 fraction) was transferred to new centrifuge tubes and then ultracentrifuged at 100 000 ◊ g for 60 min at 4OC. The pellet was resus- pended in 0.15 M KCl with Ultra Turrax IKA T10 basic homogenizer (IKA-Werke GmbH & Co. KG Staufen, Germany) and ultracentrifuged again at 100 000 ◊ g for 60 min at 4OC. The obtained pellet was dispersed in TRIS/sucrose buffer and stored at
-80OC until use.
Protein concentration in microsomal fraction was determined by Lowry protein assay (26).
Chromatographic conditions
Chromatographic separation was performed on a Kinetex (2.6 µm, PFP, 100 A, 100 ◊ 3 mm, Phenomenex, Torrance, CA, USA) analytical col- umn using UltiMate 3000 UPLC system (Dionex, USA). The mobile phase consisted of acetonitrile containing 0.1% formic acid (eluent A) and water containing 0.1% formic acid (eluent B). At the flow rate of 500 µL/min, the amount of eluent A increased lineary from 20 to 95% in 2 min, main- tained by 95% for 6 min, then returned to 20% in 3 min, and was left to equilibrate for 6 min. The total run time was 15 min.
MATERIALS AND METHODS
Chemicals and reagents
Midazolam, 1í-hydroxymidazolam MZ), 4í-hydroxymidazolam (4-OH-MZ),
LC/MS/MS conditions
The analyzed compounds were detected using TSQ Quantum Ultra triple quadrupole mass spec- trometer (Thermo Scientific, Waltham, MA, USA). Heated electrospray ionization (H-ESI) was used both in the positive and negative (for 6-hydroxy- chlorzoxazone) modes. The selected reaction moni- toring transitions and collision energies optimized for each metabolite and internal standard are listed in Table 1. The mass spectrometric conditions were as follows: needle voltage 4020 V, vaporizer tem- perature 250OC, sheath gas (nitrogen) pressure 30 psi, auxiliary gas (nitrogen) pressure 10 psi, capil-
(1-OH-
phena-
cetin, acetaminophen (ACP), tolbutamide, 4- hydroxytolbutamide (4-OH-TB), bufuralol, 1- hydroxybufuralol (1-OH-BF), S-mephenytoin, 4í- hydroxymephenytoin (4-OH-MP), chlorzoxazone, 6-hydroxychlorzoxazone (6-OH-CH), dextrophan, magnesium chloride, NADPH, Folin & Ciocalteuís phenol reagent, potassium sodium tartrate tetrahy- drate, cooper sulfate, sucrose, TRIS, sodium phos- phate dibasic and potassium chloride were pur- chased from Sigma-Aldrich (St. Louis, MO, USA).

1091
Validation of LC/MS/MS method for assessment of the in vitro activity of...
lary temperature 370OC. The argon pressure in the collision cell was approximately 2 mTorr.
dards and quality control samples were prepared by spiking 40 µL of working solution of cocktail of analytes with blank microsomal medium (without NADPH). In the next step, to stop the reaction, the mixture of acetonitrile : acetone (1 : 1, v/v) with internal standard was added. After vortex mixing and cooling down on ice, samples were centrifuged at 15 000 ◊ g for 15 min, and supernatant was ana- lyzed using LC/MS/MS method.
Standard and quality control samples
Stock solutions of metabolites of model sub- strates and dextrophan used as an internal standard (IS) were prepared at concentration of 0.5 mg/mL in methanol. Working solutions for calibration stan- dards were prepared by diluting the stocks in water to the concentrations of 0.125, 0.3125, 0.625, 1.25,
3.25, 6.25, 9.375, 12.5 and 25 µg/mL of each analyte
in cocktail, and 0.125, 0.375, 1.125, 10, 22.5 µg/mL for quality control samples. Dextrophan was diluted in acetone : acetonitrile mixture (1 : 1, v/v) to a final concentration of 250 ng/mL.
Calibration samples were made at 10, 25, 50,
100, 250, 500, 750, 1000 and 2000 ng/mL for 4í-
hydroxymephenytoin and 4-hydroxytolbutamide; 50, 100, 250, 500, 750, 1000 and 2000 ng/mL for 1-
hydoxybufuralol and 25, 50, 100, 250, 500, 750, 1000 and 2000 ng/mL for the rest of analyzed metabolites.
Quality control (QC) samples were prepared at 30, 800 and 1800 ng/mL for 4í-hydroxymepheny- toin and 4-hydroxytolbutamide, and 90, 800 and 1800 ng/mL for other substances. Calibration stan-
Method validation
Validation of the method was carried out according to the FDA and EMA bioanalytical method validation guidelines (27, 28) with respect to the linearity, selectivity, precision, accuracy, recovery, limit of detection, limit of quantification, matrix effect and post-preparative stability.
Selectivity is the ability of the analytical method for separation and quantification of the ana- lyte in the presence of other compounds in the sam- ple. It was assessed by studying peak interference from six independent sources of microsomes.
Calibration curves were constructed using seven to nine calibration standards and fitted by a weighted least-squares linear regression. To assess linearity, deviations of the back calculated concen-
Table 1. Selected reaction monitoring parameters for quantified analytes.
Table 2. Linear ranges, determination coefficients, limits of detection and limits of quantification for selected compounds.
ACP - acetaminophen; 6-OH-CH - 6-hydroxychlorzoxazone; 4-OH-MP - 4í-hydroxymephenytoin; 1-OH-BF - 1-hydroxybufuralol; 4-OH- TB - 4-hydroxytolbutamide; 1-OH-MZ - 1í-hydroxymidazolam; 4-OH-MZ - 4í-hydroxymidazolam.

ACP
6-OH-CH
4-OH-MP
1-OH-BF
4-OH-TB
1-OH-MZ
4-OH-MZ
Liner range [ng/mL]
25-2000
25-2000
10-2000
50-2000
10-2000
25-2000
25-2000
Determination coefficient (r2)
0.9994
0.9995
0.9916
0.9908
0.9965
0.9974
0.9983
LOD [ng/mL]
0.1
1
0.05
0.01
0.01
0.025
0.01
LLOQ [ng/mL]
25 ± 2.8
25 ± 2.8
10 ± 1.3
50 ± 5.9
10 ± 0.5
25 ± 3.0
25 ± 3.1
Analyte
Q1 [m/z]
Q3 [m/z]
CE [V]
Acetaminophen
152
110
14
4-Hydroxytolbutamide
287
89
38
4í-Hydroxymephenytoin
235
150
19
1-Hydroxybufuralol
278
186
14
6ñHydroxychlorzoksazone
180
120
20
1í-Hydroksymidazolam
342
324
21
4í-Hydroxymidazolam
342
297
24
Dextrophan (IS)
258
157
38

1092
KAMIL KUå et al.
Figure 1. Fragmentation mass spectra of selected metabolites with purposed fragmentation pathways: 4-hydroxytolbutamide (A), aceta- minophen (B), 1-hydroxybufuralol (C), 6-hydroxychlorzoxasone (D), 1-hydroxymidazolam (E) and 4-hydroxymidazolam (F)
trations were set at ±15% of the nominal ones, except of the LLOQ, when a deviation of ±20% was permitted.
Inter- and intra-batch precision and accuracy of the assay were verified for three concentrations of QC samples. Six replicates of sample were analyzed on the same day to determine the intra-day precision and accuracy, and on each of two separate days to determine inter-day precision and accuracy. The mean values of precision should be within 15% of the actual value, except of LLOQ, where it should not deviate more than 20%. The accuracy deter- mined at each concentration level should not exceed 85-115%, except of LLOQ, where it should not exceed 80-120%.
The limit of detection (LOD) was defined as the concentration that generates a signal-to-noise ratio of 3, whereas the limit of quantification
(LLOQ) was determined as the lowest concentration of the analytes in the sample, which can be quanti- fied with precision lower than 20%, and an accura- cy between 80% and 120%.
The stability of samples in autosampler batch was assessed by reanalyzing extracted QC samples kept under the autosampler conditions (10OC) for 24
h. An acceptable stability was defined if the initial loss of concentration was below 15%.
Recovery and absolute matrix effect were assessed at three concentrations of QC samples by comparing the peak areas of 6 runs at each concen- tration for the neat solutions of analytes and the standard solutions spiked before and after protein precipitation. Recovery was calculated by dividing the mean peak area of standards spiked before extraction by mean peak area of standards spiked after extraction.

1093
Validation of LC/MS/MS method for assessment of the in vitro activity of...
The matrix effect was evaluated by comparing the mean peak area of the analyte spiked post extrac- tion to the mean peak of its neat solutions. The absolute matrix effect was defined as (1 - signal of post-extracted spiked sample/signal of pure solution)
◊ 100%. The relative matrix effect was also estimated according to the work of Matuszewski (29) by assess- ing the variability of standard line slopes expressed as coefficient of variation. The precision of the method should not exceed 4%, and then is considered to be reliable and free from the relative matrix effect.
shaking water bath (OLS 200, Grant Instruments, Cambridge, UK). The reaction was initiated by addi- tion of NADPH (cofactor of metabolic reaction). Following a 10 min incubation, the reaction was ter- minated with 500 µL of ice cold mixture of acetoni- trile : acetone (1 : 1, v/v) containing dextrophan as an internal standard (250 ng/mL). Samples were subsequently cooled on ice for 20 min to precipitate the proteins, and then centrifuged at approximately 15 000 ◊ g for 15 min at 4OC. The supernatant (200
µL) was transferred to vials, and a volume of 20 µL was injected onto the analytical column.
CYP-mediated activities in the presence of ketoconazole were expressed as a percentage of the corresponding control values. A sigmoidal shaped curve were fitted to the data and the IC50 values for ketoconazole were calculated by fitting the Hill equation to the data using non-linear regression of the plot of percent control activity versus concentra- tion of ketoconazole using Prism (GraphPad Software, Version 6). The enzymes activities in the presence of ketoconazole were compared with the control samples without addition of ketoconazole.
Inhibition study
To the total volume (0.5 mL) of incubation mixture of rat liver microsomes (1 mg/mL) added
0.1 M phosphate buffer (pH 7.4), 10 mM MgCl2, 1 mM NADPH and substrates cocktail in concentra- tions near their Km values: phenacetin - 7.5 µg/mL, tolbutamide - 2.5 µg/mL, bufuralol - 12.5 µg/mL, chlorzoxazone - 15 µg/mL and midazolam - 7.5
µg/mL. To the incubation mixture ketoconazole was added in the concentration range of 0.3 to 37.5 µM. Mixtures were preincubated for 5 min at +37OC in a
Figure 2. Extracted SRM chromatograms of simultaneous analysis of seven metabolites of model substrates and IS in calibration standard sample. Abbreviations see Table 2

1094
KAMIL KUå et al.
Figure 3. The activity of selected isoenzymes after incubation of model substrates of CYP450 with ketoconazole in rat liver microsomes. Data are shown as the mean ± SD, n = 3. 1-OH-MZ - 1'-hydroxymidazolam; 4-OH-MZ - 4'-hydroxymidazolam
RESULTS AND DISCUSSION
Differences between e.g., human and rat enzymes concern mostly individual genes, while not only families but also subfamilies are common. Major isoenzymes of CYP450 in human liver are: CYP 1A2, CYP 2B6, CYP 2C8, CYP 2C9, CYP 2D6 and
CYP 3A4/5, while in rat are: CYP 1A1, CYP 1A2, CYP 2B1, CYP 2C6, CYP 2C11, CYP 2D1/2, CYP
3A1/2 and CYP 4A1. Enzymes CYP 2C and CYP 3A appear to be the most abundant cytochromes P450 in rat and human liver, respectively (31).
Most of the described in the literature experi- ments deal with human microsomes, while there are a lot of work performed on rat microsomes, because rats are the most commonly used model in preclini- cal studies (6, 32-34). Zhang et al. (6) studied phenacetin, tolbutamide, mephenytoin, dextro- methorphan and midazolam metabolism incubated
Method development
Metabolism studies of new
compounds that
occur with the participation of enzymes require sim- ple, quick, sensitive and reproducible analytical methods. The aim of this work was to develop and validate a simple and reliable LC/MS/MS method, which can be used to determine simultaneously seven metabolites of specific substrates of selected CYP450 isoenzymes in rat liver microsomes.
Animal models are widely used to predict kinetics and toxicity of drugs in human, and cytochrome P450 enzymes are mainly involved in the metabolism of drugs both in human and rodents. The three families: CYP1, CYP2 and CYP3 play a significant role in drug biotransformation (30).

1095
Validation of LC/MS/MS method for assessment of the in vitro activity of...
individually and in cocktail form with rat micro- somes and no significant differences were observed. Phenacetin and tolbutamide metabolism in rat liver microsomes were studied by Jurica et al. (32). Coctail assay was used to study herb-drug interac- tion in rat microsomes by Xia et al. (33). Recently, comparison between human and rat microsomal CIME cocktail assay was verified by Videau et al. (34).
In presented work, the model substrates of selected isoenzymes of CYP450 were added in the form of cocktail to rat microsomes. Substrates were selected on the basis of a number of literature reports describing the metabolism of drugs in rats (30, 31, 35ñ37) and the EMA guideline describes probe substances, which can be used for evaluation of cytochrome P450 activity in human (38). The fol- lowing model substrates were selected: phenacetin
for CYP 1A2, tolbutamide for CYP 2C9, S- mephenytoin for CYP 2C19, bufuralol for CYP 2D6, chlorzoxazone for CYP 2E1 and midazolam for CYP 3A4. Fragmentation mass spectra of select- ed metabolites and proposed fragmentation path- ways are shown in Figure 1.
The analytes were at first detected in full scan mass spectrum and then fragmented to product ions to ascertain their pseudo-molecular ions and to select product ions used in SRM mode. Except of chlorzoxazone and its metabolite, which responded better in negative ionization mode, all other sub- stances were quantified in positive ion mode.
A simple protein precipitation method was used for sample purification and different precipita- tion mixtures were verified. The results showed that the most efficient way was precipitation using a mixture of acetonitile and acetone (1 : 1, v/v).
Table 3. Intra- and inter-day precision and accuracy of the method.
Abbreviations see Table 2.
Table 4. Extraction recovery and absolute matrix effect of the method.
Abbreviations see Table 2.

ACP
6-OH-CH
4-OH-MP
1-OH-BF
4-OH-TB
1-OH-MZ
4-OH-MZ

Recovery [%]

QC low
85.4
92.0
68.4
83.5
74.1
79.4
82.1
QC medium
70.5
91.1
68.3
86.4
74.3
84.8
83.0
QC high
75.3
94.8
71.4
89.0
78.9
91.1
90.6

Mean ± SD
77.1 ± 7.6
92.6 ± 1.9
69.4 ± 1.8
86.3 ± 2.8
75.8 ± 2.7
85.1 ± 5.9
85.2 ± 4.7
Absolute matrix effect [%]
QC low
65.0
16.5
8.9
-23.7
20.2
12.3
-14.8%
QC medium
61.9
17.9
14.0
5.2
22.7
-5.4
0.2
QC high
60.1
10.2
10.1
4.0
20.8
-6.8
2.0

Mean ± SD
62.3 ± 2.5
14.9 ± 4.1
11.0 ± 2.7
-4.8 ± 16.4
21.2 ± 1.3
0.03 ± 10.6
-4.2 ± 1.2

ACP
6-OH-CH
4-OH-MP
1-OH-BF
4-OH-TB
1-OH-MZ
4-OH-MZ

Intra-day precision (RSD%)
LLOQ
11.1
10.4
13.0
19.5
4.5
12.6
13.0
QC low
3.2
8.8
4.0
7.5
4.4
7.1
5.1
QC medium
11.0
7.6
3.4
2.7
2.5
1.7
1.0
QC high
8.4
6.5
3.8
2.1
3.0
3.5
2.7

Inter-day precision (RSD%)
LLOQ
13.9
12.0
12.0
7.5
3.2
5.4
6.6
QC low
5.6
7.5
7.4
14.9
6.1
6.3
4.1
QC medium
11.7
5.9
11.0
4.0
7.5
4.7
3.0
QC high
10.8
5.4
7.5
3.5
3.5
4.0
2.4

Accuracy (%)

LLOQ
118.7
106.3
103.1
83.7
107.0
97.2
96.5
QC low
102.9
95.5
106.0
113.6
99.8
98.2
109.2
QC medium
103.3
98.1
104.6
106.7
99.8
103.8
102.7
QC high
107.4
93.0
99.7
99.6
97.6
105.4
98.2

1096
KAMIL KUå et al.
Method validation
Selectivity and linearity
No interference was found at the retention time of the analytes and the IS during analysis of blank microsomes from six independent sources (signal- to-noise ratio < 9). Extracted SRM chromatograms of seven metabolites of model substrates and IS in calibration standard sample are shown in Figure 2.
The assay was linear over the concentration range of 10-2000 ng/mL for 4í-hydroxymepheny- toin and 1-hydroxytolbutamide, 50-2000 ng/mL for 1-hydroxybufuralol, and 25-2000 ng/mL for the rest of the analyzed metabolites. Linear ranges, determination coefficients, LOD and LLOQ values are shown in Table 2. The best linear fit of the cal- ibration curve for each analyte was obtained using a weighting factor 1/x. Precision and accuracy of the LLOQ met the requirements both of FDA and
lytes. The extraction recovery was noticed to be con- sistent for each analyte over the entire QC concen- tration range, what can indicate that extraction effi- ciency of the method is reliable over the studied concentration.
The mean values of absolute matrix effect ranged from 0.03 to 62.3%, and the higher matrix effect was observed for acetaminophen (62.3%) but even so, the calibration curve covered the expected concentration range. To confirm the reliability of this method, the relative matrix effect was studied. Coefficients of variation of calibration curve slopes for compounds studied were less than 1.8% (Table 5), and according to Matuszewski indications (29), the method was free from the relative matrix effect liability.
Stability study
Analyzed aliquots were stable in autosampler batch in 10OC for at least 24 h. The mean values of the calculated concentrations were within 94-113% and therefore acceptable (Table 6). Consequently, extracts from sample incubation may be placed in autosampler rack for up to 24 h.
EMA guidelines, RSD respectively.
<
20%
and
80-120%,
Precission and accuracy
For all quantified metabolites, the intra-day coefficients of variation were less than 11% for low, medium and high QC concentrations, and less than 20% for LLOQ. The inter-day data were also repro- ducible with CV less than 12% for prepared QC samples, and less than 14% for LLOQ. The study shows good accuracy: 93-114% for low, medium and high QC concentrations, and 83-119% for LLOQ level. Precision and accuracy data for each of the analytes are summarized in Table 3.
Application of the method
The method was used to assess the strength of inhibition of ketoconazole on selected CYP450 isoenzymes. The incubation curves are shown in Figure 3 and determined IC50 values for ketocona- zole are listed in Table 7. The results showed that ketoconazole had the most potent inhibitory effect on CYP 3A4 (IC50 = 1 µM), but also suppressed activity of CYP 1A2 (IC50 = 4.4 µM), CYP 2E1 (IC50
= 8.1 µM) and CYP 2C9 (IC50 = 23.5 µM), whereas
the activity of CYP 2D6 was not changed.
Many researchers found that antifungal azoles can inhibit the activity of several cytochrome P450
Recovery and matrix effect
Table 4 shows the results of recovery
and
absolute matrix effect for all selected analytes. A simple precipitation method used in this study yield- ed a mean recovery greater than 69% for all ana-
Table 5. Individual and mean values of standard line slopes for analytes studied.
Abbreviations see Table 2.
Slope of the calibration curve

ACP
6-OH-CH
4-OH-MP
1-OH-BF
4-OH-TB
1-OH-MZ
4-OH-MZ
0.00048
7.8309E-06
0.000155
0.00275
0.00043
0.0046
0.00147
0.00048
7.8867 E-06
0.00015
0.00286
0.00044
0.0049
0.00148
0.0005
8.2371 E-06
0.000153
0.00291
0.00043
0.00478
0.00152
0.00049
8.3464 E-06
0.00015
0.00288
0.00043
0.00468
0.00153
0.00051
8.4954 E-06
0.000149
0.00286
0.00047
0.00461
0.00147
0.00052
8.2371 E-06
0.000139
0.00288
0.00042
0.00491
0.00148
Mean
0.000496
8.17227 E-06
0.000149
0.00285
0.000436
0.004747
0.001492
CV [%]
3.3
3.2
3.7
1.9
4.0
2.9
1.8

1097
Validation of LC/MS/MS method for assessment of the in vitro activity of...
Table 6. Autosampler stability data for analytes studied.
Abbreviations see Table 2.
Table 7. The influence of ketoconazole on the inhibition of selected isoenzymes.
isoenzymes. In the work (39), Paine et al. demon- strated the inhibition of intestinal CYP 1A2 by keto- conazole. Suppression of tolbutamide hydroxylase (reaction catalyzed by CYP 2C9) by azoles in in vitro experiments was studied by Back et al. (40). They found that ketoconazole inhibited CYP 2C9 in human microsomes with IC50 = 16.5 µM. The effect of non-selective cytochrome P450 inhibitors such as 1-aminobenzotriazole, SKF-525A and ketoconazole on human cytochrome P450 isoenzymes was studied by Emoto et al. (41). They found that ketoconazole exerts weak inhibitory potential on CYP 1A2 activ- ity (Ki = 120 µM), CYP 2C competitive inhibition (Ki = 7.9 µM for CYP 2C9 and Ki = 6.9 µM for CYP 2C19), mixted inhibition on CYP 2D6 activity (Ki = 12 µM) and CYP 2E1 activity (Ki = 41 µM) and competitive inhibition on CYP 3A4 (Ki = 0.02 µM). The influence of ketoconazole on amiodarone metabolism both in human and rat microsomes was studied by Elsherbiny et al. (42). Interestingly, in rat liver microsomes ketoconazole exerted higher inhibitory potential on CYP 2D enzymes than on CYP 3A family. The human isoenzymes most inhib- ited by ketoconazole were CYP 1A1, CYP 3A4, CYP 2D6 and CYP 1A2, respectively. The working groups of Kobayashi et al. (43) and Eagling et al.
(44) studied the effect of different inhibitors on cytochrome P450 activity in rats and showed that ketoconazole is not specific CYP 3A inhibitor, but it
can suppress the activity of other isoenzymes. Only in one in vivo work, it was shown that ketoconazole had no effect on cytochrome P450 activity in rats (45).
Despite the use in presented studies different substrates and different substrate concentrations it can be seen that ketoconazole has the strongest inhibitory potential on CYP3A4 but this drug is also strong inhibitor of CYP1A2 and CYP2E1 and mod- erate inhibitor of CYP2C9. Presented data are gen- erally in good agreement with earlier cited studies and confirmed that ketoconazole should not be con- sider as a selective CYP 3A4 inhibitor in rat liver microsomes.
CONCLUSIONS
A sensitive and an accurate LC/MS/MS method has been developed for the simultaneous evaluation of the activities of six cytochrome P450 izoenzymes in rat liver microsomes. The method was successfully validated following FDA and EMA guidelines that ensures the acceptability of its performance and the reliability of analytical results. The applied screening cocktail was verified with ketoconazole, a known CYP450 enzyme inhibitor. The resulting method can be used as a first and rapid screening of new compounds, drug candidates using rat liver microsomes.
Isoenzyme

Substrate

Metabolite

Ketoconazole IC50 [µM]
CYP 1A2
Phenacetin
Acetaminophen
4.40 ± 0.29
CYP 2C9
Tolbutamide
4-Hydroxytolbutamide
23.5 ± 0.05
CYP 2D6
Bufuralol
1-Hydroxybufuralol
ó-
CYP 2E1
Chlorzoxazone
6-Hydroxychlorzoxazone
8.10 ± 0.041

CYP 3A4
Midazolam
1í-Hydroxymidazolam
1.51 ± 0.024
Midazolam
4í-Hydroxymidazolam
1.02 ± 0.027
Stability of analytes [%]

ACP
6-OH-CH
4-OH-MP
1-OH-BF
4-OH-TB
1-OH-MZ
4-OH-MZ
QC low
103.2
95.7
101.3
100.4
109.5
102.2
94.9
QC medium
111.6
106.4
105.7
106.3
112.6
105.0
94.4
QC high
111.2
102.7
104.7
103.6
104.4
105.2
95.7

1098
KAMIL KUå et al.
Acknowledgment
16.
FDA: Guidance for industry: Drug me- tabolism/drug interaction studies in the drug development process: Studies in vitro, Food and Drug Administration. Rockville, MD 1997.
Yang B., Liu W., Chen K., Wang Z., Wang C.: Drug Metab. Dispos. 42, 1737 (2014).
Wang B., Yang S., Hu J., Li Y.: J. Ethnophar- macol. 155, 1473 (2014).
Chen P., Guo M., Wygle D., Edwards P., Falck
J.R. et al.: Am. J. Pathol. 166, 615 (2005).
Han Y.-L., Li D., Ren B., Jing G.-P., Meng X.-
L. et al.: J. Ethnopharmacol., 139, 104 (2012). Hu X., Huang W., Yang Y.: Eur. J. Drug Metab. Pharmacokinet. DOI: 10.1007/s13318-014- 0219-4 (2014).
De Bock L., Vande Casteele S.R.F., Mulliez S.M.N., Boussery K., Van Bocxlaer J.F.P.: J. Chromatogr. A 1218, 793 (2011).
Larsen K., Najle R., Lifschitz A., MatÈ M.L., Lanusse C., Virkel G.L.: Int. J. Toxicol. 33, 307 (2014).
Zhao X.J., Ishizaki T.: J. Pharmacol. Exp. Ther. 283, 1168 (1997).
Martignoni M., Groothuis G.M.M., de Kanter R.: Expert Opin. Drug Metab. Toxicol. 2, 875 (2006).
Lowry O.H., Rosebrough N.J., Farr A.L., Randall R.J.: J. Biol. Chem. 193, 265 (1951). FDA: Guidance for Industry, Bioanalytical Methods Validation, Rockville, MD 2001.
EMA: Guideline on bioanalytical method vali- dation, London 2012.
Matuszewski B.K.: J. Chromatogr. B. Analyt. Technol. Biomed. Life Sci. 830, 293 (2006).
Bogaards J., Bertrand M., Jackson P.CE., Oudshoorn M., Weaver R. et al.: Xenobiotica
30, 1131 (2000).
Kobayashi K., Urashima K., Shimada N., Chiba K.: Biochem. Pharmacol. 63, 889 (2002).
Jurica J., Kyr M., McCaskey Hadasova E., To- mandl J.: Neuro Endocrinol. Lett. 30, 92 (2009). Xia C., Sun J., Wang G., Shang L., Zhang X. et al.: Planta Med. 76, 245 (2010).
Videau O., Pitarque S., Troncale S., Hery P., ThÈvenot E. et al.: Xenobiotica 42, 349 (2012). Chovan J.P., Ring S.C., Yu E., Baldino J.P.: Xenobiotica 37, 459 (2007).
OíBrien P.J., Chan K., Silber P.M.: Chem. Biol. Interact. 150, 97 (2004).
Hiroi T., Chow T., Imaoka S., Funae Y.: Drug Metab. Dispos. 30, 970 (2002).
EMA: Guideline on the Investigation of Drug Interactions, London 2013.
This work was
supported by the European
Union from the resources of the European Regional Development Fund under the Innovative Economy
Programme (grant coordinated POIG.01.01.02-00-069/09).
by
JCET-UJ,
No.
17.
18.
Declaration of interest
19.
The authors report no conflict and declaration of interest.
20.
21.
REFERENCES
1.
Bertz R. J., Granneman G. R.: Clin. Pharmaco- kinet. 32, 210 (1997).
Nelson D.R., Koymans L., Kamataki T., Stege- man J.J., Feyereisen R. et al.: Pharmacogenetics 6, 1 (1996).
Zhou S., Yung Chan S., Cher Goh B., Chan E., Duan W. et al.: Clin. Pharmacokinet. 44, 279 (2005).
Kostiainen R., Kotiaho T., Kuuranne T., Auriola S.: J. Mass Spectrom. 38, 357 (2003). Ayrton J., Plumb R., Leavens W., Mallett D., Dickins M., Dear G.: Rapid Commun. Mass Spectrom. 12, 217 (1998).
Zhang S., Song N., Li Q., Fan H., Liu C.: J. Chromatogr. B. Analyt. Technol. Biomed. Life Sci. 871, 78 (2008).
Lee K. S., Kim S. K.: J. Appl. Toxicol. 33, 100 (2013).
Prot J.-M., Videau O., Brochot C., Legallais C., BÈnech H., Leclerc E.: Int. J. Pharm. 408, 67 (2011).
Zhang W., Han F., Guo P., Zhao H., Lin Z. J. et al.: J. Chromatogr. B. Analyt. Technol. Biomed. Life Sci. 878, 1169 (2010).
Testino S.A.J., Patonay G.: J. Pharm. Biomed. Anal. 30, 1459 (2003).
Liu Y., Jiao J., Zhang C., Lou J.: Biol. Pharm. Bull. 32, 717 (2009).
Dierks E.A., Stams K.R., Lim H.-K., Cornelius G., Zhang H., Ball S.E.: Drug Metab. Dispos. 29, 23 (2001).
Stewart N.A. Buch, S.C., Conrads T.P., Branch R.A.: Analyst 136, 605 (2011).
Turpault S., Brian W., Van Horn R., Santoni A., Poi- tiers F. et al.: Br. J. Clin. Pharmacol. 68, 928 (2009). De Bock L., Boussery K., Colin P., De Smet J., Tíjollyn H., Van Bocxlaer J.: Talanta 89, 209 (2012).
22.
2.
23.
3.
24.
4.
25.
5.
26.
6.
27.
28.
7.
29.
8.
30.
9.
31.
10.
32.
11.
33.
12.
34.
35.
13.
36.
14.
37.
15.
38.

1099
Validation of LC/MS/MS method for assessment of the in vitro activity of...
1. Paine M.F., Schmiedlin-ren P., Watkins P.B.: Drug Metab. Dispos. 27, 360 (1999).
2. Back D.J., Tjia J.F., Karbwang J., Colbert J.: Br. J. Clin. Pharmacol. 26, 23 (1988).
3. Emoto C., Murase S., Sawada Y., Jones B.C., Iwasaki K.: Drug Metab. Pharmacokinet. 18, 287 (2003).
1. Elsherbiny M.E., El-Kadi A.O.S., Brocks D.R.: J. Pharm. Pharm. Sci. 11, 147 (2008).
2. Kobayashi K., Urashima K., Shimada N., Chiba K.: Drug Metab. Dispos. 31, 833 (2003).
3. Eagling V.A., Tjia J.F., Back D.J.: Br. J. Clin. Pharmacol. 45, 107 (1998).
4. Uchida S., Tanaka S., Namiki N.: Biopharm. Drug Dispos. 35, 228 (2014).
Received: 16. 09. 2014