Acta Poloniae Pharmaceutica ñ Drug Research, Vol. 68 No. 4 pp. 481ñ492, 2011
ISSN 0001-6837
Polish Pharmaceutical Society
APPLICABILITY OF THE SILVER AMALGAM ELECTRODE IN VOLTAMMETRIC DETERMINATION OF ZINC
AND COPPER IN GASTRIC JUICE AND GASTRIC MUCOSA OF RATS
W£ODZIMIERZ OPOKA1*, BOGUS£AW BAå2, WITOLD RECZY—SKI2, MA£GORZATA P£ONKA3, DANUTA DROZDOWICZ4, ZBIGNIEW åLIWOWSKI4 and TOMASZ BRZOZOWSKI4
1Department of Inorganic Chemistry, Faculty of Pharmacy, Jagiellonian University Medical College, 9 Medyczna St., 30-688 KrakÛw, Poland
2Department of Analytical Chemistry, Faculty of Material Science and Ceramics, AGH University of Science and Technology, 30 Mickiewicza Av., 30-059 KrakÛw, Poland
3Department of Physiotherapy, University School of Physical Education, 78 J. Paul II Av., 31-571 KrakÛw, Poland
4Department of Physiology, Jagiellonian University Medical College, 16 GrzegÛrzecka St., 31-531 KrakÛw, Poland
Abstract: The aim of the work was to compare two analytical methods of trace analysis in respect to their appli- cability in heavy metals determination in biological samples. Atomic absorption spectrometry (AAS) may be considered as the method of choice in such analyses due to its accuracy, precision and low detection limit. On the other hand, voltammetric methods seem to be as useful, but rarely applied. Having in mind that there is no universal analytical method, we have compared two AAS and voltammetric methods as the tools for Zn and Cu determination in the samples collected from rat gastric juice and gastric mucosa. Construction of the renewable silver amalgam film electrode (Hg(Ag)FE) for stripping voltammetry was described. Detailed optimization of measurements procedure and sample preparation for differential pulse anodic stripping voltammetry (DP ASV) and AAS were also performed and presented. The obtained results of quantitative analysis of the chosen param- eters by means of both methods are discussed.
Keywords: zinc, copper, gastric mucosa, gastric juice, amalgam film electrodes, stripping voltammetry, atom- ic absorption spectrometry
Zinc is important for body metabolism, growth, gastric acid secretion, reduction and tissue healing, it activates about 300 enzymes (1ñ3) but the
mucosa was reported to delay ulcer healing but the mechanism of this action of zinc preparation has not been fully clarified. The studies in animals and humans confirmed that zinc possesses, so called, ìcytoprotectiveî properties (13).
Similarly, the anti-inflammatory action of cop- per complexes seems to contribute to their anti-ulcer effect. Part of the protective and anti-ulcer mecha- nism could be attributed to absorption and transport of copper that seems to be involved in an activation of the several copper-dependent enzymes (14ñ16). It has been confirmed that copper nicotinate exhibits a gastroprotective action against formation of gastric lesions (17).
precise mechanism of	containing compounds
Zn2+
affecting mucosal integrity, gastroprotection and ulcer healing remains unclear. It is confirmed that zinc affects healing of damage of various tissues (4). In many experiments it was shown that gastric dam- age induced in laboratory animals by various means such as ethanol, steroidal agents, strong acids (5ñ8) and stress (9, 10) can be treated using zinc-contain- ing remedies. Therefore, it was concluded that zinc exhibits protective properties and could be used in the treatment of inflammation in various form of gastrointestinal injury under different experimental conditions (11, 12). Zinc deficit in serum and gastric
Trace elements ions such as
Zn2+	Cu2+
and
could be successfully measured in the gastric juice
* Corresponding author: e-mail address: mfopoka@cyf-kr.edu.pl
481

482
W£ODZIMIERZ OPOKA et al.
and	gastric	mucosa	during	ulcer	healing.
tiometric-stripping analysis (PSA) was possible in whole blood samples (32), a time-consuming sam- ple digestion was required (ca. 3.5 h) greatly reduc- ing the attraction of this method as an analytical pro- tocol. It is clearly evident from the literature that a rapid and sensitive electrochemistry protocol is required for the determination of metal species in clinical samples.
The sensitive determination of zinc and copper traces is quite simple by ASV equipped with HMDE. However, the toxicity of mercury limits the usage of the mercury electrodes in the analytical practice. The essential group of electrodes used nowadays in voltammetry measurements are film electrodes. There are two types of such electrodes: mercury film electrodes (MFEs) (33) and bismuth film electrodes (BiFEs) (34). Film electrodes are characterized by high sensitivity and repeatability. The MFEs, due to their large surface-to-volume ratio, give sharper stripping peaks and therefore enhanced sensitivity (35). One of the alternatives to mercury are liquid and solid amalgams (36). The problem of limiting the amount of mercury or its soluble salts used to generate MFEs needed for the analytical procedure can be solved with the help of a renewable silver amalgam film electrode (Hg(Ag)FE). The principle of working and first pro- posal of a construction of the (Hg(Ag)FE) was described in (37). In this work, Hg(Ag)FE was pro- posed as a working electrode for voltammetric determination of Zn and Cu in gastric juice and gas- tric mucosa of rats.
Compounds chelating of	and
Zn2+	Cu2+
can exert
beneficial influence on the ulcer healing possibly via
Zn2+ and Cu2+ mediated increase in gastric microcir- culation but their antisecretory activity and the effect on gastrin release, which may exert trophic action on gastric mucosa with ulcer, remain unknown.
Important methods of analyzing heavy metals are mainly based on the use of graphite furnace atom- ic absorption spectrometry (GF-AAS) (18), the direct flame atomic absorption spectrometry (F-AAS) (19), inductively coupled plasma mass spectrometry (ICP- MS) (20) or inductively coupled plasma optical emis- sion spectroscopy (ICP-OES), neutron activation analysis (NAA) (21) and X-ray fluorescence analysis (22). However, these spectrometric methods are expensive, and not suitable for the in situ measure- ment and reliable measure of total metal ion concen- trations, and provide no information regarding the concentration of metals in a sample that might inter- act with a biological molecules or species.
An alternative to these spectroscopy tech- niques is electro-analytical technique ñ voltamme- try. Voltammetry is certainly a valid analytical tech- nique, very simple and suitable for metals determi- nation in multicomponent complex matrices. Voltammetric methods are widely applied in bio- medical and pharmacological analysis (23, 24) as well as in analysis of metals (25ñ27). Differential pulse voltammetry (DPV) is considered a conven- ient method because of the wide range of linearity, excellent reproducibility, low experimental cost and the attainment of low detection limit. Stripping voltammetry (SV) comprises a variety of electro- chemical approaches, having a step of preconcentra- tion onto the electrode working prior to the voltam- metric measurement. The major advantage of SV compared with direct voltammetric measurement is the preconcentration factor. In trace analysis of heavy metal ions, anodic stripping voltammetric (ASV) is the most popular SV technique (28, 29). Moreover, such a technique may be a good alterna- tive to spectroscopy, which, in the case of determi- nation of metal species, metals in complex matrices, needs expensive equipment. However, organic con- stituents in the sample matrix often passivate elec- trodes and prevent direct determination (30), such that either the use of mercury is required and/or a sample pre-treatment may need to be performed. For example, Tripathi et al. (31) have quantified zinc in infants using ASV with the use of a hanging mercu- ry drop electrode (HMDE) and with complementary AAS. While it has been shown that ASV and poten-
EXPERIMENTAL
Measuring apparatus and software
An Electrochemical Analyzer M161 (MTM- ANKO, Poland) was used in this study. The classi- cal three-electrode quartz cell of 10 mL volume, consisting of a homemade renewable silver amal- gam film electrode (Hg(Ag)FE), with a surface area of 1ñ12 mm2, as the working electrode, a double junction reference electrode Ag/AgCl/3M KCl and a platinum wire as an auxiliary electrode. All solu- tions used for analyses were purged with argon of 99.995% purity. Magnetic Teflon-coated bar was used for stirring (approx. 500 rpm) during the accu- mulation period. The MTM-ANKO EAGRAPH software enabled electrochemical measurements, data acquisition and advanced processing of the results (38ñ40). For spectroscopic determination of Zn and Cu, the Perkin Elmer spectrometer Model 3110 (USA) was used. Liquid samples were digest- ed using a UV-digester (Mineral, Poland). Solid

483
Applicability of the silver amalgam electrode in voltammetric determination of zinc...
Figure 1. (A) The principle of mechanical refreshing of the liquid amalgam film silver based electrode. The Hg(Ag)FE used in our exper- iments: (B) configuration before use, (C) configuration ready for measurement. (1) micrometric screw, (2) piston pin with Ag cylindrical electrode at the end, (3) O-ring, (4) Ag foil (0.05 mm), (5) liquid silver amalgam (10 ÏL), (6) electric contact pin
materials were digested in a microwave digestion system Anton Paar Multiwave 3000 (Switzerland). All experiments were carried out at room tempera- ture.
ples and supporting electrolyte. The standard solu- tion of zinc(II) and copper(II) were at a concentra- tion of 1000 mg/L (Merck). Solutions with lower zinc and copper concentrations were made weekly by appropriate dilution of the stock solution. Fumed silica of the specific surface area 255 m2/g (Sigma- Aldrich) was activated by roasting for 30 min at 950OC. Prior to use, all glassware and, if necessary, also the electrode body, were cleaned by immersion in a 6 M nitric acid, followed by copious rinsing in distilled water to avoid contamination.
Chemicals and glassware
All used reagents were of analytical grade. All solutions and the sample preparation were made using 4-fold distilled water (two last stages from quartz). HNO3 65%, H2O2 30% and KNO3 (Merck, SuprapurÆ) were used for the preparation of sam-

484
W£ODZIMIERZ OPOKA et al.
Construction of the renewable silver amalgam film electrode, Hg(Ag)FE
The preparation and construction of the renew- able (Hg(Ag)FE) were described in details (41, 42). The construction of the applied electrode that allows the liquid amalgam film to be refreshed before each measurement, a procedure essential for its perform- ance, is given in Figure1. Figure 1A presents how the film electrode is refreshed, whereas Figures 1B and 1C show the construction of the electrode used throughout the experiments.
The micrometric screw /1/ allows the precise setting of the part of the electrode surface /2/ that is in contact with the sample solution. Rotating the screw by a full 360O changes the area of electrode surface by 1.8 mm2. The O-ring seals /3/ made from silicon rubber, separated by a 0.05 mm thick silver foil /4/ used to remove any possible mechanical con- tamination, overlap the amalgam layer /5/ and remove the excess liquid amalgam. The procedure of refreshing the outer amalgam film involves two steps: (a) pulling up the silver electrode base inside the electrode body, through the amalgam chamber (Fig. 1B) and then (b) pushing it back outside the electrode body (Fig. 1C). The total volume of the liquid silver amalgam used to fill up the chamber does not exceed 10 L. The most recent version of amalgam film electrode applied for measurements is shown in Fig. 1B, which presents the electrode con- figuration before the electrochemical experiment. The electrode is in operation when the silver cylin- der is moved down and immersed in the solution. Figure 1C shows the Hg(Ag)FE ready for measure- ments.
cell while maintaining an argon atmosphere over the solution and the DP voltammograms were recorded. The total analytical procedure consists in carrying out two steps in succession. In the first step, zinc is determined in the range -1.10 to -0.70 V (Eacc = -
1.10 V, tacc = 30 s). In the second step, copper con-
centration is measured in the range -0.25 to 0.15 V
(Eacc = -0.25 V, tacc = 30 s). The quantitative deter- minations of zinc and copper ions were performed using the standard additions method (three concen- trations). Three curves were recorded and averaged for each concentration.
The peak current value, relevant to each addi- tion, is plotted on the y-axis, while the x-axis is graduated in terms of the amount of analyte added. The regression line is calculated and extrapolated back to the point on the x-axis at which y = 0. It is clear that this negative intercept on the x-axis corre- sponds to the amount of the analyte in the test sam- ple. It is important to highlight that such a method shows a particular advantage: the regression analyt- ical calibration function does not present matrix effects. Because of very complicated background in these experiments, baseline correction algorithm should be applied to obtain high quality calibration model (43, 44). All samples were measured under the same conditions.
Sample preparation
Experiments was performed on 15 male Wistar rats weighing 200ñ250 g. The animals have been divided into two groups. Group A (n = 6) consisted of animals with operationally implanted metal gas- tric fistula measuring gastric juice (GJ) secretion and group B (n = 9) with gastric mucosa (GM) of animals without fistula. The procedure of implanti- ng the metal cannula into the gastric wall had been conducted 14 days before the start of study. The ani- mals have been fed with standard granulated mass including indispensable nourishment components, they have had free access to water and they have stayed in a well-lit room with access to fresh air. The animals were housed in a temperature of 22 ± 1OC.
The content of zinc and copper ions in GJ has been studied in rats of group A and in GM in group
B. Before testing, all animals were food fasted for 24 h with access to water ad libitum. The animals were sacrificed and their stomachs were removed rapidly, opened along the lesser curvature, and washed in physiological saline and gastric tissue samples were taken for determination of zinc and copper concentrations. Tissue samples from each animal have been put in Eppendorf tubes, frozen and kept at ñ80OC until analysis.
Standard procedure of measurements
The stripping was performed in the differential pulse (DP) mode. Before measurements, the voltam- metric cell was conditioned in 0.1 M nitric acid, rinsed with distilled water and shortly conditioned in supporting electrolyte (0.1 M KNO3). Next, 5 mL of
0.05 M KNO3 was added in the electrochemical cell as a blank and the solution was purged with argon for at least 5ñ7 min. A potential of -0.10 V was applied to conditioning the electrode. The accumu- lation step (time and potential) was carried out from the stirred solution for a period of tacc = 30 s at an Eacc
= -1.10 V. After a rest period of 5 s, DP voltammo-
gram was recorded in the anodic direction from -
1.10 to 0.15 V with a potential scan rate of 25 mV/s and pulse amplitude of -30 mV. The voltammogram for the blank solution demonstrates electrochemical cell and supporting electrolyte purity. Then, 0.05ñ0.20 mL of sample solution was added to the

485
Applicability of the silver amalgam electrode in voltammetric determination of zinc...
All experimental procedures were approved by the Experimental Animal Research Committee of the Jagiellonian University Medical College.
tions of Zn2+ and Cu2+ were performed. For volumes of gastric juice higher than 0.1 mL, the samples were initially digested in an UV system. For the samples containing high concentrations of organic compounds, much higher amount of silica was used ñ up to 50 mg SiO2 for 1 mL of the analyzed solu- tion. As the addition of such high amount of silica directly to the electrochemical cell is not possible, the second procedure was proposed. To 20 mL of contaminated sample 0.25ñ1.0 g of fumed silica was added. The suspension was stirred and shaken for 15 min. To separate the fumed silica, the sample was centrifuged for 5ñ8 min. at 3500 rpm. An appropri- ate volume of the supernatant, free of organic inter- fering substances, was then added to the cell con- taining the supporting electrolyte. The Zn2+ and Cu2+ determination were performed in standard condi- tions.
Gastric mucosa (GM)
For DP ASV zinc and copper determination in GM, about 50 mg of dried sample was weighed and transferred into a high pressure PTFE container and treated with 4 mL of HNO3 and 2 mL of H2O2 (30%). The container was then placed into a microwave oven. Digestion of the samples was carried out with the following program: 20 min under microwave irradiation 45 min cooling time, 5 min waiting time. Digested samples were placed on a heated plate in order to evaporate the excess of reagents. The sam- ple solutions were cooled to room temperature and transferred quantitatively into volumetric flasks (5 mL) and filled up to the mark with 4-fold distilled water.
RESULTS AND DISCUSSION
Gastric juice (GJ)
GJ samples were acidified with nitric acid immediately after collection by addition of 2 L HNO3 (conc.) to each 20 L of sample. The pre- treatment, intended to destroy organic compounds (complexing agents and organic surfactants), fol- lowed the wet ashing method used in the ASV deter- mination of Zn2+ and Cu2+ in GJ. Then, samples were transferred directly into a miniature quartz tube and were digested by UV irradiation for 2 h. The quartz tubes were let to cool at room temperature. Aliquots of 0.2ñ0.5 mL of this solution were introduced into the electrochemical cell containing the supporting electrolyte.
Characteristic features of the Hg(Ag) film elec- trode
The Hg(Ag)FE maintain its perfect repeatabil- ity and reproducibility for 2ñ3 thousand cycles under condition that: a) for the Hg(Ag)FE prepara- tion, silver of fibrous texture is used, b) for the Hg(Ag)FE regeneration, silver liquid amalgam is used (5). During the performed experiments, the use of silver wire (Goodfellow Science Park, England) ensured the Hg(Ag)FE stability for 6 months of electrode tests. Silver liquid amalgam does not dis- turb the Hg(Ag)FE surface even though it is exposed to constant contact for several weeks, refreshed in it film does not change its properties for many min- utes, and the hydrogen overpotential is comparable to the mercury electrodes. The content of silver in the liquid amalgam was determined using AAS method. When electrode is not renewed before measurement, peak current drops for successive scan about 5%.
The application of fumed silica
Interfering organic substances
might be
removed from the analyzed liquid samples not only be means of digestion, but also by their adsorption on fumed silica. First Kowalski et al. (45) have pro- posed an effective method of voltammetric meas- urements, in which electrolytic process was not dis- turbed by addition of fumed silica directly to the electrochemical cell. The usefulness of that proce- dure in elimination of certain interferences in deter- mination of traces of heavy metals was observed earlier and described in our works (46, 47). In this work, organic interfering substances were removed from the samples in two different ways. The first method consisted in preparation of either 5 mL of the supporting electrolyte spiked by the 0.05ñ0.20 mL of GJ. Then, 25 mg of SiO2 was added directly to the cell with the sample. The solution was stirred and deaerated with pure argon. Next, the determina-
Influence of DP ASV parameters on technique on zinc and copper peaks
In order to adapt the DP ASV method to ppb
concentrations of	and Cu2+, three parameters were optimized: step potential (Es), pulse amplitude (E) and pulse time (tp = tw (waiting time) + ts (cur- rent sampling time)). To optimize the conditions for Zn2+ and Cu2+ measurements, the following instru- mental parameters were systematically verified: Es in the range 1ñ5 mV, E in the range 10ñ60 mV and tp from 10 to 50 ms. Changes of the potential step (in the given range) cause the increase on peaks cur-
Zn2+

486
W£ODZIMIERZ OPOKA et al.
Figure 2. DP ASV determination of (A) Cu2+ and (D) Zn2+ ions in gastric juice sample using the method of standard addition. Curves: blank;
(0) 0.05 mL of the sample and 5 mL 0.05 M KNO3 with 25 mg fumed silica; (1, 2 and 3) standard addition: 18.8, 37.6 and 56.3 g/L1 Cu2+ or Zn2+. (B) and (E) curves from part (A) and (D) after background correction. (C) and (F) calibration plots for Cu2+ and Zn2+ ions, respec- tively. DP mode: potential step, 2 mV; pulse amplitude, -30 mV; pulse width, 40 ms. Condition of electrode positing: Eacc = -1.1 V; tacc = 30 s for Zn2+ and Eacc = -0.25 V; tacc = 30 s for Cu2+
rent by 8ñ15%. However, increasing Es is accompa- nied by the background current rising. Therefore, it is not suggested to apply the Es > 2 mV. The step potential of 2 mV was applied in further work. The
best results (signal-to-background current ratio) were obtained for the pulse amplitude -30 mV (for tacc = 30 s the peak current for 20 g/L Zn2+ and Cu2+ were ~61 nA and ~37 nA, respectively). The peaks

487
Applicability of the silver amalgam electrode in voltammetric determination of zinc...
Table 1. Concentrations of Zn and Cu in the CRM No. 185, bovine liver determined by means of DP ASV and F-AAS methods (n = 5).
Table 2. Concentrations of Zn and Cu in the rat gastric juice (GJ) samples determined by means of DP ASV and F-AAS methods (n = 5).
Table. 3. Comparison of Zn and Cu concentrations in rat gastric juice (GJ) samples with and without digestion (n = 5).
Table 4. Concentrations of Zn and Cu in the rats gastric mucosa (GM) samples determined by means of DP ASV and AAS methods.
Sample
number
Zn concentration [mg/g]
Cu concentration [mg/g]
F-AAS
DP ASV
ET-AAS
DP ASV
1
13.2
12.9
1.09
0.98
2
14.2
14.3
0.49
0.52
3
14.5
13.6
0.85
0.72
4
12.0
12.2
2.53
2.72
5
12.8
13.3
2.79
2.81
6
12.2
11.8
3.47
2.98
7
14.0
13.4
0.25
0.19
8
14.4
15.9
2.77
2.31
9
13.3
12.7
1.99
2.20

Sample
Zn concentration [mg/L]
Cu concentration [mg/L]
Without digestion
Digested
Without digestion
Digested
A
3.69 ± 0.29
3.41 ± 0.17
2.53 ± 0.15
2.57 ± 0.11
B
39.70 ± 2.78
37.46 ± 1.50
1.12 ± 0.08
1.23 ± 0.07
Sample number
Zn concentration [mg/L]
Cu concentration [mg/L]
DP ASV
F-AAS
DP ASV
F-AAS
1
4.13
4.27
2.80
2.93
2
4.68
4.72
1.68
1.70
3
4.69
4.84
1.37
1.46
4
5.97
6.06
1.28
1.38
5
3.32
3.47
1.71
1.85
6
2.16
2.22
0.77
0.82
Element concentration
DP ASV
F-AAS
Zn [mg/g] (recovery, %)
Certified value: 142 ± 3

139 ± 4 (98)

135 ± 8 (95)
Cu [mg/g] (recovery, %) Certified value: 189 ± 4

187 ± 5 (99)

191 ± 11 (101)

488
W£ODZIMIERZ OPOKA et al.
2.99 ± 0.87x [nA/(g/L)] + 0.18 ± 1.1[nA]. The corre- lation coefficient was r = 0.9997. The relative standard deviation for Zn2+ determination at the concentration 20 g/L was 1.6% (n = 5). The detection limit for Zn2+ following the accumulation time of 30 s, calculated as a 3Û for the blank, was equal to 0.9 g/L.
The calibration graph for Cu2+ and tacc = 30 s was linear from 2 to 100 g/L and obeyed the equa- tion y = 1.86 ± 0.64x [nA/(g/L)] + 0.25 ± 0.8 [nA].
The correlation coefficient was r = 0.9991. The rel-
for zinc are narrower and substantially higher than those of copper. The locations of the peaks are inde- pendent of concentration of metals. The width of the peak at half-height for Zn2+ and Cu2+ are 42 ± 2 and
48 ± 2 mV, respectively. The reproducibility is below 2% for the 5ñ50 g/L concentration range of both metals. The best results (precision, repro- ducibility and the signal-to-background current ratio) were obtained for pulse width 40 ms, and this was the value chosen for the further work. In each case, tw = ts. The height of Zn and Cu peaks strongly decreases with the increasing time tp but simultane-
ously the background current decreases.
ative standard deviation for	determination at the concentration 20 g/L was 2.8% (n = 5). The detection limit for Cu2+ following the accumulation time of 30 s, calculated as a 3Û for the blank, was equal to 1.4 g/L.
For both metals the detection limit can be decreased further by prolonging the accumulation time. In the tested range, 0 s = tacc = 120 s, the rela- tion Ip-tacc is linear, with a sensitivity of ca. 2 nA/s
Cu2+
Influence of the Hg(Ag)FE surface size on zinc and copper peaks
In stripping methods, the peak current is lin- early dependent on the surface area of the working electrode. The surfaces of film electrodes are usual- ly much larger than those of mercury drop elec- trodes. When using the Hg(Ag)FE the surface of the working electrode may easily be varied in a wide range. For a surface area of 1.8 mm2, the peak cur- rent for 10 g/L Zn2+ was 9 nA and grew linearly as the surface of the working electrode increased in size. For a surface area of 11.5 mm2, the peak cur- rent of zinc was 56 nA. The parameters of the linear growth of peak current for zinc and copper vs. sur- face of working electrode are: slope, 0.5 ± 0.02 and
0.3 ± 0.01 [nA/mm2], and correlation coefficient r = 0.999 and r = 0.997, respectively. For further study, a 6 mm2 surface area Hg(Ag)FE was applied.
for 20 g/L	and 1.3 A/s for 20 g/L Cu2+.
Zn2+
Moreover, sensitivity might be twice increased by increasing the electrode surface Hg(Ag)FE from 6 to 12 mm2. Linearity range (upper limit) is the same for both analyzed metals and is equal to 0.05 g/L.
DP ASV analysis of Cu2+ in gastric juice
The GJ sample (0.05 mL) with addition of 5 mL of the supporting electrolyte (0.05 M KNO3) with 25 mg of SiO2 was analyzed by DP ASV under the described conditions. The representative voltam- metric curves for GJ sample are presented in Figure 2A, B and C.
The shape and width of the peak were similar to that obtained from a synthetic solution. The obtained value of standard (10 g/L), based on three repli- cates, was 24.9 ± 1.1 g/L of Cu2+ ions. The detection limit for the determination of copper under these conditions was estimated to be 3 g/L and is limited by the purity of the reagents used in the digestion procedure. The recovery of Cu2+ was tested by addi- tion of 20 g/L of Cu2+. The average recovery was 102 ± 5%. For comparison, the Cu2+ concentration in the sample of gastric juice, measured using HMDE was 22.6 ± 2.5 g/L. The Hg(Ag)FE was applied for Cu2+ determination in the CRM. The obtained results are presented in Table 1. The results of Cu2+ determi- nations by means of DP ASV method in the rat GJ samples are given in Table 2.
Effect of accumulation time and potential
The influence of the accumulation potential (Eacc) was studied in the range from ñ1.15 to ñ0.95 V and -0.35 to -0.15 V with 0.05 M KNO3 spiked with 20 g/L Zn2+ and Cu2+, respectively. The repeatabil- ity and the magnitude of the analytical signal were found to be independent of the accumulation poten- tial in the potential range -1.15 to -1.05 V for zinc and -0.35 to -0.15 V for copper. The accumulation potential ñ1.15 V and -0.25 V was chosen.
The accumulation time (tacc) was changed from 0 to 120 s. The peaks current for supporting elec- trolyte containing 20 g/L Zn2+ and Cu2+ increased linearly with the accumulation times. For further study, the accumulation time of 30 s was chosen. The zinc and copper peaks potential is not depend- ent on either the accumulation time and potential.
DP ASV analysis of Zn2+ in gastric juice
The GJ sample (0.05 mL) with addition of 5 mL of the supporting electrolyte (0.05 M KNO3) with 25 mg of SiO2 was analyzed by DP ASV under the described conditions. The representative voltam-
Calibration graphs
The calibration graph for Zn2+ and tacc = 30 s was linear from 1 to 100 g/L and obeyed the equation y =

489
Applicability of the silver amalgam electrode in voltammetric determination of zinc...
metric curves for GJ sample are presented in Figure 2D, E and F.
The shape and width of the peak were similar to that obtained from a synthetic solution. The obtained value of standard (10 g/L), based on three replicates, was 35.5 ± 1.6 g/L of Zn2+ ions. The detection limit for the determination of zinc under these conditions was 2 g/L and is limited by the purity of the reagents used in the digestion procedure. The recovery of Zn2+ was tested by addition of 20 g/L of Zn2+. The aver- age recovery was 103 ± 4%. For comparison, the Zn2+ concentration in the sample of gastric juice, measured using HMDE was 33.8 ± 2.2 g/L. The Hg(Ag)FE was applied for Zn2+ determination in the CRM. The obtained results are presented in Table 1. The results of Zn2+ determinations by means of DP ASV method in the rat GJ samples are given in Table 2.
Each value is the average of 4 separate deter- minations. It is clear that in the case of Cu determi- nation, the obtained values are in very good agree- ment. In the case of zinc, for both samples a slight loss of the analyte in the digested samples is noted (about 7%). This is connected with the digestion procedure and the volatility of the analyte. What is more, for both elements, analysis of the digested samples gave more precise results. Most probably, although digestion procedure can cause some errors resulting from additional operations, inhomogeneity of raw samples contributes to a decrease of determi- nation precision. Nevertheless, one can accept slightly less precise results in situation when no loss of analyte is expected and analyses might be per- formed much faster and at lower cost (no need to use an expensive equipment and reagents). In trace analysis, speed is also a factor of interest. However, it should be noted that the possibility of analysis of not treated liquid samples (as it was in the case of rat GJ) should always be considered referring to the analyte and the sample composition.
F-AAS analysis of rat gastric juice samples
Generally, one of the essential conditions refer- ring to the analytical sample is its homogeneity. The gastric juice samples did not fulfill this requirement. So, one should expect that some kind of pretreat- ment is required. On the other hand, any additional analytical operations may induce errors, resulting from the use of reagents, instruments and may lead to the sample contamination or loss of analyte, not mentioning time and money consumption.
In the first step of the presented method com- parison, it was checked whether direct analysis of intact sample is possible. In that step, only AAS was considered as the analytical method. Initial meas- urements enabled to choose the flame technique (F- AAS) using air-acetylene flame as adequate tool for Zn analysis. Two samples of the sap, differing in Zn concentrations, were prepared by combining 0.5 mL of 6 samples (total sample volume was 3 mL) and digested in four repetitions. For quantitative deter- mination of Zn and Cu, the Perkin Elmer spectrom- eter Model 3110 (USA) was used. Flame analysis was performed at 213.9 nm, slit 0.7 nm, HCL (Hollow Cathode Lamp). All the measurement parameters (gases flow, burner and lamp position, etc.) were optimized before analysis). Linear range of F-AAS zinc analysis was 0.02ñ0.75 mg/L; detec- tion limit was 0.02 mg/L. Cu determination was made at 324.8 nm, slit 0.7 nm, HCL. The same opti- mization procedure for Cu determination was per- formed. Linear range of copper analysis was 0.10ñ3.00 mg/L; detection limit was 0.10 mg/L. Under given measurement conditions for Zn and Cu determination by means of F-AAS, the results reported in Table 3 were obtained for the rat GJ samples without and after digestion.
F-AAS analysis of rat GM
Solid samples (animal or human tissues) ana- lyzed by the discussed methods in most cases should be digested. Numerous methods can be applied, however, microwave wet digestion in closed sys- tems is usually preferred. GM has been collected in order to measure zinc and copper ion content. Tissue samples from each animal have been put in Eppendorf tubes, and stored frozen and kept at ñ80OC till analysis. The samples were digested in the microwave digestion system Anton Paar Multiwave 3000 (as it was described earlier).
Zinc concentration was determined by means of F-AAS, copper was measured by means of electrother- mal technique (ET-AAS) ñ the Perkin Elmer graphite furnace HGA 600 was used. The furnace program was optimized with the use of ìMethod developmentî pro- gram. Optimal conditions were as follows: atomiza- tion: time 5 s, temperature 2350OC; pretreatment time 30 s, temperature 1050OC; pyro/platform graphite tubes were applied. Cu determination was made at 324.8 nm, slit 0.7 nm, HCL lamp. Detection limit of Cu determi- nation in ET-AAS was 0.5 mg/L. The obtained results of Zn and Cu determinations in the rat GM samples are presented in Table 4.
Estimation of the precision and accuracy of the applied methods
To determine accuracy and precision of the methods, analysis of the certified reference material was performed. The material chosen was bovine

490
W£ODZIMIERZ OPOKA et al.
Figure 3 Correlation between the results of Zn concentrations in the rat gastric juice samples measured by means of DP ASV and FF-AAS methods
Figure 4. Correlation between the results of Cu concentrations in the rat gastric mucosa samples measured by means of DP ASV and FF- AAS methods
liver CRM No. 185 (EU Community Bureau of Reference). The certified value of Zn concentration was 142 ± 3 mg/g and that of Cu was 189 ± 4 mg/g. The values obtained for the digested samples (pro- cedure described earlier in the text) and by means of both compared methods are presented in the Table 1 (for 5 repetitions of analysis).
Both methods are accurate, however, analysis of zinc and copper with the use of AAS are less pre- cise comparing to the ASV method. Low precision of Cu analysis results from the fact, that determina- tions were performed in the manual mode (the sam-
ples were pipetted into the graphite tube manually). Also, lower precision of Zn versus Cu determination for both methods is a result of the possible contami- nation of the samples with the analyte, as zinc is an element abundant in the laboratory environment.
Statistical analysis of the obtained data
From the results presented in Table 2 and 4 and in Figures 3 and 4 it is clear that high level of agree- ment was achieved. It should be noted that precision of measurements is higher for DP ASV determina- tions ñ usually this value is less than 2% (RSD). In

491
Applicability of the silver amalgam electrode in voltammetric determination of zinc...
the case of flame analysis (F-AAS) of not treated samples of rat GJ precision was in the range 4ñ7% (due to inhomogenity of the samples). It may be concluded that both methods may be used for quan- titative determination of Zn and Cu in GJ and GM samples. Obviously it implies, that other biological samples can be analyzed similarly as well. However, AAS method enables analysis of the total element concentration in the sample; with DP ASV one can measure only ionic forms of elements, thus the val- ues measured depend on the sample pretreatment procedures.
4.
Barbarino, F., Toganel E., Brilinschi C.: Clin. Pharmacol. 14, 685 (1992).
Pihan G., Regillo C., Szabo S.: Dig. Dis. Sci. 32, 1395 (1987).
Wong, S.H., CH Cho, C. W. Ogle.: Pharmacology 33, 94 (1986).
Bulbena O., Escolar G., Navarro C., Bravo L., Pfeiffer C.J.: Dig. Dis. Sci. 38, 730 (1993).
Watanabe T., Arakawa T., Fukuda T., Higuchi K., Kobayashi K.: Dig. Dis. Sci. 40, 1340 (1995). Escolar G., Navarro C., Sendros S., Bulbena O.: Arch. Int. Pharmacodyn. 290, 128 (1987).
Cho C.H., Luk C.T., Ogle C.W.: Life Sci. 49, PL189 (1991).
Brzozowski, T.: J. Phys. Pharm. 54 (Suppl. 3), 99 (2003).
Bravo, M.L., Escolar G., Navarro C., Fontarnau R., Bulbena O.: Scanning Microsc. 6, 855 (1992). Yazdanpanah K., Moghimi N., Yousefinejad V., Ghaderi E, Darvishi N.: Pak. J. Med. Sci. 25, 404 (2009).
Sorenson J.R.: Prog. Med. Chem. 26, 437 (1989).
Sorenson J.R.: J. Med. Chem. 27, 1747 (1984). Sorenson, J.R. In: Inflammatory diseases and copper. Sorenson JRJ Ed., p. 289, Humana Press, Clifton, NJ 1982.
Muobarak Jaber F., Tuorkey A., Abdul-Aziz K.K.: Biomed. Pharmacother. 63, 194 (2009). Tuzen M.: Trace Elements and Electrolytes 19, 202 (2002)
Zheng Y., Shougui J.: Guangpuxue Yu Guangpu Fenxi 13, 71 (1993).
B·r·ny E., Bergdahl I. A., Bratteby L-E., et al.: Sci. Total Environ. 286, 129 (2002).
Lavi N., Alfassi Z.B.: Analyst 115, 817 (1990). Bearse R.C., Close D.A., Malanify J.J., Umbarger C.J.: Anal. Chem. 46, 499 (1974).
Ali A.M.M., Farghaly O.A., Ghandour M.A.: Anal. Chim. Acta 412, 99 (2000).
Barthus R.C., Mazo L.H., Poppi R.J.: J. Pharm. Biomed. Anal. 38, 94 (2005).
Kruusma J., Banks C. E., Nei L., Compton R. G.: Anal. Chim. Acta 510, 85 (2004).
Mahajan R.K., Walia T.P., Sumanjit, Lobana T.S.: Talanta 67, 755 (2005).
De Vries W.T., van Dalen E.: J. Electroanal. Chem. 14, 315 (1967).
Wang J.: Stripping Analysis: Instrumentation, and Applications VCH Publishers, Deerfield Beach 1985.
Opoka W., Jakubowska M., Baú B., Sowa- KuÊma M.: Biol. Trace Elem. Res. DOI 10.1007/s12011-010-8790-2 (2010).
5.
6.
7.
8.
9.
10.
CONCLUSIONS
11.
The performed
analysis proved comparable
accuracy of measurements of both methods and higher precision of DP ASV. The DP ASV method enables speciation analysis what can not be made by AAS. Linear range of analysis is much higher for the DP ASV method than for AAS. Detection limit is also much better (of three orders of magnitude) in DP ASV than in AAS determinations.
The following facts were established:
1. DP ASV method and applied renewable silver amalgam film electrode can be readily used for quantitative determination of zinc and copper in gas- tric juice and gastric mucosa of rats;
2. the analytical procedure presented herein offers high precision and accuracy, what was confirmed by the analysis of the certified reference material and the use of an alternative method, i.e., AAS;
3. high precision, accuracy and low detection limits of DP ASV determination of Zn and Cu enable reli- able measurements and observation of subtle changes of the analytes concentrations in complex biological samples;
4. in various instances the presented analytical sys- tem might be indispensible for quantitative analysis of ionic forms of elements (speciation analysis) at a very low, ppt level of concentration;
5. the use of the Hg(Ag)FE reduces substantially mercury consumption during measurements and analyst contact with this toxic element, comparing to the mercury drop electrode.
12.
13.
14.
15.
16.
17.
18.
19.
20.
21.
22.
23.
24.
25.
26.
REFERENCES
27.
1. Kadakia S.C., Wong R.K., Maydonovitch C.L., Nelson N.R., Henkin R.I.: Dig. Dis. Sci. 37, 513 (1992).
2. Mann N.S., Mann S.K., Brawn P.N., Weaver B.: Digestion 53, 108 (1992).
3. Parkin G.: Chem. Rev. 104, 699 (2004).
28.
29.

492
W£ODZIMIERZ OPOKA et al.
30.
Hoyer B., Florence T.M.: Anal. Chem. 59, 2839 (1987).
Tripathi R.M., Raghunath R., Mahapatra S., Sadasivan S.: Sci. Total Environ. 277, 161 (2001).
Moreno M.A., Marin C., Vinagre F., Ostapczuk P.: Sci. Total Environ. 229, 209 (1999).
Florence T.M.: J. Electroanal. Chem. 27, 273 (1970).
Wang J., Lu J., Hocevar S., Farias P., Ogorevc B.: Anal. Chem. 72, 3218 (2000).
Crew A., Cowell D.C., Hart J.P.: Talanta 75, 1221 (2008).
Lovric M., Komorsky-Lovric ä., Bond A.M.: J. Electroanal. Chem. 319, 1 (1991).
Baú B., Kowalski Z.: Electroanalysis 14, 1067 (2002).
Jakubowska M., Kubiak W.W.: Anal. Chim. Acta 512, 241 (2004).
39.
Jakubowska M., Kubiak W.W.: Electroanalysis 17, 1687 (2005).
Jakubowska M.: Electroanalysis 22, 564
(2010).
Polish Patent No. P-319 984 (1997).
Baú B.: Anal. Chim. Acta 570, 195 (2006). Jakubowska M., Baú B., Ciepiela F., Kubiak W.W.: Electroanalysis 22, 1757 (2010).
Jakubowska M.: J. Hazard. Mat. 176, 540 (2010).
Kowalski, Z., Ko≥der E., Niewiara E.: Euroanalysis V, KrakÛw, Book of Abstracts I- 46 (1984).
Kubiak W.W., Kowalski Z.: Anal. Chem. 61, 1598 (1989).
Niewiara	E.,	Baú	B.,	Kubiak	W.W.: Electroanalysis 19, 2192 (2007).
31.
40.
41.
42.
43.
32.
33.
44.
34.
45.
35.
36.
46.
37.
47.
38.
Received: 14. 02. 2011