International Conference on Particle Physics and Astrophysics (ICPPA-2015)	IOP	Publishing
Journal of Physics: Conference Series 675 (2016) 012035	doi:10.1088/1742-6596/675/1/012035
# The high precision measurement of the 144Ce activity in the SOX experiment
L Di Noto1,31 on behalf of the SOX collaboration,
M Agostini2, K Althenmiiller2, S Appel2, G Bellini3, J Benziger4, N
Berton5, D Bick6, G Bonfini7, D Bravo—Berguno8, B Caccianiga3, F
Calaprice9, A Caminata1, P Cavalcante7, R Cereseto1, A
Chepurnov10, K Choi11, M Cribier5, D DAngelo3, S Davini12, A
Derbin13, I Drachnev12, M Durero5, A Etenko14, S Farinon1, V
Fischer5, K Fomenko15, D Franco16, F Gabriele7, J Gaffiot5, C
Galbiati9, C Ghiano1, M Giammarchi3, M Goeger-Neff2, A Goretti9,
M Gromov10, C Hagner6, Th Houdy5, E Hungerford17, Aldo Ianni7,
Andrea Ianni9, N Jonqueres 18, K Jedrzejczak19, M Kaiser6, V
Kobychev20, D Korablev15, G Korga7, V Kornoukhov21, D Kryn16, T
Lachenmaier22 T Lasserre5, M Laubenstein7, T Lehnert23, J Link8 E
Litvinovich14,24, F Lombardi7, P Lombardi3, L Ludhova3, G
Lukyanchenko14,24, I Machulin14,24, S Manecki8, W Maneschg25, S
Marcocci12, J Maricic11, G Mention5, E Meroni3, M Meyer6, L
Miramonti3, M Misiaszek7 19, M Montuschi26, P Mosteiro9, V
Muratova 13, R Musenich1, B Neumair2, L Oberauer2, M
Obolensky16, F Ortica27, M Pallavicini1, L Papp2, L Perasso1 , A
Pocar28, G Ranucci3, A Razeto7, A Re3, A Romani27, R Roncin7,16, C
Rossi1, N Rossi7, S Schonert2, L Scola5, D Semenov13 H Simgen25, M
Skorokhvatov14,24, O Smirnov15, A Sotnikov15, S Sukhotin14, Y
Suvorov14,29, R Tartaglia7, G Testera1, J Thurn23, M Toropova14, C
Veyssiere5, M Vivier5, E Unzhakov13, R B Vogelaar8, F
von Feilitzsch2, H Wang29, S Weinz30, J Winter30, M Wojcik19, M
Wurm 6,30, Z Yokley8, O Zaimidoroga15, S Zavatarelli1, K Zuber23 and
G Zuzel19
1. Dipartimento di Fisica, Universita degli Studi e INFN 16146 Genova, Italy
2. Physik-Department and Excellence Cluster Universe, Technische Universitat Miinchen,
85748 Garching, Germany
1. Dipartimento di Fisica, Universita degli Studi e INFN, 20133 Milano, Italy
2. Chemical Engineering Department, Princeton University, Princeton, NJ 08544, USA
3. Commisariat a l’Energie Atomique et aux Energies Alternatives, Centre de Saclay, IRFU,
   91191 Gif-sur-Yvette, France
4. Institut fur Experimentalphysik, Universitat, 22761 Hamburg, Germany
5. INFN Laboratori Nazionali del Gran Sasso, 67010 Assergi (AQ), Italy
6. Physics Department, Virginia Polytechnic Institute and State University, Blacksburg, VA
   24061, USA
7. Physics Department, Princeton University, Princeton, NJ 08544, USA
   31 To whom any correspondence should be addressed.
Content from this work may be used under the terms of the Creative Commons Attribution 3.0 licence. Any further distribution
of this work must maintain attribution to the author(s) and the title of the work, journal citation and DOI.
1
Published under licence by IOP Publishing Ltd

International Conference on Particle Physics and Astrophysics (ICPPA-2015)	IOP	Publishing
Journal of Physics: Conference Series 675 (2016) 012035	doi:10.1088/1742-6596/675/1/012035
1. Lomonosov Moscow State University Skobeltsyn Institute of Nuclear Physics, 119234
   Moscow, Russia
2. Department of Physics and Astronomy, University of Hawaii, Honolulu, HI 96822, USA
3. Gran Sasso Science Institute (INFN), 67100 L’Aquila, Italy
4. St. Petersburg Nuclear Physics Institute NRC Kurchatov Institute, 188350 Gatchina,
   Russia
5. NRC Kurchatov Institute, 123182 Moscow, Russia
6. Joint Institute for Nuclear Research, 141980 Dubna, Russia
7. AstroParticule et Cosmologie, Universite Paris Diderot, CNRS/IN2P3, CEA/IRFU,
   Observatoire de Paris, Sorbonne Paris Cite, 75205 Paris Cedex 13, France
8. Department of Physics, University of Houston, Houston, TX 77204, USA
9. Commisariat a l’Energie Atomique et aux Energies Alternatives, Centre de Saclay,
   DEN/DM2S/SEMT/BCCR, 91191 Gif-sur-Yvette, France
10. M Smoluchowski Institute of Physics, Jagiellonian University, 30059 Krakow, Poland
11. Kiev Institute for Nuclear Research, 06380 Kiev, Ukraine
12. Institute for Theoretical and Experimental Physics, 117218 Moscow, Russia
13. Kepler Center for Astro and Particle Physics, Universitat Tubingen, 72076 Tubingen,
    Germany
14. Department of Physics, Technische Universitat Dresden, 01062 Dresden, Germany
15. National Research Nuclear University MEPhI (Moscow Engineering Physics Institute),
    Kashirskoe highway 31, Moscow, 115409, Russia
16. Max-Planck-Institut fur Kernphysik, 69117 Heidelberg, Germany
17. Dipartimento di Fisica e Scienze della Terra Universita degli Studi di Ferrara e INFN, Via
    Saragat 1 - 44122, Ferrara, Italy
18. Dipartimento di Chimica, Universita e INFN, 06123 Perugia, Italy
19. Amherst Center for Fundamental Interactions and Physics Department, University of
    Massachusetts, Amherst, MA 01003, USA
20. Physics and Astronomy Department, University of California Los Angeles (UCLA), Los
    Angeles, California 90095, USA
21. Institute of Physics and Excellence Cluster PRISMA, Johannes Gutenberg-Universitaat
    Mainz, 55099 Mainz, Germany
E-mail: lea.dinoto@ge.infn.it
Abstract. In order to perform a resolutive measurement to clarify the neutrino anomalies
and to observe possible short distance neutrino oscillations, the SOX (Short distance neutrino
Oscillations with BoreXino) experiment is under construction. In the first phase, a 100 kCi
144Ce-144Pr antineutrino source will be placed under the Borexino detector at the Laboratori
Nazionali del Gran Sasso (LNGS), in center of Italy, and the rate measurement of the
antineutrino events, observed by the very low radioactive background Borexino detector, will be
compared with the high precision (< 1%) activity measurement performed by two calorimeters.
The source will be embedded in a 19 mm thick tungsten alloy shield and both the calorimeters
have been conceived for measuring the thermal heat absorbed by a water flow. In this report
the design of the calorimeters will be described in detail and very preliminary results will be
also shown.
## 1. Introduction
Since the neutrino anomalies leave room to the existence of sterile neutrino and to the short
distance neutrino oscillations [1-3], the SOX experiment was proposed [4] with the goal of
observing the neutrino events emitted by a high activity 144Ce-144Pr artificial source. The
source will be placed in a tunnel, 8 m far from the center of Borexino detector and thanks to the
very low radioactive background and to the detector properties [5,6], short distance oscillations
of eV mass sterile neutrino might be observed and at the same time a disappearance experiment
can be performed, if the source activity is also measured with less than 1% precision [4,7].
2

International Conference on Particle Physics and Astrophysics (ICPPA-2015)	IOP	Publishing
Journal of Physics: Conference Series 675 (2016) 012035	doi:10.1088/1742-6596/675/1/012035
Figure 1. a) Section of the source with the tungsten alloy shield embedded inside the INFN-
TUM calorimeter. The main parts of the apparatus are clearly visible. b) Preliminary power
measurement performed by using a mock-up electrical source. In the picture the parameter as
resulted by fitting the data with the exponential function of equation 2, are reported as well.
As the size of the source embedded in the shield is really close to the available room in the
tunnel, the project will be developed in two different phases: the calorimetric measurement of the
activity, just before the insertion of the source inside the pit under Borexino (1 week duration)
and the second phase, when the source will be installed inside the tunnel and the Borexino data
taking will occur (1.5 year duration). Other calorimetric measurements in the between or at the
end of the data taking might be also performed in order to increase the accuracy of the final
results.
# 2. Description of the calorimetric measurement
The 144Ce-144Pr artificial source will be produced with around 100 kCi activity at the reactor
of the Mayak Production Association in Russia after a chemical extraction of the 144CeO2 from
exhausted nuclear fuel. The radioactive powder will be sealed inside a double stainless steel
container and then embedded in a 19 mm thick safety shield made of a tungsten alloy (see figure
1a). After the production and the transportation to the LNGS, the calorimetric measurement
will occur thanks to two different calorimeters, that are under construction and testing by the
INFN-TUM and the CEA groups within the SOX collaboration. Both the calorimeters have been
conceived to measure the source activity with high precision by knowing the power released by
the radiation in the tungsten shield and absorbed by a water flow.
In the INFN-TUM calorimeter the water flows inside a copper heat exchanger, in contact with
the tungsten shield (see figure 1a), while in the CEA calorimeter the source with the tungsten is
directly immersed in the water inside the calorimeter chamber. In both the systems, neglecting
the heat losses Piost, the power P is achieved by measuring with high accuracy the mass flow
m and the temperatures Tout and Tin of the water outgoing and entering in the copper heat
exchanger, according to the relation:
where c is the specific heat of the water averaged in the Tin-Tout temperature range. As it will
be shown in the following, the calorimeters have been designed in order to minimize Piost that
will be estimated during the calibration and testing phase and reduced as much as possible.
In particular in the INFN-TUM calorimeter, for reducing the convection, the copper heat
exchanger is kept in vacuum at a pressure lower than 5 • 10-5 mbar, while the source inside
the tungsten, separated through a vacuum flange, is at atmospheric pressure. Secondly for
minimizing irradiation, two stages of 10 layers each one of super insulators are placed between
the copper and the vacuum tank and at the same time hot water flows inside a pipe built
directly on the tank itself (see figure 1a) in order to thermalize the chamber close to the copper
3
P = mc(Tout - Tin)	(1)

International Conference on Particle Physics and Astrophysics (ICPPA-2015)
IOP Publishing
Journal of Physics: Conference Series 675 (2016) 012035
doi:10.1088/1742-6596/675/1/012035
temperature. Finally for reducing the conduction, the source inside the copper heat exchanger
is placed on an hanging platform, suspended by three kevlar ropes with low heat conductivity.
In both the systems a closed water loop is foreseen and in the INFN-TUM calorimeter the water
loop was designed in order to accurately control the massflow and the temperature of the water
just before the entrance in the vacuum tank. Thanks to a proportional valve controlled by a
feedback loop, the water flow, coming from the chiller, can be tunable in the range of 2-15 g/s
and stabilized at 0.3% of the nominal value, heavily reducing the oscillations due to the chiller
pump. At the same time also the temperature of the water entering inside the calorimeter can
be stabilized within 0.1 K in the 13-17 °C range, compensating any environmental temperature
fluctuation or chiller instability during all the phases of the measurement. In particular during
the calorimetric measurement a massflow of 10 g/s was chosen, giving a difference of temperature
of Tout — Tin ~ 28 °C, for a source power of 1200 W correspondent to 100 kCi activity.
On the contrary, during the second phase, since the source will be inserted in the pit inside
the copper heat exchanger of the INFN-TUM calorimeter, the water loop will be necessary to
cool down the apparatus and the controls feedback will be aimed to stabilize the temperature in
the tunnel in order to prevent turbulence effect in the scintillator liquid of the Borexino detector.
# 3. Preliminary results and conclusions
In the calibration phase, before the source delivery, the set up can be tested with a mock-up
source made by a copper cylinder with the same dimensions of the source container, embedded
in an Aluminum body that replaces the tungsten shield. Inside the copper cylinder six electrical
heaters are placed and they can be turned on with a total maximum power of 1200 W (the
same power dissipated by a 100 kCi Ce source), measured very precisely by an amperometer
and a voltmeter directly connected with the heaters. The temperature of the copper and of
the aluminum part (on the top, on the side and on the bottom) can be monitored and the
temperatures of many critical points inside the vacuum chamber (as the platform, the copper
jacket and the internal chamber) are acquired as well.
As first test a rough measurement performed with a simple copper coil directly at contact
with the aluminum was done at atmospheric pressure with a set power of 497 W. The plot of
the measured power calculated by the relation 1 as a function of the time is shown in figure 1b.
The data were acquired after the heaters switching on and the fit was done with the exponential
function:
/	x-P2 \
P(t) = po h — e piJ	(2)
where the time scale is in UNIX time date format. From the fit the measured power at the
equilibrium was found around 491 W, less than the set value since the heat losses due to the
convection can not be neglected and a first estimation of the time constant p1 ~ 9735 s = 2, 7
hours was achieved for the first time.
In conclusion two calorimeters have	been built for measuring the	source	activity	with	a
precision <	1% and now they are under	testing. In the next months many calibrations	will	be
performed with the mock-up source, while the final measurement with the Ce source and the
starting of the Borexino data taking is fixed in the autumn 2016.
# References
1. Mention G et al. 2011 Phys. Rev. D 83 073006
2. Aguilar A	et al. (LSND Collaboration) 2001	Phys. Rev. D 64112007
3. Aguilar A	et al. (MiniBooNe Collaboration)	2013 Phys. Rev. Lett. 110 161801
4. Bellini	G	et al.	(Borexino	Collaboration)	2013 JHEP	08 038
5. Bellini	G	et al.	(Borexino	Collaboration)	2014 Phys.	Rev. D 89	112007
6. Bellini	G	et al.	(Borexino	Collaboration)	2010 Phys.	Lett. B 687	299
7. Palazzo A 2013	Mod. Phys. Lett. A 28 1330004
4